## Sheet1
prob_id | created_date | solve_hours | micro | unit | status | status_chk | prob_type_chk | prob_type_detail | description
133417 | 42661.99 | 0.75 | LI23 | 41 | 4 | HARDWARE | HW Problem | Klys 23-4 has AMJ and PHJ problems that are pretty visible on the beam
132731 | 42593.6022 | 0.5 | LI27 | 21 | 4 | HARDWARE | HW Problem | 27-2 CAM status on scp
132923 | 42615.136 | 1.5 | LI22 | 71 | 4 | HARDWARE | HW Problem | 22-7 is experiencing large phase excursions. See https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=860967
132930 | 42615.9412 | 1.5 | LI27 | 41 | 4 | HARDWARE | HW Problem | Klystron 27-4 FLK without remote reset. When reset the station repeatedly trips on EVOC and goes to lock out.
131833 | 42517.8302 |  | LI22 | 61 | 4 | HARDWARE | HW Problem | 22-6 has had many vacuum faults since 5/27 turn on.
131871 | 42520.2464 |  | LI24 | 11 | 4 | HARDWARE | HW Problem | 24-1 PAD shows NaN phase and 0 ampl after power dip. Cannot use this in feedback.attempted IPL'ing PIOP, no luck and I could not locate the PAD in EPICS.
131871 | 42520.2464 |  | LI24 | 11 | 4 | HARDWARE | HW Problem | 24-1 PAD shows NaN phase and 0 ampl after power dip. Cannot use this in feedback.attempted IPL'ing PIOP, no luck and I could not locate the PAD in EPICS.
132130 | 42540.4729 | 1 | LI24 | 61 | 4 | HARDWARE | HW Problem | RFS 24-6 FLK no remote reset
132419 | 42564.3357 | 1 | LI23 | 11 | 4 | HARDWARE | HW Problem | LI23/25/27/29 LCLS KLystron bulk supplies have very dirty and clogged air filters that need to be replace
132423 | 42564.3441 | 1 | LI27 | 21 | 4 | HARDWARE | HW Problem | 27-2 Kly trim supply power on lamp is burnt out and should be replace (pref with x10 which is lighter supply)
132452 | 42564.6336 | 1 | LI30 | 71 | 4 | HARDWARE | HW Problem | 30-7: BV OOT at 344kV. Adjust to 350kV. Check DQ'g current and adjust as needed.
132524 | 42571.5663 | 1 | LI20 | 81 | 4 | HARDWARE | HW Problem | 20-8 Has 3 overheated D-Qing resistor jumpers that needs replace on the north side.
132562 | 42573.7118 | 0.5 | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X (21-2) has increased phase jitter, possibly only while running the day exp't configuration. Jumping from 0.12 degX to over 0.20 degX. First upstream PAD/PAC signal with similar increase in instability is the klystron BV (KLYS:LI21:K2:L1X_2_S_SACT). Went away quickly when night config was loaded. First increase 07/21 at 11:00 drops off at 21:05. Increases quickly again 07/22 at 09:30 so far pattern same as yesterday. http://physics-elog.slac.stanford.edu/lclselog/show.jsp?dir=/2016/29/22.07&pos=2016-07-22T16:58:54
132729 | 42593.594 |  | LI21 | 71 | 4 | HARDWARE | HW Problem | 21-7 tripping frequently on HV over current faults.
132740 | 42594.3136 | 1 | LI29 | 51 | 4 | HARDWARE | HW Problem | station 29-5 is showing 4 Water Faults in last 24 hours on Error count
132748 | 42594.7331 | 1.5 | LI24 | 21 | 4 | HARDWARE | HW Problem | Klystron LI24-21 has tripped off repeatedly over the last few hours. Upcoming program requires all klystrons to be available and stable.
132787 | 42599.3854 |  | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3: Beam voltage lowered due to IKOC/EOLC faults.
133500 | 42671.0085 | 1.4 | LI24 | 51 | 4 | HARDWARE | HW Problem | 24-5 has bad phase jitter.
133548 | 42676.2628 | 0.75 | LI22 | 71 | 4 | HARDWARE | HW Problem | station 22-7 has growing BeamVolts jitter
133653 | 42684.373 | 2.5 | LI24 | 81 | 4 | HARDWARE | HW Problem | 24-8 install and runup new deuterium thyratron from spares inventory, update depot
133800 | 42702.0901 | 1 | LI24 | 21 | 4 | HARDWARE | HW Problem | 24-2 is MOD faulted. Touch panel has a "Host Communication err" popup and is unresponsive.
133804 | 42702.4133 |  | LI20 | 81 | 4 | HARDWARE | HW Problem | 20-8 (L0B) Core bias power supply experienced current noise after Nov 13th. Open the cabinet and inspect possible noise source (arcing somewhere) including pulse cable.
133806 | 42702.4664 | 2.5 | LI20 | 71 | 4 | HARDWARE | HW Problem | Klystron 20-7 trips incessantly on MOD fault lockout. We can reset it temporarily, but it trips off again a minute later.
133850 | 42704.9823 | 1 | LI22 | 61 | 4 | HARDWARE | HW Problem | 22-6 Showing frequent bouts of intense phase jitter. Virtually unusable.
133088 | 42635.4076 | 1.5 | LI25 | 41 | 4 | HARDWARE | HW Problem | 25-4 PIOP shows Focus Monitor out of tols. Jumps from 12.97 to 13.04. Checked FIAT rack t.p. and it looks very steady. MKSU?
133106 | 42637.2737 | 1.5 | LI20 | 51 | 4 | HARDWARE | HW Problem | 20-5 FLK no reset.
133176 | 42642.6034 | 1.5 | LI16 | 31 | 4 | HARDWARE | HW Problem | ION pump controller on spare klystron at 16-3 is dead
133204 | 42644.9759 | 1 | LI23 | 21 | 4 | HARDWARE | HW Problem | klys 23-2 tripped, no reset.
131887 | 42521.4778 | 1 | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7: Seeing some beam volts jitter.
132103 | 42537.4782 | 2 | LI30 |  | 4 | HARDWARE | HW Problem | 30-6: Seeing some beam volts jitter.
132210 | 42544.2932 | 1.5 | LI30 | 71 | 4 | HARDWARE | HW Problem | 30-7 mod fault
132235 | 42547.464 | 2.5 | LI30 | 71 | 4 | HARDWARE | HW Problem | Klys 30-7 keeps going into heater delay.
132510 | 42570.8262 | 1 | LI22 | 51 | 4 | HARDWARE | HW Problem | 22-5 tripping frequently. HVR faults.
132510 | 42570.8262 | 1 | LI22 | 51 | 4 | HARDWARE | HW Problem | 22-5 tripping frequently. HVR faults.
132197 | 42543.6122 |  | LI26 | 21 | 4 | HARDWARE | HW Problem | 26-2 MOD fault, no remote reset.
132410 | 42564.1603 |  | LI24 | 81 | 4 | HARDWARE | HW Problem | 24-8 window temp (KLYS:LI24:81:WNDW2_TEMP) has a low alarm. The temperature oscillates, at its lowest it is just barely lower than the limit (current temp is 90.5 degF, limit is 91).
132459 | 42565.3109 |  | LI20 | 81 | 4 | HARDWARE | HW Problem | 20-8: BV OOT at 346kV. Adjust to 350kV. Check DQ'g current and adjust as needed.
132544 | 42572.5182 |  | LI24 | 21 | 4 | HARDWARE | HW Problem | Time Accounting: Klys 24-1 AMPL dropped and moved for about 15 minutes. Recovered on its own, problem was not easily found. EDIT (7/22 10:36 BRIPMAN) - This problem returned the following day and was traced to large amplitude jitter on 23-1. We suspect that the jumps observed on 24-1 and/or 24-2 were caused by the longitudinal feedback reacting to 23-1 amplitude jumps. After taking the station off the beam and waiting a few minutes, the amplitude jitter is no longer visible. Either it went away on its own or it was somehow related to having the station on accelerate.
132544 | 42572.5182 |  | LI24 | 21 | 4 | HARDWARE | HW Problem | Time Accounting: Klys 24-1 AMPL dropped and moved for about 15 minutes. Recovered on its own, problem was not easily found. EDIT (7/22 10:36 BRIPMAN) - This problem returned the following day and was traced to large amplitude jitter on 23-1. We suspect that the jumps observed on 24-1 and/or 24-2 were caused by the longitudinal feedback reacting to 23-1 amplitude jumps. After taking the station off the beam and waiting a few minutes, the amplitude jitter is no longer visible. Either it went away on its own or it was somehow related to having the station on accelerate.
132566 | 42574.0817 | 1 | LI27 | 81 | 4 | HARDWARE | HW Problem | 27-8 MOD fault immediately after being turned on.
133419 | 42662.1375 |  | LI22 | 71 | 4 | HARDWARE | HW Problem | 22-7 exhibiting single shot phase excursions.
133513 | 42671.8456 | 1 | LI24 | 51 | 4 | HARDWARE | HW Problem | 24-5 ACOC faults on turn on.
133513 | 42671.8456 | 0.2 | LI24 | 51 | 4 | HARDWARE | HW Problem | 24-5 ACOC faults on turn on.
131838 | 42518.6175 | 1.5 | LI21 | 21 | 4 | HARDWARE | HW Problem | Kly Mod 21-2 (L1X): bouts of drop outs and phase jitter.
132216 | 42544.7615 | 2 | LI25 | 31 | 4 | HARDWARE | HW Problem | Klystron LI25-31 amplitude is bouncing between 25 and 69. Behavior started around 3:35 pm and has continued until now (2.5 hrs.). See MCC elog entry for 06/23/2016 18:13 for plot.
132486 | 42568.8937 | 1 | LI22 | 71 | 4 | HARDWARE | HW Problem | 22-7 FTPs show occasional amplitude dropouts (nothing visible on beam volts).
133030 | 42628.6241 | 0.5 | LI19 | 11 | 4 | HARDWARE | HW Problem | 19-1B klystron 617A shows some elevated vacuum - reportedly 1x10^-7
132880 | 42610.5193 |  | LI28 | 51 | 4 | HARDWARE | HW Problem | 28-5 is tripping on "ACC 1" water faults (external faults on mod)
132882 | 42610.7771 |  | LI21 | 61 | 4 | HARDWARE | HW Problem | Station 21-6 has a CAMAC fault; does not come back up after trying reset of PIOP or after IPL.
132882 | 42610.7771 | 1.5 | LI21 | 61 | 4 | HARDWARE | HW Problem | Station 21-6 has a CAMAC fault; does not come back up after trying reset of PIOP or after IPL.
132883 | 42611.2633 |  | LI29 | 21 | 4 | HARDWARE | HW Problem | 29-2 faulted many times on 8/29 owl shift. Occasionally showed external fault but generally no indication from ACR of the cause.
132883 | 42611.2633 | 0.75 | LI29 | 21 | 4 | HARDWARE | HW Problem | 29-2 faulted many times on 8/29 owl shift. Occasionally showed external fault but generally no indication from ACR of the cause.
132896 | 42611.9344 | 3 | LI27 | 41 | 4 | HARDWARE | HW Problem | 27-4 TBR persistent klystron heater delay. OPS swapped fuse supplying power to heater circuit under phone guidance from PEM Hovey/Tagle. EDIT: After swapping fuse and letting time out, it now trips off almost immediately when trying to turn the modulator on. Beam volts look very high when it is briefly on, if I trust the readback.
132896 | 42611.9344 |  | LI27 | 41 | 4 | HARDWARE | HW Problem | 27-4 TBR persistent klystron heater delay. OPS swapped fuse supplying power to heater circuit under phone guidance from PEM Hovey/Tagle. EDIT: After swapping fuse and letting time out, it now trips off almost immediately when trying to turn the modulator on. Beam volts look very high when it is briefly on, if I trust the readback.
132896 | 42611.9344 | 5 | LI27 | 41 | 4 | HARDWARE | HW Problem | 27-4 TBR persistent klystron heater delay. OPS swapped fuse supplying power to heater circuit under phone guidance from PEM Hovey/Tagle. EDIT: After swapping fuse and letting time out, it now trips off almost immediately when trying to turn the modulator on. Beam volts look very high when it is briefly on, if I trust the readback.
132143 | 42541.2968 | 0.7 | LI27 | 21 | 4 | HARDWARE | HW Problem | 27-2 Had a water-leak that rendered the PAD unfunctional. Need to replace PAD and run a new POLY.
132523 | 42571.565 | 1 | LI20 | 61 | 4 | HARDWARE | HW Problem | 20-6 Has overheated dequing resistor long wire the drops to the inside of the modulator that needs replace.
132565 | 42573.9133 | 1 | LI25 | 71 | 4 | HARDWARE | HW Problem | 25-7 tripped on FLK; trips almost immediately when we turn back on.
132745 | 42594.4831 | 1.5 | LI25 | 61 | 4 | HARDWARE | HW Problem | 25-6 is tripping on HVOC.
132873 | 42608.5139 | 1 | LI21 | 61 | 4 | HARDWARE | HW Problem | Klys 21-6 trips at least 3 or more times an hour. Has been this way for at least 2 months.
131732 | 42509.596 | 1 | LI25 | 41 | 4 | HARDWARE | HW Problem | 25-4 Mod has a burnt/broken Carborundum resistor
131831 | 42517.8063 | 0.5 | LI23 | 41 | 4 | HARDWARE | HW Problem | 23-4 has had lots of over-voltage faults after 5/27 turn on.
131853 | 42519.3735 |  | LI23 | 41 | 4 | HARDWARE | HW Problem | Klys 23-4 tripped, no remote reset: station still in local
131892 | 42521.559 | 2 | LI22 | 31 | 4 | HARDWARE | HW Problem | 22-3 PFN needs tuning. Slope -0.8
132248 | 42548.8253 | 1 | LI25 | 21 | 4 | HARDWARE | HW Problem | Klystron LI25-21 trips off on EOLC fault immediately after reset and turn on.
132985 | 42622.2165 | 2 | LI25 | 61 | 4 | HARDWARE | HW Problem | 25-6 faulting on over-voltage.
132912 | 42614.3115 | 1 | LI30 | 41 | 4 | HARDWARE | HW Problem | 30-4: DeQing current low = 0.5A.
132922 | 42614.9982 | 2 | LI28 | 41 | 4 | HARDWARE | HW Problem | Klys 28-4 reads "MOD," but other numbers look good and appears to be contributing to the beam. Tried IPL-ing, with no luck
132937 | 42618.294 |  | LI27 | 41 | 4 | HARDWARE | HW Problem | 27-4 is tripped, can't reset remotely
132937 | 42618.294 | 2 | LI27 | 41 | 4 | HARDWARE | HW Problem | 27-4 is tripped, can't reset remotely
132987 | 42622.2828 | 0.5 | LI22 | 71 | 4 | HARDWARE | HW Problem | 22-7 is experiencing amplitude dropouts about once every 3-4 seconds.
132598 | 42577.5355 | 1 | LI29 | 71 | 4 | HARDWARE | HW Problem | 29-7: T20 Compartment fan fault scrolling on PLC fault log .
133672 | 42685.7511 |  | LI29 | 71 | 4 | HARDWARE | HW Problem | Klys 29-7 suddenly showing frequent large beam volts dropout. Wouldn't be surprised if it goes away on its own at some point as it started very suddenly.
133719 | 42691.2767 | 1 | LI27 | 41 | 4 | HARDWARE | HW Problem | Klystron 27-4 and 29-1 MOD faulted and won't reset remotely
131897 | 42521.6392 | 1.5 | LI23 | 41 | 4 | HARDWARE | HW Problem | 23-4 uk came back low after turn on. Please check klystron heater values.
131898 | 42521.6755 | 2 | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 L1X periods of high AMPL/PHAS jitter (particularly when hot) Decker analysis shows this is coming from the RF drive/TWT.
132389 | 42562.6219 | 2 | LI21 | 51 | 4 | HARDWARE | HW Problem | station 21-5 Klystron Vacuum Fault
132469 | 42565.5913 | 1 | LI30 | 51 | 4 | HARDWARE | HW Problem | 30-5 tripping on External interlock when on the beam. Have seen KMC faults on the MKSU when this happens.
132469 | 42565.5913 | 2 | LI30 | 51 | 4 | HARDWARE | HW Problem | 30-5 tripping on External interlock when on the beam. Have seen KMC faults on the MKSU when this happens.
132469 | 42565.5913 | 1 | LI30 | 51 | 4 | HARDWARE | HW Problem | 30-5 tripping on External interlock when on the beam. Have seen KMC faults on the MKSU when this happens.
133107 | 42637.5718 | 1 | LI22 | 81 | 4 | HARDWARE | HW Problem | Klystron 22-8 24V battery fault
133175 | 42642.5619 |  | LI22 | 71 | 4 | HARDWARE | HW Problem | 22-7 BV FTPs not working (show 0 BV, station is at nominal voltage).
133000 | 42625.3103 | 2 | LI27 | 31 | 4 | HARDWARE | HW Problem | Klys 27-3 amplitude went from 54 -> 47 in a discrete step at 06:30
133021 | 42628.199 | 1 | LI27 | 41 | 4 | HARDWARE | HW Problem | Klys 27-4 keeps going FLK
133021 | 42628.199 | 3 | LI27 | 41 | 4 | HARDWARE | HW Problem | Klys 27-4 keeps going FLK
133079 | 42633.7792 | 2 | LI22 | 71 | 4 | HARDWARE | HW Problem | Klys 22-7 phase trim not zeroing phase. I can see the KPHR move the amount one would expect, but the phase readback doesn't change.
133131 | 42639.6736 | 1 | LI28 | 21 | 4 | HARDWARE | HW Problem | Klys 28-2 FLK, no reset.
133131 | 42639.6736 | 1 | LI28 | 21 | 4 | HARDWARE | HW Problem | Klys 28-2 FLK, no reset.
133167 | 42642.1691 | 1.5 | LI21 | 31 | 4 | HARDWARE | HW Problem | TBR Kly. LI 21 - 3 tripped on "MSKU" fault. Checking the fault it show the drive fault light is on.
133167 | 42642.1691 | 5 | LI21 | 31 | 4 | HARDWARE | HW Problem | TBR Kly. LI 21 - 3 tripped on "MSKU" fault. Checking the fault it show the drive fault light is on.
133173 | 42642.5432 | 1.2 | LI23 | 41 | 4 | HARDWARE | HW Problem | 23-4 has high beam volts jitter.
133179 | 42642.6316 |  | LI24 | 31 | 4 | HARDWARE | HW Problem | Klys 24-3 has been tripping frequently on this shift
133185 | 42643.0247 |  | LI22 | 71 | 4 | HARDWARE | HW Problem | 22-7 FTPs for BV and BI show enormous (false) shot to shot jitter.
133211 | 42645.5991 | 0.5 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 tripping to HVR
131731 | 42509.592 | 1 | LI25 | 51 | 4 | HARDWARE | HW Problem | 25-5 Mod has a broken Despiking coil resistor
131981 | 42528.2021 | 7 | LI30 | 71 | 4 | HARDWARE | HW Problem | Klys 30-7 has frequent HVR faults
132030 | 42531.332 | 1 | LI25 | 81 | 4 | HARDWARE | HW Problem | Klystron 25-8 MOD HVR incessantly when placed on the beam
132030 | 42531.332 | 1 | LI25 | 81 | 4 | HARDWARE | HW Problem | Klystron 25-8 MOD HVR incessantly when placed on the beam
132030 | 42531.332 | 1 | LI25 | 81 | 4 | HARDWARE | HW Problem | Klystron 25-8 MOD HVR incessantly when placed on the beam
132030 | 42531.332 |  | LI25 | 81 | 4 | HARDWARE | HW Problem | Klystron 25-8 MOD HVR incessantly when placed on the beam
132030 | 42531.332 |  | LI25 | 81 | 4 | HARDWARE | HW Problem | Klystron 25-8 MOD HVR incessantly when placed on the beam
132192 | 42543.5898 | 1 | LI20 | 51 | 4 | HARDWARE | HW Problem | 20-5 Mod T20 input snubber wire/lug is burnt and needs to be R/R
132645 | 42583.2999 | 1 | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X (21-2) phase jitter is excessive, and is clearly affecting the beam quality.
132645 | 42583.2999 | 1 | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X (21-2) phase jitter is excessive, and is clearly affecting the beam quality.
132998 | 42622.6348 | 2 | LI27 | 31 | 4 | HARDWARE | HW Problem | Klystron 27-3 MOD faulted and resets remotely, but trips off again a couple minutes later.
133038 | 42629.4228 |  | LI24 | 11 | 4 | HARDWARE | HW Problem | 24-1 and 24-2 - BCS window for standby time is 10uS, should be 20uS.
133049 | 42630.5345 | 1 | LI24 | 41 | 4 | HARDWARE | HW Problem | Kly. LI 24 - 4 tripped 3 times on "Electro-magnetic" fault
133094 | 42636.3527 |  | LI26 | 21 | 4 | HARDWARE | HW Problem | 26-2 showing beam volt jitter and amplitude jitter on FTPs, affects beam quality when it's on the beam.
133094 | 42636.3527 |  | LI26 | 21 | 4 | HARDWARE | HW Problem | 26-2 showing beam volt jitter and amplitude jitter on FTPs, affects beam quality when it's on the beam.
133205 | 42645.0387 | 0.5 | LI23 | 21 | 4 | HARDWARE | HW Problem | Kly. LI 23 - 2 tripped on 1 hour Klystron heater delay
133778 | 42697.3784 |  | LI22 | 71 | 4 | HARDWARE | HW Problem | Duplicate of 132111 - KLYS:LI22:71:WNDW1_TEMP Thermal couple looks to be broken. Read back is -10982.3
133781 | 42697.5689 | 0.5 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 is tripping on klystron water faults.
133789 | 42698.442 | 0.75 | LI25 | 31 | 4 | HARDWARE | HW Problem | Kly. LI 25 - 3 trips off when I tried to turn it on.
133794 | 42700.2771 | 1 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3 has tripped off at least three times on owl shift; EDM Error panel shows 6 'CURR' faults.
133794 | 42700.2771 | 3 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3 has tripped off at least three times on owl shift; EDM Error panel shows 6 'CURR' faults.
132175 | 42542.5981 | 1 | LI26 | 31 | 4 | HARDWARE | HW Problem | 26-3 WGV flt
132203 | 42544.0351 | 1 | LI26 | 11 | 4 | HARDWARE | HW Problem | station 26-1 FLK -won't reset remotely
132301 | 42555.3151 | 0.2 | LI24 | 11 | 4 | HARDWARE | HW Problem | 24-1: DQing current is overshot from 22A to 27.5A on 7/1 to present.
132444 | 42564.5744 |  | LI24 | 11 | 4 | HARDWARE | HW Problem | station 24-1 has Water Waveguide 2 fault
133220 | 42647.0011 | 1 | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3 HVOC and ACOC. tripped off countless times during swing shift and continues to trip off on HVR. Resets each time but then trips again later.
133266 | 42650.4121 | 0.1 | LI26 | 31 | 4 | HARDWARE | HW Problem | LI 26 Kly. - 3 tripped on WGV. informed PEM
133348 | 42656.4933 | 1.5 | LI23 | 11 | 4 | HARDWARE | HW Problem | 23-1 HVOC's possibly caused by PFN capacitor going bad
132042 | 42532.7921 | 0.5 | LI29 | 21 | 4 | HARDWARE | HW Problem | Klys 29-2, has bouts of constant tripping.
132821 | 42602.847 |  | LI21 | 61 | 4 | HARDWARE | HW Problem | Klys 21-6 appears to be tripping off and having a CAM error simultaneously. Brought back with an IPL once, but unable to repeat the trick.
133039 | 42629.4278 |  | LI15 | 41 | 4 | HARDWARE | HW Problem | 614A at station 15-4C shows a high tube pressure at 8 x 10-8
133076 | 42633.6066 | 2 | LI21 | 11 | 4 | HARDWARE | HW Problem | 21-1 occasional faults for SCR pan fan
133206 | 42645.0682 | 2 | LI21 | 11 | 4 | HARDWARE | HW Problem | Klys. LI 21 - 1 tripping off mult. times. PIOP message show Focus P.S. faults. PEM suspects MSKU unit maybe at fault.
133943 | 42712.4155 | 1.2 | LI27 | 81 | 4 | HARDWARE | HW Problem | station 27-8 has High Beam Volts Jitter --see mccElog entry today at 10:00 for pictures
132234 | 42547.4376 | 1 | LI29 | 11 | 4 | HARDWARE | HW Problem | Klys 29-1 showing BVLT jitter.
132705 | 42591.4857 | 1.5 | LI30 | 41 | 4 | HARDWARE | HW Problem | 30-4: Seeing some beam volts jitter.
132756 | 42595.3787 | 1 | LI24 | 81 | 4 | HARDWARE | HW Problem | 24-8 TBR Arcing in the klystron tank..
132756 | 42595.3787 | 2 | LI24 | 81 | 4 | HARDWARE | HW Problem | 24-8 TBR Arcing in the klystron tank..
132756 | 42595.3787 | 1.5 | LI24 | 81 | 4 | HARDWARE | HW Problem | 24-8 TBR Arcing in the klystron tank..
132871 | 42608.3851 | 1 | LI29 | 11 | 4 | HARDWARE | HW Problem | 29-1 has significant phase jitter.
133070 | 42632.8102 | 1 | LI26 | 71 | 4 | HARDWARE | HW Problem | 26-7 tripped on KLV with elevated klystron ion pump current readout.
133366 | 42657.6373 | 6 | LI21 | 21 | 4 | HARDWARE | HW Problem | Time Accounting: Had to replace cable for L1X jitter.
133394 | 42660.4131 | 4 | LI24 | 31 | 4 | HARDWARE | HW Problem | Klys 24-3 cycling freqently, error count worthless.
133407 | 42661.4189 | 1.5 | LI21 | 41 | 4 | HARDWARE | HW Problem | 21-4 is having lots of BVLT drop outs.
133422 | 42662.2584 | 0.5 | LI25 | 11 | 4 | HARDWARE | HW Problem | 25-1 is MOD faulted, trips off immediately when it is turned on.
133429 | 42662.4133 | 1.5 | LI22 | 71 | 4 | HARDWARE | HW Problem | 22-7 Beam Current and Beam voltage fast time plots are in-operative, suspect bad PIOP. This problem began after PIOP swap in August
132271 | 42551.2577 | 2 | LI30 | 61 | 4 | HARDWARE | HW Problem | Klystron 30-6 MOD faulted 3 times over the last 5 minutes. Delta temperature.
132271 | 42551.2577 | 0.5 | LI30 | 61 | 4 | HARDWARE | HW Problem | Klystron 30-6 MOD faulted 3 times over the last 5 minutes. Delta temperature.
134154 | 42768.4818 |  | LI26 | 51 | 4 | HARDWARE | HW Problem | 26-5: Dead man switch failure at door 2.
134164 | 42771.3164 | 0.5 | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2: Remove tail clipper assembly.
134165 | 42771.3187 | 0.8 | LI21 | 31 | 4 | HARDWARE | HW Problem | 21-3: Remove tail clipper assembly.
132131 | 42540.5503 | 1 | LI28 | 71 | 4 | HARDWARE | HW Problem | RFS 28-7 tripping upon turn on with ACOC faults.
132579 | 42576.5758 | 1 | LI23 | 61 | 4 | HARDWARE | HW Problem | 23-6 has a damaged/over heated DQing load resistor jumper on the north side, far right bottom jumper
132580 | 42576.5767 | 1 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3 has a damaged/over heated DQing load resistor jumper on the south side, far right, lower connection
133200 | 42644.2882 | 2 | LI29 | 11 | 4 | HARDWARE | HW Problem | Kly LI 29 - 1 tripped off. Now on 1 hour heater delay.
133230 | 42648.2684 |  | LI23 | 31 | 4 | HARDWARE | HW Problem | Klys 23-3 lost AMPL/PHAS readback.
133249 | 42649.4666 | 1 | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3 focus monitor current varying +/- 1 amp, noisy on history plot
133292 | 42652.7371 |  | LI23 | 51 | 4 | HARDWARE | HW Problem | Quad LI24 713 (QM21) won't trim after power dip.
133309 | 42654.2273 |  | LI21 | 41 | 4 | HARDWARE | HW Problem | 21-4 amplitude dropped out momentarily about half a dozen times on 10/11 owl shift. Does not register on the sector error counts and it appears that the beam volts don't drop during these temporary excursions.
133318 | 42654.4782 | 2 | LI26 | 61 | 4 | HARDWARE | HW Problem | Kly. LI 26 - 6 tripped the VVS breaker.
133439 | 42663.2294 | 1 | LI26 | 61 | 4 | HARDWARE | HW Problem | 26-6 MOD/HVR faults, no remote reset.
133494 | 42670.5578 |  | LI25 |  | 4 | HARDWARE | HW Problem | Sector 25, Need to run POLYs on these Stations: 25-5,6,8 before running a Poly on the SBST. Put into CKP when complete.
133499 | 42671.0076 | 1.2 | LI29 | 51 | 4 | HARDWARE | HW Problem | 29-5 has bad phase and amplitude jitter.
132735 | 42593.9481 |  | LI25 | 31 | 4 | HARDWARE | HW Problem | Klys 25-3 tripping constantly on RE faults. Keeping us from running high energy.
132735 | 42593.9481 |  | LI25 | 31 | 4 | HARDWARE | HW Problem | Klys 25-3 tripping constantly on RE faults. Keeping us from running high energy.
133043 | 42629.568 |  | LI24 | 41 | 4 | HARDWARE | HW Problem | Klystron LI24-41 is tripped off on electro-magnet current fault and will not reset remotely.
133120 | 42639.2674 |  | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 seeing WKL FS trips when LI29 KCS LCW is taxed by additional heat load (as in when SBST 28 trips off).
133387 | 42660.2742 | 1 | LI27 | 41 | 4 | HARDWARE | HW Problem | 27-4 has had excessive HVOC trips
134287 | 42796.5002 | 1 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 has water leaking out the magnet weep hole
132761 | 42596.2217 | 1 | LI24 | 11 | 4 | HARDWARE | HW Problem | Klystron 24-1 HVR faulted and won't reset remotely.
133337 | 42655.6675 | 1 | LI22 | 61 | 4 | HARDWARE | HW Problem | Kly LI 22 - 6 tripped on "MOD" fault and won't reset.
133339 | 42655.7357 | 1 | LI20 | 81 | 4 | HARDWARE | HW Problem | 20-8 has Keep Alive undercurrent fault. Can't bring it on remotely.
133526 | 42673.3971 | 3 | LI23 | 31 | 4 | HARDWARE | HW Problem | Klys 23-3 trips frequently.
131828 | 42517.5912 |  | LI29 | 81 | 4 | HARDWARE | HW Problem | 29-8 Water summary flt on. Water fault indicator shows WG1 flow switch faulty.
131912 | 42522.4656 | 0.75 | LI23 | 81 | 4 | HARDWARE | HW Problem | 23-8: Seeing some beam volts jitter.
132125 | 42539.3101 | 1 | LI25 | 61 | 4 | HARDWARE | HW Problem | station 25-6 mod faulted
132126 | 42539.4073 | 2 | LI25 | 11 | 4 | HARDWARE | HW Problem | RFS 25-1 trips off immediately upon turn on.
132832 | 42605.1528 | 2 | LI24 | 21 | 4 | HARDWARE | HW Problem | 24-2 tripping approx. every ten minutes.
132832 | 42605.1528 | 1 | LI24 | 21 | 4 | HARDWARE | HW Problem | 24-2 tripping approx. every ten minutes.
133420 | 42662.1572 | 0.75 | LI22 | 51 | 4 | HARDWARE | HW Problem | 22-5 has relatively high beam volts jitter, which is in turn causing phase jitter.
131796 | 42516.5412 |  | LI27 | 11 | 4 | HARDWARE | HW Problem | 27-1 Water leaks from the klystron hose.
133896 | 42708.7086 | 1 | LI23 | 41 | 4 | HARDWARE | HW Problem | 23-4 showing lots of phase jitter causing FEL spikes down
133897 | 42708.7103 | 1 | LI23 | 51 | 4 | HARDWARE | HW Problem | 23-5 showing lots of phase jitter causing FEL spikes down
133899 | 42709.0752 | 0 | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3 has occasional BV dropouts, bad amplitude jitter.
133900 | 42709.0771 | 1.2 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3 has high amplitude jitter.
131962 | 42525.0698 | 5 | LI21 | 11 | 4 | HARDWARE | HW Problem | L1S is exhibiting single shot occasional amplitude dropouts. See https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=847700
131990 | 42528.6119 | 2 | LI20 | 81 | 4 | HARDWARE | HW Problem | Kly Mod 20-8 (L0B): station AC controllers unstable. Core bias and thyratron filament become unstable above 20KV from HVPS. Causes repeated trips on core bias UC and, mostly, thyratron filament OCs.
131993 | 42528.9786 |  | LI20 | 81 | 4 | HARDWARE | HW Problem | Downtime accounting: Klys 20-8 (L0B), tripping frequently
132034 | 42531.4896 | 1 | LI23 | 11 | 4 | HARDWARE | HW Problem | 23-1 has developed intermittent bouts of extremely severe phase and amplitude jitter.
132034 | 42531.4896 | 1 | LI23 | 11 | 4 | HARDWARE | HW Problem | 23-1 has developed intermittent bouts of extremely severe phase and amplitude jitter.
132213 | 42544.516 | 1 | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7: Tripping quickly after turn-on. Seeing HVOV and EOLC on the fault log.
132213 | 42544.516 |  | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7: Tripping quickly after turn-on. Seeing HVOV and EOLC on the fault log.
132213 | 42544.516 |  | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7: Tripping quickly after turn-on. Seeing HVOV and EOLC on the fault log.
133105 | 42637.1642 |  | LI22 | 71 | 4 | HARDWARE | HW Problem | 22-7 no phase trim. See MCCELOG entry: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=864148
133601 | 42681.3052 | 1 | LI21 | 81 | 4 | HARDWARE | HW Problem | station 21-8 high Amplitude Jitter
133601 | 42681.3052 |  | LI21 | 81 | 4 | HARDWARE | HW Problem | station 21-8 high Amplitude Jitter
133601 | 42681.3052 | 1 | LI21 | 81 | 4 | HARDWARE | HW Problem | station 21-8 high Amplitude Jitter
133649 | 42683.9823 |  | LI23 | 71 | 4 | HARDWARE | HW Problem | KLYS:LI23:71:Phase Jitter is 1.5 %. Was 1.0% before the PAMM.
133656 | 42684.4616 |  | LI26 | 31 | 4 | HARDWARE | HW Problem | 26-3 WGV fault. No observable vacuum issues, probably just needs gauge juiced.
133670 | 42685.5928 |  | LI24 | 31 | 4 | HARDWARE | HW Problem | 24--3 shows no amplitude. It is not on the beam, but it is behaving exactly as 24-1 did earlier (see CATER 133667). Initially there was zero amplitude shown, tried to IPL PIOP, after the station has zero amplitude but this does not propagate an error up to the klystron CUD, and there is no phase readback.
132451 | 42564.6322 | 2 | LI28 | 51 | 4 | HARDWARE | HW Problem | 28-5: BV OOT at 331kV. Adjust to 335kV. Check DQ'g current and adjust as needed.
132451 | 42564.6322 | 2 | LI28 | 51 | 4 | HARDWARE | HW Problem | 28-5: BV OOT at 331kV. Adjust to 335kV. Check DQ'g current and adjust as needed.
132451 | 42564.6322 | 1 | LI28 | 51 | 4 | HARDWARE | HW Problem | 28-5: BV OOT at 331kV. Adjust to 335kV. Check DQ'g current and adjust as needed.
133053 | 42631.3896 | 0.5 | LI29 | 51 | 4 | HARDWARE | HW Problem | 29-5 has high phase jitter.
133053 | 42631.3896 | 1 | LI29 | 51 | 4 | HARDWARE | HW Problem | 29-5 has high phase jitter.
133260 | 42650.2604 | 1 | LI26 | 11 | 4 | HARDWARE | HW Problem | Klystron 26-1 HVR and won't reset remotely. It comes on briefly, but trips off again afterwards.
133388 | 42660.3241 | 1 | LI24 | 51 | 4 | HARDWARE | HW Problem | 24-5 The PFN need tuning, slope =-0.6
133996 | 42717.4495 | 1 | LI23 | 41 | 4 | HARDWARE | HW Problem | 23-4 has large Beam Volts jitter
131870 | 42520.242 |  | LI20 | 71 | 4 | HARDWARE | HW Problem | L0A PAC waveforms are not present in plots. I and Q vals are to be expected, L0A ampl is correct and phase scan shows the phase is also correct. This happened after a power dip MON morning at around 5am
131879 | 42520.5641 | 1 | LI26 | 31 | 4 | HARDWARE | HW Problem | 26-3 has no DQing current
132178 | 42542.8487 |  | LI26 | 11 | 4 | HARDWARE | HW Problem | 26-1 intermittently reporting false WKL faults. Station stays on when this happens.
132436 | 42564.5565 | 1 | LI26 | 11 | 4 | HARDWARE | HW Problem | LI26/27 Mod DQing resistors need to be inspected for burnt Jumper links and cater for repair
132440 | 42564.5647 | 0.1 | LI29 | 51 | 4 | HARDWARE | HW Problem | station 29-5 has klystron Vacuum fault
133674 | 42686.0675 | 1.4 | LI29 | 41 | 4 | HARDWARE | HW Problem | 29-1 WGV. Locally vacuum gauge controller is in "Protect" mode reporting to control system a non-plausible 2e-4 torr.
133716 | 42690.8122 |  | LI20 | 81 | 4 | HARDWARE | HW Problem | Klystron Li20-81 (L0-B) shows intermittent Mod Fault but klystron is still functioning. Bad status.
133725 | 42691.9485 |  | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 CBL fault; resetting/IPLing PIOP doesn't help.
133728 | 42692.5255 |  | LI27 | 31 | 4 | HARDWARE | HW Problem | Kly. LI 27 - 3 tripped "MOD' fault. Tried to reset 3 times but faulted each time.
131726 | 42509.4836 |  | LI27 | 61 | 4 | HARDWARE | HW Problem | 27-6: Broken/Burnt carborundum resistors found during PFN cap type inspection. APS locked applied.
132004 | 42529.5306 | 1 | LI29 | 21 | 4 | HARDWARE | HW Problem | 29-2: Seeing some beam volts jitter. Also tune PFN.
132004 | 42529.5306 | 2 | LI29 | 21 | 4 | HARDWARE | HW Problem | 29-2: Seeing some beam volts jitter. Also tune PFN.
132186 | 42543.4336 | 1 | LI30 | 71 | 4 | HARDWARE | HW Problem | 30-7 Modulator control power will not activate after PAMM inspection. (flickering F12 LED on the control board )
132186 | 42543.4336 | 1 | LI30 | 71 | 4 | HARDWARE | HW Problem | 30-7 Modulator control power will not activate after PAMM inspection. (flickering F12 LED on the control board )
132420 | 42564.3411 | 1 | LI29 | 71 | 4 | HARDWARE | HW Problem | 29-7 Kly trim supply power on lamp is burnt out and should be replace (pref with x10 which is lighter supply)
132427 | 42564.3787 | 2 | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 SCRTD EOLC LED is not indicating a signal, Could be burnt out. Check for continuity from CT to SCRTD.
132427 | 42564.3787 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 SCRTD EOLC LED is not indicating a signal, Could be burnt out. Check for continuity from CT to SCRTD.
132450 | 42564.6305 | 1 | LI23 | 81 | 4 | HARDWARE | HW Problem | 23-8: BV OOT at 342kV. Adjust to 350kV. Check DQ'g current and adjust as needed.
132487 | 42568.9239 | 1 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3 has had multiple over-current faults this shift.
133939 | 42711.6991 |  | LI30 | 71 | 4 | HARDWARE | HW Problem | Kly. LI 30 - 7 BVLT. came back 10 volts lower than when the unit went down on Mon.
131877 | 42520.5554 | 0.5 | LI23 | 41 | 4 | HARDWARE | HW Problem | 23-4 BVJ
131978 | 42528.0098 |  | LI25 | 51 | 4 | HARDWARE | HW Problem | 25-5 MOD fault no reset
131978 | 42528.0098 | 1.5 | LI25 | 51 | 4 | HARDWARE | HW Problem | 25-5 MOD fault no reset
132200 | 42543.7533 | 1 | LI23 | 71 | 4 | HARDWARE | HW Problem | Klystron LI23-71 has a Fire Wire fault and will not reset.
132246 | 42548.6069 | 1 | LI29 | 61 | 4 | HARDWARE | HW Problem | 29-6 Mod had several ACOC faults on turn on Mod Main Contactor and Assoc equip needs to be Inspected for possible cause.
133770 | 42697.032 | 1 | LI23 | 81 | 4 | HARDWARE | HW Problem | Klys 23-8 is cycling/tripping frequently. Error count only shows AMPL errors.
131789 | 42516.4367 | 1 | LI28 | 71 | 4 | HARDWARE | HW Problem | 28-7 Low Voltage PS interlocks and Thyratron filament / reservoir fail to engerize
132002 | 42529.4749 | 1 | LI24 | 11 | 4 | HARDWARE | HW Problem | 24-1 has high beam volts jitter.
132055 | 42534.4198 | 2 | LI25 | 51 | 4 | HARDWARE | HW Problem | 25-5 has massive amplitude jitter and beam volts are 20 kV low.
132166 | 42542.1712 | 1 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3 has high amplitude jitter.
132422 | 42564.3431 | 1 | LI27 | 81 | 4 | HARDWARE | HW Problem | 27-8 Kly trim supply power on lamp is burnt out and should be replace (pref with x10 which is lighter supply)
132431 | 42564.4647 | 0.3 | LI24 | 11 | 4 | HARDWARE | HW Problem | RFS 24-1 no drive power
132433 | 42564.5538 | 0.5 | LI28 | 71 | 4 | HARDWARE | HW Problem | Klystron 28-7 MOD faulted and won't reset remotely. Station is in local mode.
131723 | 42509.4147 | 1 | LI25 | 71 | 4 | HARDWARE | HW Problem | 25-7 Mod has a severely burnt EOLC stack
133922 | 42709.7754 | 0.5 | LI24 | 21 | 4 | HARDWARE | HW Problem | 24-2 having amplitude and phase dropouts, enough to trip off the beam often. Put to standby.
133928 | 42710.5644 | 0.2 | LI22 | 51 | 4 | HARDWARE | HW Problem | Klys 22-5 is WGV
133928 | 42710.5644 | 0.2 | LI22 | 51 | 4 | HARDWARE | HW Problem | Klys 22-5 is WGV
132199 | 42543.7251 | 0.5 | LI20 | 61 | 4 | HARDWARE | HW Problem | Klys 20-6 modulator won't turn on, FLK won't reset.
132418 | 42564.331 | 2 | LI23 | 11 | 4 | HARDWARE | HW Problem | LI23/24 Power ten supplies are running warm due to dust/dirt blocking air vents in cover and need to be brushed/vacuumed off.
132438 | 42564.5588 | 1 | LI30 | 11 | 4 | HARDWARE | HW Problem | LI30 Mod DQing resistors need to be inspected for burnt Jumper links and cater for repair
132503 | 42569.8504 |  | LI25 | 31 | 4 | HARDWARE | HW Problem | Multiple reflected energy faults from 25-3. Ops turned down drive as a temporary solution. Edit 8/24/16, 20:16 -- Station is tripping frequently and reading low amplitude when it is not tripped. Looking at Klystron Fault Count I see 19 reflected energy faults since 08:00 today, the latest 2 minutes ago.
133984 | 42715.8873 |  | LI30 | 71 | 4 | HARDWARE | HW Problem | 30-7 BVLT. came back at 337. Was at 347 before last down time.
131791 | 42516.4459 | 1 | LI22 | 41 | 4 | HARDWARE | HW Problem | 22-4 thyratron reservoir & filament failed to power up when restoring power to modulator
131846 | 42518.9713 |  | LI30 | 61 | 4 | HARDWARE | HW Problem | 30-6 RF amplitude low with drive at (or near) saturation.
131862 | 42519.7414 | 1.5 | LI28 | 51 | 4 | HARDWARE | HW Problem | 28-5 has unreasonable values for forward and reflected power.
132067 | 42535.3154 | 1 | LI28 | 11 | 4 | HARDWARE | HW Problem | Klystron 28-1 MOD faulted and won't reset remotely
132075 | 42535.5448 | 2 | LI28 | 11 | 4 | HARDWARE | HW Problem | Klystron LI28-11 is tripped off with a Mod Fault and will not reset remotely.
132526 | 42571.5676 | 1.5 | LI21 | 61 | 4 | HARDWARE | HW Problem | 21-6 Has 1 overheated D-Qing resistor jumpers that needs replace on the south side. Associated resistor needs to be connected to the standoff and tightened.
132564 | 42573.9107 | 1 | LI30 | 61 | 4 | HARDWARE | HW Problem | 30-6 tripped on mod fault; can't get on locally.
132568 | 42574.8127 | 1 | LI20 | 81 | 4 | HARDWARE | HW Problem | L0B (20-8) tripped off. OPS unable to reset remotely or locally. Called in Stiles from APS. -2h downtime.
132568 | 42574.8127 | 4 | LI20 | 81 | 4 | HARDWARE | HW Problem | L0B (20-8) tripped off. OPS unable to reset remotely or locally. Called in Stiles from APS. -2h downtime.
131800 | 42516.6094 |  | LI29 | 41 | 4 | HARDWARE | HW Problem | 29-4 has a blown klystron filament supply fuse. The fuse holder is stuck and the fuse can not be replaced.
132136 | 42540.6481 | 1 | LI29 | 21 | 4 | HARDWARE | HW Problem | RFS 29-2 trips upon turn on ACOCs
132269 | 42550.7636 | 1 | LI25 | 61 | 4 | HARDWARE | HW Problem | Klys 25-6 trips frequently, about 2-3 times an hour.
132296 | 42552.8796 | 0.8 | LI27 | 41 | 4 | HARDWARE | HW Problem | 27-4 HVR no remote reset.
132687 | 42589.4111 | 1.5 | LI24 | 61 | 4 | HARDWARE | HW Problem | Klys 24-6 trips more frequently after BVLTs raised to 330 on 8/7. Took off beam as is now not useful.
132691 | 42589.625 | 2 | LI22 | 81 | 4 | HARDWARE | HW Problem | Klys 22-8 is showing high AMPL jitter.
132700 | 42591.3678 | 1 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 Attempt to raise beam voltage to 350KV. Raise voltage slowly to prevent vacuum trips.
134234 | 42787.2843 | 2 | LI22 | 81 | 4 | HARDWARE | HW Problem | 22-8: MKSU problem causing PPS red lights to flash in sector 22.
131961 | 42524.8472 | 1 | LI30 | 81 | 4 | HARDWARE | HW Problem | 30-8 fire wire tripped. No external indication of fire.
132040 | 42532.4717 | 1.5 | LI27 | 61 | 4 | HARDWARE | HW Problem | 27-6 faulted to KHD
132041 | 42532.7109 | 1 | LI28 | 71 | 4 | HARDWARE | HW Problem | 28-7 MOD faults repeatedly.
132573 | 42576.367 | 1 | LI28 | 11 | 4 | HARDWARE | HW Problem | 28-1 Klystron Heater delay, not clearing
132578 | 42576.5737 | 1 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-5 has a damaged/over heated DQing load resistor jumper on the south side, upper left jumper arcing to resistor to the right
132584 | 42576.6317 | 1 | LI26 | 41 | 4 | HARDWARE | HW Problem | 26-4 BV FTPs do not look normal. PIOP swap did not fix the problem.
132611 | 42578.1697 | 1 | LI24 | 21 | 4 | HARDWARE | HW Problem | 24-2 has some kind of communication error with the PLC. PLC reads "Host Communication err Error Station No:00h.
134271 | 42792.3888 | 0.5 | LI27 | 41 | 4 | HARDWARE | HW Problem | 27-4: Anode reactor broken on PFN.
134272 | 42792.4093 | 1.2 | LI21 | 41 | 4 | HARDWARE | HW Problem | 21-4: TB7 in thyratron compartment broken.
131730 | 42509.5909 | 1 | LI26 | 61 | 4 | HARDWARE | HW Problem | 26-6 Mod has a burnt/broken carborundum resistor. HV cable between the thyratron and despiking coil needs to be replaced due to broken lug
132190 | 42543.587 | 1.5 | LI20 | 71 | 4 | HARDWARE | HW Problem | 20-7 Mod T20 input snubber wire/lug burnt needs R/R T20 compartment light bulb is burnt out R/R
132204 | 42544.0359 | 0.5 | LI30 | 51 | 4 | HARDWARE | HW Problem | station 30-5 HVR --Can't turn on remotely
132415 | 42564.3273 | 1 | LI29 | 11 | 4 | HARDWARE | HW Problem | LI29/30 Power ten supplies are running warm due to dust/dirt blocking air vents in cover and need to be brushed/vacuumed off.
134138 | 42763.6288 | 2 | LI30 | 11 | 4 | HARDWARE | HW Problem | Kly Mod 30-1: PFN cabinet ground hook broken. PFN capacitor #4 of upper string (cap is upper right hand in upper string or string #1) leaking from bushing cap. Replace capacitor and ground hook.
132701 | 42591.3777 |  | LI22 | 81 | 4 | HARDWARE | HW Problem | DUPLICATE of CATER 132693 - 22-8 FTPs coming from MKSU (BVOLT, BCURR, FE, RE) all look distorted. MKSU has been swapped. Request PIOP be replaced.
132113 | 42538.2073 |  | LI30 | 61 | 4 | HARDWARE | HW Problem | Klys 30-6 showing Beam Volts jitter.
132837 | 42605.5352 | 1.2 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 tripping frequently 10 trips in last half hour.
132846 | 42606.0483 | 1 | LI29 | 61 | 4 | HARDWARE | HW Problem | 29-6 has high BV jitter.
132875 | 42609.5634 | 1 | LI24 | 41 | 4 | HARDWARE | HW Problem | 24-4 Tripping on ACOC and HVOC
132941 | 42619.4666 | 3 | LI20 | 71 | 4 | HARDWARE | HW Problem | Kly Mod 20-7(L0A): replace L4888 thyratron and thyratron stand with E2V CX1937AX deuterium thyratron and modified stand with trigger box and new blower.
132941 | 42619.4666 | 3 | LI20 | 71 | 4 | HARDWARE | HW Problem | Kly Mod 20-7(L0A): replace L4888 thyratron and thyratron stand with E2V CX1937AX deuterium thyratron and modified stand with trigger box and new blower.
132941 | 42619.4666 | 7 | LI20 | 71 | 4 | HARDWARE | HW Problem | Kly Mod 20-7(L0A): replace L4888 thyratron and thyratron stand with E2V CX1937AX deuterium thyratron and modified stand with trigger box and new blower.
132941 | 42619.4666 | 7.5 | LI20 | 71 | 4 | HARDWARE | HW Problem | Kly Mod 20-7(L0A): replace L4888 thyratron and thyratron stand with E2V CX1937AX deuterium thyratron and modified stand with trigger box and new blower.
132951 | 42619.8588 | 1.5 | LI22 | 61 | 4 | HARDWARE | HW Problem | 22-6 Control system reporting elevated klystron vacuum.
132958 | 42620.6088 | 1.2 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3: Slope=-0.6 PFN needs tuning.
131733 | 42509.5976 | 1 | LI24 | 51 | 4 | HARDWARE | HW Problem | 24-5 Thyratron compartment De Qing cap has badly burnt lugs
132254 | 42550.2418 | 0.5 | LI23 | 21 | 4 | HARDWARE | HW Problem | 23-2 WGV fault
132277 | 42551.4903 | 1 | LI29 | 11 | 4 | HARDWARE | HW Problem | 29-1: Seeing some beam volts jitter.
132297 | 42552.9959 | 1.5 | LI21 | 31 | 4 | HARDWARE | HW Problem | 21-3 FLK. Reset locally by ACR operators. After ACR ops reset the fault the station seemed to cycle itself a few times before staying on.
132567 | 42574.2918 | 1 | LI26 | 21 | 4 | HARDWARE | HW Problem | 26-2 has EOLC faults
132670 | 42586.0058 | 1 | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3 BV jitter FTPs occasionally have a (presumably spurious) dip at the beginning of the plot. This is probably what's causing the station to show BVJ (beam volts jitter) on the Klystron CUD.
132670 | 42586.0058 | 1 | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3 BV jitter FTPs occasionally have a (presumably spurious) dip at the beginning of the plot. This is probably what's causing the station to show BVJ (beam volts jitter) on the Klystron CUD.
132670 | 42586.0058 | 1 | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3 BV jitter FTPs occasionally have a (presumably spurious) dip at the beginning of the plot. This is probably what's causing the station to show BVJ (beam volts jitter) on the Klystron CUD.
132670 | 42586.0058 | 1 | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3 BV jitter FTPs occasionally have a (presumably spurious) dip at the beginning of the plot. This is probably what's causing the station to show BVJ (beam volts jitter) on the Klystron CUD.
132885 | 42611.3851 | 2 | LI21 | 61 | 4 | HARDWARE | HW Problem | Klystron 21-6 beam volts too low
132898 | 42612.073 | 1 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4 FTP reveals a horrible beam volts oscillation.
132828 | 42604.5507 | 1 | LI23 | 21 | 4 | HARDWARE | HW Problem | Klys 23-2 showing couple of degrees of phase jitter. Don't see this on AMPL/BVLT, kind of odd.
132853 | 42606.5998 | 1 | LI27 | 41 | 4 | HARDWARE | HW Problem | 27-4 trips on turn on ACOC
132878 | 42610.3707 | 1.5 | LI30 | 61 | 4 | HARDWARE | HW Problem | Klys 30-6 showing bad cable status.
132884 | 42611.3539 |  | LI30 | 61 | 4 | HARDWARE | HW Problem | 30-6 Klystron vacuum readback is NFG.
131832 | 42517.8066 |  | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7 has had many vacuum faults since 5/27 turn on.
132807 | 42600.4525 | 0.5 | LI22 | 21 | 4 | HARDWARE | HW Problem | 22-2 WGV appears to need filmanent juiced.
133273 | 42650.5887 | 1 | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2 counting up WA#2 trips.
133144 | 42640.4943 | 2.5 | LI30 | 41 | 4 | HARDWARE | HW Problem | 30-4 having end of line clipper faults.
133144 | 42640.4943 | 1.5 | LI30 | 41 | 4 | HARDWARE | HW Problem | 30-4 having end of line clipper faults.
131857 | 42519.4662 | 1 | LI20 | 51 | 4 | HARDWARE | HW Problem | 20-5 Modulator is off, please turn on.
131895 | 42521.5842 | 2 | LI22 | 21 | 4 | HARDWARE | HW Problem | 22-2 Modulator HV dc filter cap is swollen and needs equip Manager evaluation for disposition
132088 | 42536.6186 | 2 | LI25 | 31 | 4 | HARDWARE | HW Problem | 25-3 no deQ current when modulator is above 330 kV. Would like to run at 350 kV. Tune PFN.
132088 | 42536.6186 | 1.5 | LI25 | 31 | 4 | HARDWARE | HW Problem | 25-3 no deQ current when modulator is above 330 kV. Would like to run at 350 kV. Tune PFN.
132261 | 42550.3911 | 1 | LI29 | 41 | 4 | HARDWARE | HW Problem | 29-4: Seeing some beam volts jitter.
132262 | 42550.4298 | 1 | LI29 | 81 | 4 | HARDWARE | HW Problem | 29-8: Seeing some beam volts jitter.
132310 | 42556.443 | 1 | LI27 | 11 | 4 | HARDWARE | HW Problem | 27-1: there is water leaking from the ACS supply strainer above the fiat rack. There is a bucket there but I don't see an open Cater.
132443 | 42564.5739 | 2 | LI20 | 21 | 4 | HARDWARE | HW Problem | LI20 Kly trim power supply filters located in LI19 need to be cleaned or replaced
132458 | 42565.242 | 0.5 | LI26 | 31 | 4 | HARDWARE | HW Problem | 26-3 WGV fault
133359 | 42657.0818 | 1.2 | LI23 | 21 | 4 | HARDWARE | HW Problem | 23-2 is experiencing beam volts dropouts.
133360 | 42657.0822 | 1.2 | LI23 | 41 | 4 | HARDWARE | HW Problem | 23-4 is experiencing beam volts dropouts.
131913 | 42522.4781 | 0 | LI28 | 21 | 4 | HARDWARE | HW Problem | Kly. LI 28 - 2 keeps tripping off. Had a EOV fault, a TOC fault and a MSKU fault but most of time it just trips off.
132377 | 42560.2891 | 1.5 | LI20 | 51 | 4 | HARDWARE | HW Problem | 20-5 at KHD. Fault log had klystron filament O/C fault. Reset and had normal current but voltage low. With test load box normal reading. Klystron filament connector at tank hot.
133159 | 42641.1047 | 0.5 | LI25 | 11 | 4 | HARDWARE | HW Problem | 25-1 FLK. I reset it and it came on but very unstable on BV FTP and within about 20 seconds tripped again.
131951 | 42524.5717 | 0.5 | LI28 | 41 | 4 | HARDWARE | HW Problem | Kly. LI 28 - 4 tripped on "Electro - Magnetic" Current and won't reset. Install klixon tell-tale.
132026 | 42531.2684 | 1 | LI23 | 21 | 4 | HARDWARE | HW Problem | 23-2 WGV flt
132038 | 42532.2034 | 4 | LI27 | 11 | 4 | HARDWARE | HW Problem | water leak in KLYS gallery sector 27 ACS fitting dripping onto FIAT rack 27-1B
132038 | 42532.2034 | 4 | LI27 | 11 | 4 | HARDWARE | HW Problem | water leak in KLYS gallery sector 27 ACS fitting dripping onto FIAT rack 27-1B
132094 | 42537.2894 |  | LI29 | 11 | 4 | HARDWARE | HW Problem | station 29-1 reporting CAMAC error. resetting and IPLing the PIOP doesn't help. Suggest replacing PIOP next.
132628 | 42579.6146 | 0.75 | LI28 | 81 | 4 | HARDWARE | HW Problem | Kly. LI 28 - 8 Beam voltage has jitter 0.45 - 0.5 %. Has increased last 2 hours.
132635 | 42580.5214 | 1 | LI30 | 61 | 4 | HARDWARE | HW Problem | 30-6: Seeing some beam volts jitter.
132635 | 42580.5214 | 2.5 | LI30 | 61 | 4 | HARDWARE | HW Problem | 30-6: Seeing some beam volts jitter.
132693 | 42590.1333 | 2 | LI22 | 81 | 4 | HARDWARE | HW Problem | 22-8 FTPs coming from MKSU (BVOLT, BCURR, FE, RE) all look ridiculously jittery while the FTPs from the PAD (AMPL, PHASE) show the station is running fine.
133524 | 42673.3196 | 0.8 | LI25 | 31 | 4 | HARDWARE | HW Problem | 25-3 HVR with no remote reset. Comes back on then immediately trips again.
133536 | 42674.8834 | 0.8 | LI22 | 81 | 4 | HARDWARE | HW Problem | KLYS:LI22:81 Beam voltage jitter is raising again up to 0.35
131916 | 42522.5524 | 1 | LI25 | 21 | 4 | HARDWARE | HW Problem | Kly. LI 25 - 2 tripped on WGV and won't reset.
132252 | 42549.3977 | 1.5 | LI30 | 51 | 4 | HARDWARE | HW Problem | 30-5: Reported "Phase Loss" requiring cycling CB7 and the main contactor to clear. Should inspect those components for burning/damage.
132668 | 42585.6104 | 1 | LI28 | 11 | 4 | HARDWARE | HW Problem | Klystron LI28-11 has a klystron vacuum fault that needs a local reset.
133643 | 42683.6486 | 1.2 | LI29 | 81 | 4 | HARDWARE | HW Problem | Mod. 29-7 Touch panel showing SCR fan fault. On the PLC relay PCB, the rearmost SCR fan indicator LED is blinking. Troubleshoot and repair as necessary.
133684 | 42688.8227 | 0.5 | LI29 | 11 | 4 | HARDWARE | HW Problem | Klystron LI29-11 trips off on "Mod Fault" immediately after resetting.
131902 | 42522.1974 | 1 | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2 tripping frequently; took off beam.
131902 | 42522.1974 | 1.5 | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2 tripping frequently; took off beam.
132095 | 42537.3375 | 2 | LI24 | 81 | 4 | HARDWARE | HW Problem | 24-8 modulator does not regulate when changes made to beam voltage (maybe loss of all deQ); DeQ divider may be damaged due to being connected to top of thyratron for many months as part of stability experiment
132263 | 42550.4318 | 1 | LI24 | 41 | 4 | HARDWARE | HW Problem | 24-4: Seeing some beam volts jitter.
132439 | 42564.561 | 1 | LI21 | 11 | 4 | HARDWARE | HW Problem | LI21/22 Kly Trim supplies at end of LI21 have very dirty air filters' That need to be replaced or cleaned with air hose.
132577 | 42576.5472 | 0.3 | LI26 | 31 | 4 | HARDWARE | HW Problem | Kly LI 26 - 3 tripped on WGV. Won't reset.
131718 | 42508.6203 | 1 | LI30 | 21 | 4 | HARDWARE | HW Problem | 30-2: PFN strap needs replacing.
133589 | 42679.7724 | 0.75 | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3 trips immediately upon turn on.
133591 | 42679.9713 | 1 | LI21 | 11 | 4 | HARDWARE | HW Problem | 21-1 L1S seeing spikes (increase) pulses in MOD waveform. These trip the AMO beam. See lclselog time stamp of this CATER as well as MCCELOG time of this CATER.
133598 | 42681.2694 |  | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3 thyratron appears to have occasional miss-fires
133598 | 42681.2694 | 1 | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3 thyratron appears to have occasional miss-fires
131819 | 42517.3928 | 0.5 | LI27 | 41 | 4 | HARDWARE | HW Problem | 27-4 KMC fault. FIAT cb closed, Trim p/s on. MKSU?
131881 | 42520.5716 | 1.5 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3 HVR
132378 | 42560.7622 | 3 | LI21 | 51 | 4 | HARDWARE | HW Problem | 21-5 has extremely high phase and amplitude jitter. Beam volts look okay.
132528 | 42571.5712 | 1 | LI26 | 81 | 4 | HARDWARE | HW Problem | 26-8 Has 2 overheated D-Qing resistor jumpers that needs replace on the north side.
132610 | 42578.1687 | 1 | LI23 | 11 | 4 | HARDWARE | HW Problem | 23-1 having intense phase jitter. No other L2 spares, creating almost unusable beam conditions.
132813 | 42601.1686 | 1.5 | LI21 | 51 | 4 | HARDWARE | HW Problem | 21-5 tripped and now trips immediately when we try to reset it.
133051 | 42630.7463 | 2 | LI27 | 61 | 4 | HARDWARE | HW Problem | Kly. LI 27 - 6 tripped on "MOD" fault and failed to reset.
133229 | 42648.2672 | 0.5 | LI26 | 31 | 4 | HARDWARE | HW Problem | 26-3 WGV fault
132043 | 42533.3497 | 1.01 | LI21 | 31 | 4 | HARDWARE | HW Problem | Temperature readback from KLYS:LI21:31:WNDW2_TEMP is implausibly high.
132676 | 42586.4968 | 0.1 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4: Retrieve the klystron klixon tell-tale box.
133031 | 42628.6414 | 2 | LI21 | 71 | 4 | HARDWARE | HW Problem | Station 21-7 throwing intermittent positive phase flyers disrupting beam energy.
133065 | 42632.5674 | 1 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4 Downtime TBR KMC fault, no remote reset.
133616 | 42682.3596 | 2 | LI29 | 11 | 4 | HARDWARE | HW Problem | 29-1 is MOD faulting frequently. (Not easy to tell given the LI29 CAMAC IOC problem but every time I visit the modulator panel I find it faulted.)
133619 | 42682.4759 | 1.3 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28 - 3 measure klystron heater AC waveform, voltage and current, with AC power analyzer
133622 | 42682.5965 |  | LI29 | 81 | 4 | HARDWARE | HW Problem | 29 - 8 (or nearby) suspect failing PFN cap due to signature odor
133625 | 42682.9783 |  | LI21 | 11 | 4 | HARDWARE | HW Problem | Downtime accounting: L1S (Klys 21-1) BVLT jitter made beam unusable. Had to swap out thyratron from klys 24-8.
133667 | 42685.4155 |  | LI24 | 11 | 4 | HARDWARE | HW Problem | 24-1 has zero amplitude readback. Station appears to be contributing energy to the beam. Recommend IPL, IOC. OPs tried IPL, this did not recover the amplitude but did result in the amplitude error not propagating to the klystron CUD or other displays. The phase readback also became nan after IPL. Station seems to be still triggering and contributing expected amount of energy. The same symptoms have since arisen for 24-3 as well (CATER 133669)
131776 | 42515.4457 | 1 | LI21 | 61 | 4 | HARDWARE | HW Problem | 21-6 Genesys P/S 300V 17A needs replaced due to the meter display blinking and purple status on PIOP display.
132386 | 42562.4686 | 1.5 | LI24 | 61 | 4 | HARDWARE | HW Problem | station 24-6 Mod faulted
132386 | 42562.4686 |  | LI24 | 61 | 4 | HARDWARE | HW Problem | station 24-6 Mod faulted
133290 | 42652.5952 | 1 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 tripping to HVR again
133291 | 42652.7364 |  | LI23 | 51 | 4 | HARDWARE | HW Problem | Klystron 23-5 KMC after power dip. No remote reset.
133291 | 42652.7364 | 1 | LI23 | 51 | 4 | HARDWARE | HW Problem | Klystron 23-5 KMC after power dip. No remote reset.
133395 | 42660.4218 | 1.5 | LI24 | 11 | 4 | HARDWARE | HW Problem | Mod 24-1 Thyratron needs replacement: PFN and Beam volts (and other things) unstable. Ranged thy. res. from 6.35 to 7.3 Volts. Some improvement but still not good.
133418 | 42661.9911 | 0.75 | LI23 | 21 | 4 | HARDWARE | HW Problem | Klys 23-2 has BVJ and PHJ problems. Think it is getting unwholesome ideas from 23-4
133489 | 42670.336 |  | LI24 | 31 | 4 | HARDWARE | HW Problem | Klystron LI24-31 amplitude has suddenly dropped from 69.2 to 49.5. Zimmer edit 10/27 12:40: This may have been a rate dependent issue. It seems to somewhat correlate with us having the full machine at 10Hz rate for a while (instead of 120Hz).
133527 | 42673.5164 | 1.5 | LI27 | 11 | 4 | HARDWARE | HW Problem | KLYS:LI27:1 is showing 1.2 % phase jitter.
133879 | 42706.2754 | 1.2 | LI30 | 41 | 4 | HARDWARE | HW Problem | station 30-4 resets but won't turn on Edit 12/3/2016 20:49- this appears to be the case again, station immediately trips off when we try to bring it back on.
133879 | 42706.2754 | 1.2 | LI30 | 41 | 4 | HARDWARE | HW Problem | station 30-4 resets but won't turn on Edit 12/3/2016 20:49- this appears to be the case again, station immediately trips off when we try to bring it back on.
133883 | 42706.5724 | 1 | LI20 | 81 | 4 | HARDWARE | HW Problem | L0B, klys (20-8) suddenly dropped into heater delay.
133883 | 42706.5724 | 1 | LI20 | 81 | 4 | HARDWARE | HW Problem | L0B, klys (20-8) suddenly dropped into heater delay.
133884 | 42706.5854 | 2 | LI21 | 41 | 4 | HARDWARE | HW Problem | Klys 21-4 AMPL dropped to 3% suddenly. AMRF recommends changing PIOP.
133887 | 42706.893 | 1.5 | LI25 | 31 | 4 | HARDWARE | HW Problem | Station 25-3 mod faulted; won't reset. Edit 12/3/2016 20:27- this station is off again, and now immediately trips off when we try to reset it.
133887 | 42706.893 | 1 | LI25 | 31 | 4 | HARDWARE | HW Problem | Station 25-3 mod faulted; won't reset. Edit 12/3/2016 20:27- this station is off again, and now immediately trips off when we try to reset it.
133887 | 42706.893 | 1.5 | LI25 | 31 | 4 | HARDWARE | HW Problem | Station 25-3 mod faulted; won't reset. Edit 12/3/2016 20:27- this station is off again, and now immediately trips off when we try to reset it.
133888 | 42706.9675 | 2 | LI26 | 81 | 4 | HARDWARE | HW Problem | 26-8 started tripping off constantly, took off the beam. See many amplitude drops and one reflected power error on EDM error display.
133892 | 42707.6043 | 2 | LI24 | 81 | 4 | HARDWARE | HW Problem | 24-8 Thyratron needs replace due to instability
133892 | 42707.6043 | 1.5 | LI24 | 81 | 4 | HARDWARE | HW Problem | 24-8 Thyratron needs replace due to instability
133892 | 42707.6043 | 2.3 | LI24 | 81 | 4 | HARDWARE | HW Problem | 24-8 Thyratron needs replace due to instability
133893 | 42707.7024 |  | LI23 | 61 | 4 | HARDWARE | HW Problem | 23-6 shows beam volts dropouts on FTP; observed correlation between phase jitter and LCLS xray jitter.
132039 | 42532.3298 | 1.5 | LI27 | 51 | 4 | HARDWARE | HW Problem | Klystron 27-5 MOD VVS fault
133112 | 42638.382 | 12 | LI29 | 11 | 4 | HARDWARE | HW Problem | station 29-1 reports Klys Heater Delay,
133112 | 42638.382 | 1 | LI29 | 11 | 4 | HARDWARE | HW Problem | station 29-1 reports Klys Heater Delay,
133191 | 42643.4261 |  | LI26 | 21 | 4 | HARDWARE | HW Problem | 26-2 MOD with no remote reset. Went into KHD.
133210 | 42645.5732 | 1 | LI21 | 31 | 4 | HARDWARE | HW Problem | Klystron LI21-31 is tripped off on Mod fault and will not reset remotely.
133731 | 42692.8263 |  | LI21 | 31 | 4 | HARDWARE | HW Problem | 21-3 shows no amplitude and phase reads NaN. Doesn't seem to be producing any actual amplitude as tested with beam. Resetting/IPLing PIOP doesn't help. PAD has been replaced two times and the problem is still there. Recommend to replace PIOP. Refer to the follwing E-logs https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=873779 https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=873777
132380 | 42561.2749 |  | LI24 | 11 | 4 | HARDWARE | HW Problem | 24-1 showing phase jitter, some amplitude jitter. Beam volts look ok. See cater 132544.
132632 | 42580.3559 | 0.75 | LI29 | 31 | 4 | HARDWARE | HW Problem | Kly. LI 29 - 3 BVJitter is increasing 0.5%
133084 | 42634.6099 | 1 | LI24 | 11 | 4 | HARDWARE | HW Problem | Klys 24-1 HVR, no remote reset.
133741 | 42694.9191 | 0.75 | LI21 | 71 | 4 | HARDWARE | HW Problem | 21-7 having BVLT drop outs that are impeding lasing.
133951 | 42712.6704 | 2 | LI20 | 61 | 4 | HARDWARE | HW Problem | INJ valves IN. Checked gauges and pressure looked normal so we attempted to pull the valves but they just fell back IN. Experts arrived and also noticed no obvious vac activity that would cause the valves to trip IN so after a few more attempts of pulling valves we decided to turn the GUN HV OFF then pull valves. With RF off, the valves stay open. We then lowered the HV for 20-6 to 15KV (from 17.45KV) and then lowered to AMPL setpoint in the LLRF VME display to 1MV. Next we energized the gun and ramped up the HV and AMPL setpoint slowly while watching vac activity in the INJ. We then increased the ampl until the setpoint reaches its max allowed then increased the HV until we got to 17.45KV. (-2H) temps stabilized, beam was delivered. Seems to be caused by GUN RF but no smoking gun found.
133953 | 42712.7214 | 1.2 | LI23 | 61 | 4 | HARDWARE | HW Problem | 23-6 has phase, amplitude and beam voltage jitter + 0.2 to 2.5 jitter. (see pic.s in MCC log)
131924 | 42523.3049 | 1 | LI22 | 51 | 4 | HARDWARE | HW Problem | 22-5 HVR, showing reflected power faults.
131924 | 42523.3049 | 1 | LI22 | 51 | 4 | HARDWARE | HW Problem | 22-5 HVR, showing reflected power faults.
132545 | 42572.5792 | 0.5 | LI22 | 51 | 4 | HARDWARE | HW Problem | 22-5: WG vacuum gauge locked up during BV adjustment. Vacuum higher than surrounding stations.
133140 | 42640.3878 | 1.2 | LI24 | 11 | 4 | HARDWARE | HW Problem | 24-1 HVR faults frequently.
133213 | 42646.1338 | 1 | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3 tripped on "MOD" fault. Won't reset.
133227 | 42647.783 | 2 | LI25 | 31 | 4 | HARDWARE | HW Problem | Klystron LI25-31 cycles and trips off repeatedly. At this time it shows 10 voltage faults and 3 reflected energy faults. Resets okay when it trips, but frequency of trips makes it nearly unusable for production running.
133776 | 42697.3541 |  | LI21 | 81 | 4 | HARDWARE | HW Problem | Kly. LI 21 - 8 Phase jitter is jumping 3.0 deg.s (see pic.s in MCC log.)
133790 | 42698.5076 | 1.5 | LI22 | 51 | 4 | HARDWARE | HW Problem | Kly. LI 22 - 5 tripped Mod - Fault Lockout. Won't reset.
133867 | 42705.6145 | 0.5 | LI30 | 41 | 4 | HARDWARE | HW Problem | Klystron 30-4 has an Electro-Magnet Current fault
133898 | 42708.9589 | 1.2 | LI30 | 71 | 4 | HARDWARE | HW Problem | 30-7 tripped, can't reset remotely.
133923 | 42710.2981 | 1 | LI27 | 61 | 4 | HARDWARE | HW Problem | 27-6 is WGV faulted.
134004 | 42717.8546 |  | LI21 | 41 | 4 | HARDWARE | HW Problem | Kly. LI 21 - 4 Amplit. drop from 50.1 to 44.1, causing the phase to change to - 39 deg.s from 0.0 .The FEL dropped to 20 from 41. Trimmed phase back up to 0.0 but Amplit. still low at 43.6. FEL back to 4.0
134004 | 42717.8546 | 1.2 | LI21 | 41 | 4 | HARDWARE | HW Problem | Kly. LI 21 - 4 Amplit. drop from 50.1 to 44.1, causing the phase to change to - 39 deg.s from 0.0 .The FEL dropped to 20 from 41. Trimmed phase back up to 0.0 but Amplit. still low at 43.6. FEL back to 4.0
133336 | 42655.6649 | 1 | LI26 | 11 | 4 | HARDWARE | HW Problem | Kly LI 26 - 1 tripped on "MOD" fault and won't reset.
133553 | 42676.4269 |  | LI23 | 11 | 4 | HARDWARE | HW Problem | 23-1 seems to have lost accelerate triggers suddenly. Was on the beam and stopped contributing energy altogether, and showed no accelerate triggers. Swapped with 23-3 for now. 11/17/2016- happened again on swing shift tonight for a few hours 6/3/2017- happened again on swing shift at 20:28
133747 | 42695.3897 |  | LI21 | 61 | 4 | HARDWARE | HW Problem | 21-6 and 21-7 aren't getting any triggers at standby time. This isn't a major issue, because the stations still get triggers when they are on the beam, and we almost always use them. A fix for this can probably wait until a POMM or PAMM.
134037 | 42720.6857 | 1 | LI30 | 71 | 4 | HARDWARE | HW Problem | Kly. LI 30 - 7 BVLT. failed to come back were it was on the previous on state. (see plot. in MCC log)
134039 | 42721.4402 | 1.3 | LI26 | 11 | 4 | HARDWARE | HW Problem | 26-1 Trips on WGV when the DRIVE is at 50%
134041 | 42722.0639 | 2 | LI30 | 21 | 4 | HARDWARE | HW Problem | 30-2 is wave guide vacuum fault. Op juiced filament to no avail.
132281 | 42552.1422 | 0.5 | LI26 | 81 | 4 | HARDWARE | HW Problem | Klystron 26-8 MOD faulted and won't reset remotely.
132300 | 42555.3134 | 1 | LI22 | 11 | 4 | HARDWARE | HW Problem | 22-1: DQing current is dropped from 20A to 13A on 5/27 to present.
133174 | 42642.5437 | 1 | LI28 | 61 | 4 | HARDWARE | HW Problem | 28-6 has high beam volts jitter.
133208 | 42645.5327 | 0.5 | LI26 | 11 | 4 | HARDWARE | HW Problem | 26-1 is MOD fault and no reset
133298 | 42653.4468 | 1 | LI29 | 11 | 4 | HARDWARE | HW Problem | 29-1 showing DQing under current and high jitter
133745 | 42695.3752 | 2 | LI27 | 41 | 4 | HARDWARE | HW Problem | 27-4 is FLK
131931 | 42523.4205 | 1 | LI26 | 61 | 4 | HARDWARE | HW Problem | 26-6: Seeing some beam volts jitter.
133642 | 42683.6295 |  | LI29 | 61 | 4 | HARDWARE | HW Problem | Following VVS 15 runup after the 11/9 PAMM, klystrons 29-6, 29-7, 29-8, 30-4, and 30-7 are all off and refusing to turn on.
133650 | 42684.0024 |  | LI23 | 41 | 4 | HARDWARE | HW Problem | KLYS:LI23:41:Phase Jitter is 0.75 %. Was 0.75% before the PAMM.
134009 | 42718.479 | 1 | LI24 | 31 | 4 | HARDWARE | HW Problem | station 24-3 has high Beam Volts jitter. Has risen since 10:00 today.
134012 | 42718.5549 | 0.5 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4 shows Beam Volt jitter, steps of 1kV. Please have a look, res voltage may need an adjustment.
134016 | 42718.8168 | 1 | LI26 | 61 | 4 | HARDWARE | HW Problem | Kly. LI 26 - 6 tripped it's VVS breaker. Need PEM / electricians to reset.
134024 | 42719.4561 | 2 | LI30 | 61 | 4 | HARDWARE | HW Problem | 30-6 went into heater delay spontaneously.
132905 | 42612.8534 | 1.5 | LI26 | 81 | 4 | HARDWARE | HW Problem | Klys 26-8 is MOD and won't reset remotely.
133459 | 42666.4142 |  | LI23 | 11 | 4 | HARDWARE | HW Problem | 23-1 lost accelerate triggers suddenly while on the beam. Klys status showed No Accelerate Triggers" even though the station was active on beam.
133460 | 42667.2094 | 3 | LI28 | 11 | 4 | HARDWARE | HW Problem | 28-1 has been in KHD for three hours.
132988 | 42622.2844 | 1 | LI24 | 41 | 4 | HARDWARE | HW Problem | 24-4 has high beam volts jitter.
132988 | 42622.2844 | 2 | LI24 | 41 | 4 | HARDWARE | HW Problem | 24-4 has high beam volts jitter.
133068 | 42632.676 |  | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3 PIOP bad CBL status
134198 | 42775.2769 | 1.5 | LI26 | 41 | 4 | HARDWARE | HW Problem | 1 spare klystron in sector 26 has a dead KVGC
131939 | 42524.2364 | 0.5 | LI21 | 51 | 4 | HARDWARE | HW Problem | 21-5 is MOD faulted, won't reset.
132098 | 42537.4078 | 1 | LI24 | 11 | 4 | HARDWARE | HW Problem | Klystron LI24-11 has a WG2 water summary fault and will not reset remotely.
132106 | 42537.5491 | 0.1 | LI23 | 11 | 4 | HARDWARE | HW Problem | Klystron LI23-11 is tripped off on HVR and will not reset remotely.
132194 | 42543.5961 | 1.5 | LI30 | 61 | 4 | HARDWARE | HW Problem | 30-6 Mod Despiking coil resistor bracket is showing signs of arcing and needs to be replaced before it breaks down and arcs over to cabinet
132196 | 42543.5985 | 1 | LI30 | 21 | 4 | HARDWARE | HW Problem | 30-2 HV rect stack small guage wire looks overheated and worn and should be replaced.
132960 | 42620.6168 | 1.5 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6: Slope=-0.6 PFN needs tuning.
131829 | 42517.7667 |  | LI20 | 81 | 4 | HARDWARE | HW Problem | 20-8 thyratron filament unstable above 18KV of HV.
132574 | 42576.3729 | 2 | LI26 | 21 | 4 | HARDWARE | HW Problem | Kly. LI 26 - 2 Beam Volt jitter 0.7 to 1.0 ( 0.3) is normal.
132712 | 42592.4808 | 1 | LI23 | 61 | 4 | HARDWARE | HW Problem | 23-6 dropped amplitude and jumped phase.
134055 | 42723.7084 | 0 | LI25 | 41 | 4 | HARDWARE | HW Problem | LI25/26 focus supply bulk tripped off due to a ground current fault. PEM technician determined problem was with individual 25-4 klystron system. (DT SS 12/16/16 -0.3 hr MFX/CXI/XPP, 1.0 hr MEC)
132232 | 42546.8138 | 2 | LI23 | 11 | 4 | HARDWARE | HW Problem | 23-1 has high phase and amplitude jitter
132990 | 42622.2876 | 1 | LI29 | 11 | 4 | HARDWARE | HW Problem | 29-1 has high beam volts jitter.
133073 | 42633.4216 |  | LI26 | 21 | 4 | HARDWARE | HW Problem | Klys 26-2 MOD fault no remote reset. Duplicate cater 133022.
133125 | 42639.4729 | 4 | LI30 | 41 | 4 | HARDWARE | HW Problem | TBR 30-4 klystron magnet is leaking water at the weep hole on the klystron
133314 | 42654.4258 | 1.5 | LI24 | 11 | 4 | HARDWARE | HW Problem | 24-1 Klystron focus readback (FMON) is noisey, jumps between 13 and 15 Amps. Expected is 14.5 Amps. This started when ferrites were added to the focus readback cable.
131792 | 42516.4486 | 1 | LI21 | 31 | 4 | HARDWARE | HW Problem | 21-3 Trim Power supply will not turn on. All other units on the rack are on.
131855 | 42519.4228 |  | LI24 | 11 | 4 | HARDWARE | HW Problem | Water Waveguide #2 fault on klys 24-1. Not critical at the moment, but we will need the station when we go to high energy.
131909 | 42522.3195 |  | LI24 | 51 | 4 | HARDWARE | HW Problem | 24-5 showing amplitude jitter.
132108 | 42537.5993 | 5 | LI28 |  | 4 | HARDWARE | HW Problem | 28-1, 2, 3, 4: Install bypass cap 0.1uF on HV regulator board to improve the PFN voltage stability and calibrate the HV board.
132135 | 42540.6206 | 3 | LI26 | 61 | 4 | HARDWARE | HW Problem | RFS 26-6 FLK no reset counters show ACOC faults
132421 | 42564.3421 | 1 | LI29 | 51 | 4 | HARDWARE | HW Problem | 29-5 Kly trim supply power on lamp is burnt out and should be replace (pref with x10 which is lighter supply)
132969 | 42621.3539 | 1.2 | LI30 | 61 | 4 | HARDWARE | HW Problem | 30-6 Control system is reporting OOT klystron vacuum, 5.1E-7.
133012 | 42626.5352 | 2 | LI24 | 11 | 4 | HARDWARE | HW Problem | 24-1 has tripped several times during day shift. Klystron fault count shows 5 REFL faults.
133160 | 42641.267 | 1 | LI28 | 51 | 4 | HARDWARE | HW Problem | 28-5: DQing cur is unstable. Sometimes, it dropped to 0A and then shot back up to 18A in few minutes.
133184 | 42643.0208 | 6 | LI26 | 11 | 4 | HARDWARE | HW Problem | 26-1 no klystron filament heater voltage or current
133184 | 42643.0208 |  | LI26 | 11 | 4 | HARDWARE | HW Problem | 26-1 no klystron filament heater voltage or current
133723 | 42691.4583 | 1.3 | LI23 | 41 | 4 | HARDWARE | HW Problem | Kly. LI 23 - 4 showing Beam volt jitter. KLYS:LI23:41:BVLTHSTBPLOT shows 2 - 3 volt drops which might be causing this. KLYS:LI23:41:BVJTHSTBPLOT doesn't show any jitter.
132230 | 42546.7215 | 2 | LI25 | 31 | 4 | HARDWARE | HW Problem | Klystron Li25-31 BVLT and AMPL are cycling. See plot in MCC log 25-Jun-16 17:15.
132428 | 42564.3834 | 1 | LI20 | 51 | 4 | HARDWARE | HW Problem | LI20/21 Modulator DQ ing resistors need to be inspected for badly burnt jumper cables and catered for repair.
132448 | 42564.5923 | 1 | LI25 | 21 | 4 | HARDWARE | HW Problem | 25-2 WGV flt.
132481 | 42568.0945 | 1 | LI23 | 81 | 4 | HARDWARE | HW Problem | 23-8 tripping frequently on HV Over Current faults.
132481 | 42568.0945 | 1.5 | LI23 | 81 | 4 | HARDWARE | HW Problem | 23-8 tripping frequently on HV Over Current faults.
132496 | 42569.3999 | 1.5 | LI22 | 51 | 4 | HARDWARE | HW Problem | Klystron 22-5 MOD faulting frequently
132517 | 42571.3899 |  | LI27 | 51 | 4 | HARDWARE | HW Problem | 27-5 is showing Water Accelerator #1 fault in addition to fault lock out error.
132527 | 42571.5699 | 1 | LI26 | 11 | 4 | HARDWARE | HW Problem | 26-1 Has 1 overheated D-Qing resistor jumper that needs replace on the south side.
131795 | 42516.5392 |  | LI28 | 11 | 4 | HARDWARE | HW Problem | 28-1 Water leaks from the klystron hose.
132752 | 42595.1518 | 1 | LI20 | 71 | 4 | HARDWARE | HW Problem | Klystron 20-7 (L0-A) has intermittent bouts of jitter on the beam volts. It seems to have been occurring since 8/11/16 It appears to affect the pulse intensity, but it's not that obvious. See attached figures.
132959 | 42620.6142 | 1.2 | LI24 | 31 | 4 | HARDWARE | HW Problem | 24:3: Slope=0.8. PFN needs tuning.
133310 | 42654.2313 | 1.5 | LI24 | 11 | 4 | HARDWARE | HW Problem | 24-1 had many RE faults on 10/11 owl shift. Ops lowered the drive request from 100% --> 89% as a temporary solution.
133788 | 42698.2018 | 2.3 | LI20 | 71 | 4 | HARDWARE | HW Problem | Klys 20-7 (L0-A) is showing AMPL dropouts. I think this might be on beam volts as well, but I couldn't get those to plot for some reason.
133798 | 42701.3254 | 1.5 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 tripped; can't reset remotely.
132001 | 42529.4439 | 1 | LI25 | 51 | 4 | HARDWARE | HW Problem | 25-5 massive AMJ. deact'd PEM please have alook. thanks
132001 | 42529.4439 | 1 | LI25 | 51 | 4 | HARDWARE | HW Problem | 25-5 massive AMJ. deact'd PEM please have alook. thanks
132081 | 42536.2425 | 5.5 | LI24 | 81 | 4 | HARDWARE | HW Problem | TCAV3 (24-8) HTR delay 1 HR timeout. can't continue MD program. (-1HR)
132193 | 42543.5936 | 1.5 | LI30 | 81 | 4 | HARDWARE | HW Problem | 30-8 Oil filled divider is loose due to broken nylon screw Metal clip for thy comp door is broken
132363 | 42559.3013 | 1.5 | LI29 | 41 | 4 | HARDWARE | HW Problem | Kly. LI 29 - 4 has both Amplt. and Beam Voltage jitter. Seeing Beam volts drop out.
132766 | 42597.3499 | 1 | LI24 | 11 | 4 | HARDWARE | HW Problem | 24-1 HVR with no remote reset. Modulator remote statuses show connectivity issues. Note from Sonya: 24-8 as well K.Luchini: Unable to ping HPRF plc's for both stations. plc-li24-hprf01 plc-li24-hprf08
133080 | 42634.1961 | 1 | LI22 | 41 | 4 | HARDWARE | HW Problem | 22-4 is tripping frequently on reflected energy faults.
133081 | 42634.2771 | 1 | LI30 | 41 | 4 | HARDWARE | HW Problem | 30-4: DQ current is low = 7A. It needs to be calibrated
133117 | 42639.1947 | 0.1 | LI24 | 41 | 4 | HARDWARE | HW Problem | 24-4 tripped on KMC fault.
133237 | 42648.6407 | 1 | LI24 | 31 | 4 | HARDWARE | HW Problem | station 24-3 trips repeatedly
131836 | 42518.5447 | 1 | LI30 | 81 | 4 | HARDWARE | HW Problem | 30-8 Modulator faults on EOLC when turned on.
131841 | 42518.833 |  | LI25 | 31 | 4 | HARDWARE | HW Problem | Station 25-3 phase not trimming.
131885 | 42521.3461 | 1 | LI26 | 41 | 4 | HARDWARE | HW Problem | Kly. LI 26 - 4 tripped on WGV. Won't reset. Informed PEM
131676 | 42503.4273 | 2 | LI30 | 71 | 4 | HARDWARE | HW Problem | May 2016 HPRF outage preparation and recovery
132775 | 42598.1586 |  | LI23 | 31 | 4 | HARDWARE | HW Problem | Klystron 23-3 MOD faulting incessantly. It resets, but then trips off again a couple minutes later.
133005 | 42625.4859 | 1 | LI27 | 21 | 4 | HARDWARE | HW Problem | 27-2 elevated BV jitter occasional small drop outs of BV jitter FTP.
133050 | 42630.6853 | 1 | LI23 | 31 | 4 | HARDWARE | HW Problem | Kly. LI 24 - 3 tripped off 4 times, three time it tripped after we punched it back on. Swapped out with Kly. LI 24 - 1.
132000 | 42529.4154 | 3 | LI24 | 81 | 4 | HARDWARE | HW Problem | 24-8 (aka TCAV3): small T20 oil leak reported.
132119 | 42538.4618 | 1 | LI30 | 61 | 4 | HARDWARE | HW Problem | 30-6 The focus monitor does not track the focus control. Should read 14.5A, reads 15.0A
133726 | 42692.1475 | 1 | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 has faulted "CBL" status which does not remedy with PIOP IPL.
134072 | 42739.6396 | 1 | LI19 | 11 | 4 | HARDWARE | HW Problem | 19-1 Remove battery ion pump supply and restore to AC when power is restored to sector 19 racks.
131878 | 42520.5576 | 1 | LI28 | 51 | 4 | HARDWARE | HW Problem | 28-5 MOD fault
131949 | 42524.5129 | 1 | LI20 | 81 | 4 | HARDWARE | HW Problem | Kly. LI 20 - 8 ( L0-B) tripped and won't reset.
131949 | 42524.5129 | 1 | LI20 | 81 | 4 | HARDWARE | HW Problem | Kly. LI 20 - 8 ( L0-B) tripped and won't reset.
131959 | 42524.7997 |  | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 HVOC faults too frequent to use station
132189 | 42543.5852 | 2 | LI20 | 81 | 4 | HARDWARE | HW Problem | 20-8 PFN strap is damaged and needs to be replaced (U,RHS, #1 ) T20 input snubber wire/lug needs to be replaced
132191 | 42543.5884 | 1 | LI20 | 61 | 4 | HARDWARE | HW Problem | 20-6 Mod Despiking coil has a broken resistor that needs to R/R
132483 | 42568.5763 | 2 | LI27 | 21 | 4 | HARDWARE | HW Problem | 27-2 having frequent WA2 faults.
132511 | 42571.1067 | 1 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3 showing high amplitude jitter
132513 | 42571.3242 | 1 | LI28 | 61 | 4 | HARDWARE | HW Problem | 28-6 has overheated dequing resistor jumper
132549 | 42573.3687 | 1.5 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4 thyratron near EOL
132606 | 42577.6966 |  | LI20 | 81 | 4 | HARDWARE | HW Problem | Time accounting: L0B tripped due to excessive heat load. Caused 40 minutes of downtime for LCLS MD.
132624 | 42579.3841 | 0.5 | LI23 | 11 | 4 | HARDWARE | HW Problem | Kly. LI 23 - 1 tripped on MOD fault and won't reset.
132658 | 42584.0287 | 2.5 | LI23 | 11 | 4 | HARDWARE | HW Problem | 23-1 had a rash of single shot dropouts that tripped the MPS repeatedly. Could not find any evidence of this behavior on FTPs after de-acting the station. Possibly something that happens on accelerate only?
132658 | 42584.0287 | 1 | LI23 | 11 | 4 | HARDWARE | HW Problem | 23-1 had a rash of single shot dropouts that tripped the MPS repeatedly. Could not find any evidence of this behavior on FTPs after de-acting the station. Possibly something that happens on accelerate only?
133732 | 42692.8584 |  | LI23 | 61 | 4 | HARDWARE | HW Problem | 23-6 not contributing energy to beam, similar to ongoing problem with 23-1. Doesn't seem to be firing on accelerate. We are now down two stations in sector 23.
131863 | 42519.7637 | 0.5 | LI26 | 71 | 4 | HARDWARE | HW Problem | 26-7 modulator is in local.
131894 | 42521.5824 | 2 | LI22 | 81 | 4 | HARDWARE | HW Problem | 22-8 De spiking coil resistor assem is loose due to missing Hardware.
132424 | 42564.345 | 1 | LI24 | 41 | 4 | HARDWARE | HW Problem | 24-4 Kly trim supply power on lamp is burnt out and should be replace (pref with x10 which is lighter supply)
132476 | 42566.5017 | 1 | LI26 | 51 | 4 | HARDWARE | HW Problem | 26-5: Seeing some beam volts jitter.
132529 | 42571.5737 | 1 | LI30 | 81 | 4 | HARDWARE | HW Problem | 30-8 Has 1 overheated D-Qing resistor jumpers that needs replace on the south side.
132640 | 42581.2675 | 1.2 | LI21 | 21 | 4 | HARDWARE | HW Problem | L1-X (21-2) tripped with no reset.
131690 | 42505.5733 | 2 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 is FLK
133025 | 42628.4047 |  | LI21 | 11 | 4 | HARDWARE | HW Problem | 21-1 showing symptoms similar to stations that fault when assigned accelerate triggers.
133347 | 42656.4833 | 1 | LI29 | 61 | 4 | HARDWARE | HW Problem | KLYS:LI29:61:Beam voltage jitter is spiking from 0.1 to 0.5. Some drop outs on the Beam voltage read back.
131888 | 42521.4794 | 1 | LI26 | 61 | 4 | HARDWARE | HW Problem | 26-6: Seeing some beam volts jitter.
132011 | 42530.2279 | 1 | LI23 | 81 | 4 | HARDWARE | HW Problem | Klys 23-8 is FLK; no remote reset. It also cycles incessantly when it is on the beam
132011 | 42530.2279 | 1 | LI23 | 81 | 4 | HARDWARE | HW Problem | Klys 23-8 is FLK; no remote reset. It also cycles incessantly when it is on the beam
132217 | 42545.2118 | 1 | LI27 | 81 | 4 | HARDWARE | HW Problem | Klystron 27-8 HVR faulted and won't reset remotely.
132581 | 42576.5781 | 1 | LI23 | 11 | 4 | HARDWARE | HW Problem | 23-1has a damaged/over heated DQing load resistor jumper on the south side, lower, far left connection
132672 | 42586.2181 | 2.5 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 multiple DeQuing under current faults.
132674 | 42586.3974 | 0.5 | LI29 | 21 | 4 | HARDWARE | HW Problem | Klys 29-2 was tripping frequently, including one MKT fault. Trigger problems?
132681 | 42587.4547 | 1 | LI23 | 41 | 4 | HARDWARE | HW Problem | 23-4: Seeing some beam volts jitter.
132681 | 42587.4547 | 0.5 | LI23 | 41 | 4 | HARDWARE | HW Problem | 23-4: Seeing some beam volts jitter.
132727 | 42593.5863 | 1 | LI22 | 61 | 4 | HARDWARE | HW Problem | 22-6 Tripping on HVOCs.
132734 | 42593.8579 | 1 | LI24 | 61 | 4 | HARDWARE | HW Problem | Klys 24-6 is has quit a bit of BVLT jitter after raising BVLTs.
132759 | 42596.019 | 1 | LI23 | 11 | 4 | HARDWARE | HW Problem | Klystron 23-1 does not have accelerate triggers. Tried to IPL the PIOP, reset the PDU, and resetting the PIOP without success. Refer to 132658
132981 | 42622.1214 | 0.5 | LI26 | 31 | 4 | HARDWARE | HW Problem | 26-3 WGV, no remote reset.
133058 | 42632.2465 | 1.5 | LI25 | 11 | 4 | HARDWARE | HW Problem | 25-1 frequently going into HVR/FLK and has full BVLT dropouts often. Resets remotely but continuously cycles. Has been happening since about 3:20 am.
133072 | 42633.3842 | 2 | LI25 | 21 | 4 | HARDWARE | HW Problem | Kly Mod 25-2: remove bad T20 to shop for oil removal and disposal of shorted T20. PCB content of transformer oil unknown. Will have to treat oil, disposal of waste solids, and disposal of xformer as containing PCBs.
133072 | 42633.3842 | 2 | LI25 | 21 | 4 | HARDWARE | HW Problem | Kly Mod 25-2: remove bad T20 to shop for oil removal and disposal of shorted T20. PCB content of transformer oil unknown. Will have to treat oil, disposal of waste solids, and disposal of xformer as containing PCBs.
133072 | 42633.3842 | 3 | LI25 | 21 | 4 | HARDWARE | HW Problem | Kly Mod 25-2: remove bad T20 to shop for oil removal and disposal of shorted T20. PCB content of transformer oil unknown. Will have to treat oil, disposal of waste solids, and disposal of xformer as containing PCBs.
134080 | 42740.3089 |  | LI24 | 41 | 4 | HARDWARE | HW Problem | Klystron 24-4 shows signs of water puddling. Inspect klystron water system and weep hole.
132253 | 42549.5295 | 0.8 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 klys vacuum fault Update: vac gauge LI24 709 (right by klys 24-6) tripped, causing 0.1 hrs of downtime
132719 | 42592.8545 | 0 | LI23 | 31 | 4 | HARDWARE | HW Problem | Klys 23-3 tripping on Over current and EOLC frequently. Important to have for evening of 8/11.
132719 | 42592.8545 | 1 | LI23 | 31 | 4 | HARDWARE | HW Problem | Klys 23-3 tripping on Over current and EOLC frequently. Important to have for evening of 8/11.
132719 | 42592.8545 | 1 | LI23 | 31 | 4 | HARDWARE | HW Problem | Klys 23-3 tripping on Over current and EOLC frequently. Important to have for evening of 8/11.
132379 | 42560.9788 | 1 | LI22 | 51 | 4 | HARDWARE | HW Problem | 22-5 high phase and amplitude jitter; no beam volts jitter.
132395 | 42563.1926 |  | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 modulator frequently going FLK (once every fifteen minutes or so). No faults shown on fault count display.
132396 | 42563.2748 | 0.5 | LI25 | 61 | 4 | HARDWARE | HW Problem | 25-6 The stored Klystron at this station has a ion power supply blown fuse.
132592 | 42577.3078 |  | LI22 | 51 | 4 | HARDWARE | HW Problem | 22-5 showing phase excursions, once every minute or so. I do see single pulse dropouts (HV) in the history and Low RF Power" on the CUD.
133054 | 42631.431 | 2 | LI25 | 21 | 4 | HARDWARE | HW Problem | 25-2 is MOD faulted.
133054 | 42631.431 | 5 | LI25 | 21 | 4 | HARDWARE | HW Problem | 25-2 is MOD faulted.
133056 | 42631.6486 | 1 | LI28 | 41 | 4 | HARDWARE | HW Problem | Kly. LI 28 - 4 KMC fault. Won't reset.
133299 | 42653.4495 | 1 | LI30 | 41 | 4 | HARDWARE | HW Problem | Modulator 30-4: core bias voltage transducer readback is about half of what it should be.
133341 | 42655.9284 | 2 | LI20 | 81 | 4 | HARDWARE | HW Problem | L0B has a core bias fault.
133771 | 42697.0327 | 1 | LI23 | 61 | 4 | HARDWARE | HW Problem | Klys 23-6 is cycling frequently. Tend to see MOD HVOC fault.
133771 | 42697.0327 | 1 | LI23 | 61 | 4 | HARDWARE | HW Problem | Klys 23-6 is cycling frequently. Tend to see MOD HVOC fault.
134079 | 42740.3077 | 1 | LI22 | 41 | 4 | HARDWARE | HW Problem | Klystron 22-4 shows signs of water puddling. Inspect klystron water system and weep hole.
126784 | 42242.5514 | 0.5 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 reported water leak from green hose in Fiat rack.
104518 | 41305.493 | 0.5 | LI20 | 31 | 4 | HARDWARE | HW Problem | 20-3B: Thyratron removed.
107093 | 41381.5002 | 0.5 | LI09 | 41 | 4 | HARDWARE | HW Problem | 9-4: Bad T20 reported.
106694 | 41368.984 |  | LI19 | 51 | 4 | HARDWARE | HW Problem | 19-5: Thyratron is missing.
112460 | 41582.6024 | 0.5 | LI16 | 11 | 4 | HARDWARE | HW Problem | OFF LINE: 16-1: T20 removed
112126 | 41570.8299 |  | LI20 | 81 | 4 | HARDWARE | HW Problem | Water leaking from pipe between 20-8 modulator and klystron.
123094 | 42060.8 | 1 | LI23 | 11 | 4 | HARDWARE | HW Problem | 23-1 BV jitters
123094 | 42060.8 | 2 | LI23 | 11 | 4 | HARDWARE | HW Problem | 23-1 BV jitters
123105 | 42061.3212 | 0.25 | LI13 | 21 | 4 | HARDWARE | HW Problem | Tried 3 times but no reset.
123113 | 42061.4882 | 2 | LI21 | 31 | 4 | HARDWARE | HW Problem | 21-3 Klystron beam current reads 110Amps low. FEPK and PIOP ratios unchanged suggesting a readback issue.
115070 | 41701.7678 | 1 | LI12 | 31 | 4 | HARDWARE | HW Problem | TBR - 12-3 is showing an EVC fault and the modulator won't reset remotely. Klystron filament appears to be shorted.
115070 | 41701.7678 | 0.75 | LI12 | 31 | 4 | HARDWARE | HW Problem | TBR - 12-3 is showing an EVC fault and the modulator won't reset remotely. Klystron filament appears to be shorted.
115070 | 41701.7678 | 0.8 | LI12 | 31 | 4 | HARDWARE | HW Problem | TBR - 12-3 is showing an EVC fault and the modulator won't reset remotely. Klystron filament appears to be shorted.
115070 | 41701.7678 | 2 | LI12 | 31 | 4 | HARDWARE | HW Problem | TBR - 12-3 is showing an EVC fault and the modulator won't reset remotely. Klystron filament appears to be shorted.
115070 | 41701.7678 | 4 | LI12 | 31 | 4 | HARDWARE | HW Problem | TBR - 12-3 is showing an EVC fault and the modulator won't reset remotely. Klystron filament appears to be shorted.
114033 | 41660.6385 | 2.5 | LI14 | 81 | 4 | HARDWARE | HW Problem | 14-8 - New MKSU installed. Check PIOP ratios.
114814 | 41691.4808 |  | LI19 | 41 | 4 | HARDWARE | HW Problem | 19-4: "AMM" low. BV and timing OK. Did not attempt saturation.
115116 | 41703.4382 | 1 | LI05 | 21 | 4 | HARDWARE | HW Problem | 5-2: "FLK" on what error counter says was RE faults. Reset and on OK for now.
117530 | 41789.7263 | 0.5 | LI16 | 51 | 4 | HARDWARE | HW Problem | station 16-5 scp history shows 11 REFL faults since 08:00 today (9.5 hour time span).
117530 | 41789.7263 | 1 | LI16 | 51 | 4 | HARDWARE | HW Problem | station 16-5 scp history shows 11 REFL faults since 08:00 today (9.5 hour time span).
117319 | 41782.5035 | 0.5 | LI12 | 81 | 4 | HARDWARE | HW Problem | 12-8: "Bad Timing Results" due to PFN needs tuned.
117513 | 41789.3591 | 1 | LI20 | 41 | 4 | HARDWARE | HW Problem | 20-4B (XTCAV) is getting a lot of reflected energy faults. Here's what I saw: - Originally MOD faulted - Reset and turned on - Drive RF ramped up to 100 percent - Reflected energy fault occurred and drive went back to 0 - Repeat Drive Ramp - Reflected energy fault cycle Turned station off for now.
117513 | 41789.3591 | 0.5 | LI20 | 41 | 4 | HARDWARE | HW Problem | 20-4B (XTCAV) is getting a lot of reflected energy faults. Here's what I saw: - Originally MOD faulted - Reset and turned on - Drive RF ramped up to 100 percent - Reflected energy fault occurred and drive went back to 0 - Repeat Drive Ramp - Reflected energy fault cycle Turned station off for now.
117513 | 41789.3591 |  | LI20 | 41 | 4 | HARDWARE | HW Problem | 20-4B (XTCAV) is getting a lot of reflected energy faults. Here's what I saw: - Originally MOD faulted - Reset and turned on - Drive RF ramped up to 100 percent - Reflected energy fault occurred and drive went back to 0 - Repeat Drive Ramp - Reflected energy fault cycle Turned station off for now.
117320 | 41782.5057 | 1 | LI14 | 31 | 4 | HARDWARE | HW Problem | 14-3: "Bad Timing Results" due to PFN needs tuned.
117320 | 41782.5057 | 2 | LI14 | 31 | 4 | HARDWARE | HW Problem | 14-3: "Bad Timing Results" due to PFN needs tuned.
118366 | 41823.4301 | 6 | LI03 | 21 | 4 | HARDWARE | HW Problem | 3-2 cycling often. RE faults.
116820 | 41765.3925 |  | LI24 | 71 | 4 | HARDWARE | HW Problem | 24-7. Move Modulator Chassis from LINAC to Building 015. Will be used for the AIP build.
121685 | 41989.0657 | 1 | LI26 | 41 | 4 | HARDWARE | HW Problem | Klystron 26-4 has tripped off approximately 20 times in the last day or so.
125692 | 42168.1209 | 1.3 | LI24 | 41 | 4 | HARDWARE | HW Problem | 24-4 HVR resets and trips back to HVR immediately
125703 | 42170.3722 | 0.8 | LI20 | 51 | 4 | HARDWARE | HW Problem | Check with MCC EOIC to get approval to work on this This station (TCAV0) is used infrequently, and is normally on in Standby. Need improved stability by Tuesday 6/16 MD.
117316 | 41782.4994 | 0.5 | LI05 | 21 | 4 | HARDWARE | HW Problem | 5-2: "Bad Timing Results" due to PFN needs tuned.
116366 | 41747.5467 | 6 | LI19 | 41 | 4 | HARDWARE | HW Problem | 19-4: Low forward energy.
121624 | 41984.6857 | 1 | LI23 | 21 | 4 | HARDWARE | HW Problem | 23-2 high phase jitter unrelated to beam voltage. Seen on IPA output (klystron drive) and possibly on SSSB output. Jitter started after power outage on Nov. 7.
121624 | 41984.6857 | 1 | LI23 | 21 | 4 | HARDWARE | HW Problem | 23-2 high phase jitter unrelated to beam voltage. Seen on IPA output (klystron drive) and possibly on SSSB output. Jitter started after power outage on Nov. 7.
121624 | 41984.6857 | 1 | LI23 | 21 | 4 | HARDWARE | HW Problem | 23-2 high phase jitter unrelated to beam voltage. Seen on IPA output (klystron drive) and possibly on SSSB output. Jitter started after power outage on Nov. 7.
118774 | 41853.5341 |  | LI00 | 20 | 4 | HARDWARE | HW Problem | KLYS LI00 20 S-Band Pre Buncher PAD has a blown fuse due to water-leak in Area 8. Waterleak was repaired.
118725 | 41850.109 | 1 | LI26 | 51 | 4 | HARDWARE | HW Problem | 26-5 HVR fault. Won't reset.
123123 | 42061.7119 | 0.75 | LI10 | 71 | 4 | HARDWARE | HW Problem | 10-7 MOD fault, no remote reset
123124 | 42061.8041 | 0.75 | LI09 | 21 | 4 | HARDWARE | HW Problem | 9-2 MOD fault, no remote reset. This is the energy feedback station.
123138 | 42062.4723 | 1 | LI28 | 51 | 4 | HARDWARE | HW Problem | PEM says the station will stay on, and the fault has been coming and going periodically for at least a week. It would be nice if it didn't alarm in this state.
123144 | 42062.5015 |  | LI05 | 21 | 4 | HARDWARE | HW Problem | 5-2 MOD fault, won't reset.
120389 | 41935.0557 | 0.8 | LI01 | 51 | 4 | HARDWARE | HW Problem | 1-5 PHJ See the pic.s in the MCC log entry. Owl Shift, Thu, 23-Oct-14 01:33
120390 | 41935.0739 |  | LI05 | 71 | 4 | HARDWARE | HW Problem | Kly. LI05-7 tripped on "MOD" fault and won't reset.
116196 | 41742.6369 |  | LI29 | 11 | 4 | HARDWARE | HW Problem | The DLWG temperature for 29-1 is reading 3 degrees above its high alarm limit. It looks like it has been slightly oscillating (following outside temperature) a few degrees above its high alarm temperature for at least he past two weeks (see attached plot).
119646 | 41912.3119 | 2 | LI12 | 21 | 4 | HARDWARE | HW Problem | 12-2 WGV faults, VGC gauge fail & Over Pressure fail to fault
119648 | 41912.3142 | 1 | LI13 | 51 | 4 | HARDWARE | HW Problem | 13-5 VGC over pressure interlock fails to fault
119681 | 41913.3105 | 1 | LI15 | 81 | 4 | HARDWARE | HW Problem | 15-8 Modulator is at ready green but the HVOC light remains on.
119727 | 41913.6537 |  | LI25 | 11 | 4 | HARDWARE | HW Problem | Kly. LI 25 - 1 cycling "MOD"fault to ok, then back to "MOD"fault. False mod fault status
119831 | 41917.6287 |  | LI24 | 41 | 4 | HARDWARE | HW Problem | station 24-4 has 140 AMPL errors since 08:00 this morning
121193 | 41965.0596 | 1 | LI22 | 51 | 4 | HARDWARE | HW Problem | 22-5 has a waveguide vacuum fault
121389 | 41974.5967 | 1.5 | LI15 | 61 | 4 | HARDWARE | HW Problem | 15-6 T20 Network resistor is badly burnt
121395 | 41974.7542 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X needs new thyratron, current one was killed" running at over 10V reservoir on Nov 20th, see:
121399 | 41975.4084 |  | LI04 | 61 | 4 | HARDWARE | HW Problem | 4-6 MOD fault, won't reset.
121402 | 41975.4475 | 1 | LI12 | 81 | 4 | HARDWARE | HW Problem | 12-8 bad HPA, saturation FTP successful, but starts at 40%.
121433 | 41976.4748 | 0.5 | LI05 | 71 | 4 | HARDWARE | HW Problem | 5-7 MOD fault, no remote reset
121459 | 41977.6048 | 1 | LI21 | 61 | 4 | HARDWARE | HW Problem | 21-6 T20 Network filter cap needs to be turned around to correct phase connections
119608 | 41911.0079 |  | LI27 | 41 | 4 | HARDWARE | HW Problem | 27-4 failed remote turn-on
119608 | 41911.0079 | 0.5 | LI27 | 41 | 4 | HARDWARE | HW Problem | 27-4 failed remote turn-on
120098 | 41926.4224 | 1 | LI10 | 21 | 4 | HARDWARE | HW Problem | 10-2 HVOCs, already verified NO ground hooks & lowered Vres 0.2V-> no go
119824 | 41917.3613 |  | LI12 | 51 | 4 | HARDWARE | HW Problem | 12-5 fails NIRP test
120711 | 41947.9307 |  | LI26 | 41 | 4 | HARDWARE | HW Problem | Klys 26-4 is cycling frequently. It has far more AMPL errors than most stations.
120761 | 41949.3445 | 0 | LI28 | 51 | 4 | HARDWARE | HW Problem | Expected to be 360 max.
120865 | 41952.252 |  | LI15 | 31 | 4 | HARDWARE | HW Problem | Klystron 15-3 (KLYS:LI15:31) has a waveguide vacuum fault
119451 | 41905.7825 | 0.3 | LI21 | 11 | 4 | HARDWARE | HW Problem | 21-1 is responding very strongly to rate transitions, suggesting that the standby timing is incorrect.
125272 | 42146.554 | 0.5 | LI26 | 51 | 4 | HARDWARE | HW Problem | 26-5: Light bulb is burnt in T20 compartment
121357 | 41971.347 | 1 | LI22 | 81 | 4 | HARDWARE | HW Problem | Klystron 22-8 mod fault.
121551 | 41982.0518 |  | LI14 | 31 | 4 | HARDWARE | HW Problem | Won't reset.
121552 | 41982.0557 |  | LI11 | 21 | 4 | HARDWARE | HW Problem | Won't reset.
120036 | 41925.1616 |  | LI21 | 51 | 4 | HARDWARE | HW Problem | 21-5 MOD fault no reset.
120222 | 41929.1313 | 0.5 | LI20 | 71 | 4 | HARDWARE | HW Problem | 20-7 L0A MOD fault no reset
120923 | 41955.277 |  | LI30 | 81 | 4 | HARDWARE | HW Problem | We see intermittent periods when 30-8 beam volts are unstable - plots in mccelog at https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=748792
120925 | 41955.297 |  | LI29 | 61 | 4 | HARDWARE | HW Problem | 29-6 tripped ~5 times on owl shift 11/12.
120994 | 41957.3009 | 0.3 | LI18 | 21 | 4 | HARDWARE | HW Problem | Klystron 18-2 waveguide vacuum fault.
121012 | 41957.5724 | 0.5 | LI27 | 81 | 4 | HARDWARE | HW Problem | 27-8 DQ current at 1A
119879 | 41918.8592 | 1 | LI21 | 31 | 4 | HARDWARE | HW Problem | 21-3 tripping frequently "volt" errors counting up on error display.
120228 | 41929.3655 |  | LI17 | 61 | 4 | HARDWARE | HW Problem | Klystron 17-6 mod fault. Mod fault lockout.
120590 | 41942.7168 | 0.2 | DR13 | 1 | 4 | HARDWARE | HW Problem | Klys DR13-1 (NRTL compressor) has bad AMPL/PHAS/BV jitter.
120763 | 41949.4206 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 is running 380 KV. Last run it was 368.5. Raised BVMX and BCMX to 10 and 20 above current value for now.
121194 | 41965.0647 | 1 | DR13 | 1 | 4 | HARDWARE | HW Problem | NRTL Compressor Klystron MOD Fault
121538 | 41981.6256 | 0.2 | LI16 | 81 | 4 | HARDWARE | HW Problem | 16-8 has a WGV fault
121553 | 41982.1211 | 1 | LI28 | 81 | 4 | HARDWARE | HW Problem | Won't reset
121581 | 41983.1801 | 1 | LI27 | 41 | 4 | HARDWARE | HW Problem | Klystron 27-4 has tripped off approximately 15-20 times the past 8 hours, but keeps coming back.
120612 | 41943.7558 | 0.75 | LI16 | 51 | 4 | HARDWARE | HW Problem | 16-5 MOD fault, won't reset
120705 | 41947.6395 |  | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 beam voltage is low due to radiation. Changes to the station might have reduced radiation sufficiently to allow higher voltage operation.
120727 | 41948.5917 |  | LI18 | 41 | 4 | HARDWARE | HW Problem | 18-4: CBL fault; PAD and MKSU have been checked, maybe PIOP 11/09/14, LEB: Klystron showing CBL fault again.
120727 | 41948.5917 |  | LI18 | 41 | 4 | HARDWARE | HW Problem | 18-4: CBL fault; PAD and MKSU have been checked, maybe PIOP 11/09/14, LEB: Klystron showing CBL fault again.
120785 | 41949.6712 | 0.5 | LI26 | 61 | 4 | HARDWARE | HW Problem | 26-6 WGV fault.
120839 | 41950.9355 | 0.5 | LI05 | 41 | 4 | HARDWARE | HW Problem | Craig Butler has reported that klystron 5-4 has a mod fault and needs to be looked at by PEM.
120844 | 41951.3081 |  | LI08 | 81 | 4 | HARDWARE | HW Problem | Klystron 12-8 (KLYS:LI08:81) has an MKSU Trigger Enable Fault.
121093 | 41962.1959 |  | LI00 | 11 | 4 | HARDWARE | HW Problem | LI00 Klys 11 (1-1A) tripping on reflected energy. Trips about once an hour with 17 reflected energy faults in last 20 hours.
121104 | 41962.5374 | 1 | LI09 | 21 | 4 | HARDWARE | HW Problem | 9-2: Replace Core Bias variac
121394 | 41974.6975 | 0.8 | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 beam voltage jitter. Thyratron needs adjustment. Due to large beam voltage changes that were observed in B921 during recent thyratron adjustments, this work occur during a POMM with careful attention.
121394 | 41974.6975 | 0.5 | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 beam voltage jitter. Thyratron needs adjustment. Due to large beam voltage changes that were observed in B921 during recent thyratron adjustments, this work occur during a POMM with careful attention.
120094 | 41926.4182 | 2 | LI09 | 61 | 4 | HARDWARE | HW Problem | 9-6 No turn on with green light & no EXT faults
120094 | 41926.4182 | 1 | LI09 | 61 | 4 | HARDWARE | HW Problem | 9-6 No turn on with green light & no EXT faults
121083 | 41961.5066 | 2 | DR13 | 1 | 4 | HARDWARE | HW Problem | We see significant jitter on the beam volts and phase for the north compressor (DR13 KLYS 1). This is causing the positron beam to be very unstable.
121094 | 41962.2608 |  | LI22 | 81 | 4 | HARDWARE | HW Problem | 22-8 MOD fault, won't reset.
121150 | 41963.7284 | 0.75 | LI10 | 61 | 4 | HARDWARE | HW Problem | Won't reset.
121152 | 41963.7345 | 0.75 | LI06 | 51 | 4 | HARDWARE | HW Problem | Won't reset.
120691 | 41947.5524 | 1.25 | LI10 | 61 | 4 | HARDWARE | HW Problem | 10-6: "Bad Timing Results" due to PFN needs tuned.
120856 | 41951.6319 | 1 | LI04 | 51 | 4 | HARDWARE | HW Problem | 4-5 MOD fault, won't reset.
121226 | 41966.9722 |  | LI29 | 51 | 4 | HARDWARE | HW Problem | 29-5 MOD fault no reset
121228 | 41967.1627 | 0.5 | LI29 | 51 | 4 | HARDWARE | HW Problem | 29-5 MOD fault
121252 | 41967.6667 | 0.75 | LI06 | 31 | 4 | HARDWARE | HW Problem | Klystron 6-3 mod fault.
120514 | 41940.6851 | 1.5 | LI21 | 51 | 4 | HARDWARE | HW Problem | 21-5 thyratron change due to EOL
121137 | 41963.5966 | 1 | LI27 | 61 | 4 | HARDWARE | HW Problem | 27-6 T20 input snubber resistore are badly burnt
121651 | 41986.314 | 1 | LI28 | 61 | 4 | HARDWARE | HW Problem | Klystron 28-6 has tripped off 5 times over the last 8 hours
121341 | 41969.5586 | 1 | LI29 | 81 | 4 | HARDWARE | HW Problem | 29-8: Beam Volts Jitter
121726 | 41990.5038 | 1 | LI13 | 41 | 4 | HARDWARE | HW Problem | 13-4: Beam Voltage Jitter
121288 | 41968.8222 | 0.75 | LI29 | 81 | 4 | HARDWARE | HW Problem | Still won't reset, Inform PEM.
121294 | 41968.9028 | 0.2 | LI17 | 61 | 4 | HARDWARE | HW Problem | 17-6 Tripped on waveguide vac, won't reset
119826 | 41917.3731 | 0.5 | LI30 | 61 | 4 | HARDWARE | HW Problem | Klystron LI30-61 is tripped off and will not reset remotely.
120468 | 41938.492 |  | LI27 | 81 | 4 | HARDWARE | HW Problem | Kly LI27-8 has jitter on beam volt and current plus phase and and amplitude. + 1.0 sigma.
120779 | 41949.5947 | 0.2 | LI29 | 51 | 4 | HARDWARE | HW Problem | Tried to reset the station twice but it failed to reset.
121316 | 41969.4039 | 0.5 | LI19 | 41 | 4 | HARDWARE | HW Problem | 19-4 MOD fault
121630 | 41985.3954 |  | LI26 | 81 | 4 | HARDWARE | HW Problem | 26-8: Beam Volts Jitter
120019 | 41923.1485 | 1 | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7: MOD fault, no remote reset
120129 | 41926.963 | 1 | LI17 | 41 | 4 | HARDWARE | HW Problem | 17-4 klystron vacuum faults at mod on.
120289 | 41931.3218 |  | LI09 | 31 | 4 | HARDWARE | HW Problem | 9-3 MOD fault. Won't reset.
120629 | 41945.3834 | 0.3 | LI27 | 41 | 4 | HARDWARE | HW Problem | Klystron 27-5 has an HVR fault and won't turn on remotely.
120729 | 41948.5967 | 2 | LI01 | 21 | 4 | HARDWARE | HW Problem | 1-2 Mod has a burnt/broken T20 network resistor K26 (B") contacts need to be replaced."
121393 | 41974.6896 | 0.8 | LI21 | 31 | 4 | HARDWARE | HW Problem | 21-3 is currently MOD fault lockout and just cycles through many faults when we try to reset it. Very difficult to make good FEL without this station.
121393 | 41974.6896 | 1.5 | LI21 | 31 | 4 | HARDWARE | HW Problem | 21-3 is currently MOD fault lockout and just cycles through many faults when we try to reset it. Very difficult to make good FEL without this station.
121415 | 41975.6183 | 0.5 | LI11 | 21 | 4 | HARDWARE | HW Problem | 11-2 MOD Fault, won't reset.
121450 | 41977.3884 | 1 | LI25 | 41 | 4 | HARDWARE | HW Problem | 25-4 Low klystron vacuum. 1.7E-7 per PIOP..
120226 | 41929.2653 | 1 | LI23 | 51 | 4 | HARDWARE | HW Problem | 23-5 MOD fault no reset
120253 | 41929.6156 | 0.25 | LI20 | 51 | 4 | HARDWARE | HW Problem | 20-5 MOD Fault. Won't reset.
120267 | 41930.3122 | 0.5 | LI05 | 71 | 4 | HARDWARE | HW Problem | 5-7 has a MOD fault, it will not reset in the SCP.
123850 | 42091.3078 | 1 | LI13 | 11 | 4 | HARDWARE | HW Problem | 13-1 Mod Flt will not remote on
123852 | 42091.6137 | 1 | LI30 | 71 | 4 | HARDWARE | HW Problem | Kly Mod 30-7: trips repeatedly on reflected energy faults until fault lockout occurs.
123853 | 42091.6488 | 1 | LI23 | 31 | 4 | HARDWARE | HW Problem | tried to reset 3 times but it still faulted.
123854 | 42091.6594 | 0.5 | LI28 | 41 | 4 | HARDWARE | HW Problem | Won't reset.
119540 | 41907.9605 | 72 | LI21 | 11 | 4 | HARDWARE | HW Problem | Klystron LI21-11 (L1S) output is jittery and unstable after thyratron swap on 9/25/14.
119647 | 41912.3133 | 2.5 | LI13 | 31 | 4 | HARDWARE | HW Problem | 13-3 VGC over pressure interlock fails to fault
119820 | 41917.1425 | 3 | LI22 | 81 | 4 | HARDWARE | HW Problem | Klystron LI22-81 is tripped off on Mod Fault and will not reset remotely.
119820 | 41917.1425 | 2 | LI22 | 81 | 4 | HARDWARE | HW Problem | Klystron LI22-81 is tripped off on Mod Fault and will not reset remotely.
120269 | 41930.4286 | 1.5 | LI30 | 41 | 4 | HARDWARE | HW Problem | Klystron 30-4 mod fault.
120636 | 41945.8425 | 0.75 | LI12 | 21 | 4 | HARDWARE | HW Problem | Klystron 12-2 mod fault. Unable to reset by remote.
120775 | 41949.5135 | 1 | LI23 | 11 | 4 | HARDWARE | HW Problem | Unit still had a "EXTERNAL" fault showed a "Electro - Magnet current fault" and a "Electro - Magnet current Tol. fault" The "Electro - Magnet Breaker fault" cycled 3 - 4 times.
121582 | 41983.3657 | 2 | LI03 | 51 | 4 | HARDWARE | HW Problem | 3-5 is MOD faulted
121586 | 41983.4642 | 3 | LI20 | 81 | 4 | HARDWARE | HW Problem | 20-8 klystron voltage jumps 8 ns in time. Timing jumps were measured between: MKSU and thyratron grid trigger MKSU and SCRTD front monitor Timing jumps were not visible between MKSU and scope trigger Also, the thyratron anode delay exhibits a voltage drop before the thyratron breaks down (observed at thyratron grid trigger) which is unusual. it was observed that the entire grid trigger jumps at 8 ns, rather than just the thyratron breakdown.
121589 | 41983.54 | 4 | LI26 | 81 | 4 | HARDWARE | HW Problem | 26-8: Seeing some Beam Volts Jitter.
121589 | 41983.54 | 0.8 | LI26 | 81 | 4 | HARDWARE | HW Problem | 26-8: Seeing some Beam Volts Jitter.
121646 | 41985.7029 | 1.5 | LI02 | 71 | 4 | HARDWARE | HW Problem | WGV fault on 2-7 (KLYS:LI02:71)
121648 | 41986.2492 | 0.5 | LI22 | 41 | 4 | HARDWARE | HW Problem | Klystron 22-4 has a Waveguide Vacuum fault.
119685 | 41913.3294 |  | LI19 | 11 | 4 | HARDWARE | HW Problem | 19-1 has a kly stron vacuum fault
119716 | 41913.5683 | 0.5 | LI28 | 41 | 4 | HARDWARE | HW Problem | Kly. 28 - 4 is tripped on Electro-Magnet Current, Breaker, Current Tol.s Won't reset.
119738 | 41913.8265 |  | LI24 | 31 | 4 | HARDWARE | HW Problem | Duplicate 24-3 HVOCs with klystron vacuum bursts
121023 | 41957.9684 | 1 | LI14 | 61 | 4 | HARDWARE | HW Problem | 14-6 MOD fault
121269 | 41968.4298 | 1 | LI29 | 51 | 4 | HARDWARE | HW Problem | Klystron Li29-51 tripped off and will not reset remotely.
121409 | 41975.5395 | 0.5 | LI24 | 21 | 4 | HARDWARE | HW Problem | 24-2: K13 ERR. Needs to reprogram.
121699 | 41989.4698 | 0.5 | LI27 | 71 | 4 | HARDWARE | HW Problem | 27-7: Beam Volts Jitter
119520 | 41907.309 |  | LI03 | 51 | 4 | HARDWARE | HW Problem | 3-5 appears to have an intermittent fire alarm, twice, maybe three times in the last 24 hours there has been a "Supervisory Alarm" from this modulator. When we arrive, no trouble is found other than the Fire Alarm Latching box shows a latched fault. We need to troubleshoot the fire wire / fire alarm circuit in the modulator.
119642 | 41912.3002 |  | LI01 | 51 | 4 | HARDWARE | HW Problem | 1-5 has a ACC2 water fault
119830 | 41917.5725 | 0.5 | LI20 | 81 | 4 | HARDWARE | HW Problem | station 20-8 mod fault
119906 | 41919.9068 | 0.5 | LI25 | 21 | 4 | HARDWARE | HW Problem | KLYS:LI25:21:WATRERRT shows over 290 fault today. Looks like the Flow Switch is to close to it's trip point or it might be getting flaky.
120126 | 41926.8195 |  | LI09 | 61 | 4 | HARDWARE | HW Problem | 9-6 had bad T20 compartment fan.
120500 | 41940.2886 | 1 | LI21 | 31 | 4 | HARDWARE | HW Problem | 21-3 is unstable
120500 | 41940.2886 | 1 | LI21 | 31 | 4 | HARDWARE | HW Problem | 21-3 is unstable
120678 | 41947.3823 | 0.5 | LI21 | 21 | 4 | HARDWARE | HW Problem | Kly. LI21 - 2 phase and amplt. jumping but beam volts are stable.
119795 | 41915.6514 | 0.5 | LI26 | 51 | 4 | HARDWARE | HW Problem | Klystron 26-5 keeps MOD faulting. We can temporarily reset it remotely, but it goes MOD fault again a minute or two afterwards.
119832 | 41917.6456 | 0.5 | LI23 | 51 | 4 | HARDWARE | HW Problem | 23-5: MOD fault, no remote reset. This station was showing a false MOD status earlier this weekend, but now the AMPL is reading 0.0 and beam volts are 12.6.
119955 | 41921.2086 | 1 | LI22 | 41 | 4 | HARDWARE | HW Problem | 22-4: Klystron phase does not respond properly to phase shifter commands. Did a successful FOX Home (JC).
120060 | 41925.576 |  | LI28 | 41 | 4 | HARDWARE | HW Problem | Klystron 28-4 electro-magnet current. BVLT low.
120067 | 41925.7337 |  | LI01 | 51 | 4 | HARDWARE | HW Problem | Klys LI01 - 5 PAD DAC Status = "False".
120638 | 41945.9691 | 0.5 | LI22 | 81 | 4 | HARDWARE | HW Problem | Klystron 22-8 mod fault.
120733 | 41948.6022 | 0.2 | LI29 | 31 | 4 | HARDWARE | HW Problem | 29-3 WGV no reset.
120754 | 41949.2993 |  | LI20 | 93 | 4 | HARDWARE | HW Problem | PEM had a look and even swapped the MKSU to assure that the problem was with the LLRF. No call back from AMRF.
119817 | 41916.7618 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 L1X BV dropouts probably timing.
119900 | 41919.756 | 1.5 | LI22 | 81 | 4 | HARDWARE | HW Problem | 22-8 needs PFN tune. 340KV with 12 amps dequing after T20 replacement. Taps at 100%
119917 | 41920.3487 |  | LI21 | 51 | 4 | HARDWARE | HW Problem | 21-5 K26 contacts need to be replaced. T20 input snubber assem wire badlt burnt Cable between despiking coil and thyratron needs to be replaced.
120247 | 41929.5147 | 1 | LI24 | 41 | 4 | HARDWARE | HW Problem | 24-4 has faulted several times this shift. Errors display shows two over-voltage faults and a lot of AMPL complaints.
120380 | 41934.5354 | 1 | LI03 | 11 | 4 | HARDWARE | HW Problem | 3-1 HVR fault. Won't reset.. Came up several times with "OV" faults as well.
119796 | 41915.655 | 0.8 | LI30 | 71 | 4 | HARDWARE | HW Problem | Klystron 30-7 keeps MOD faulting. We try to reset it remotely, but the beam volts do not reach the correct value and it MOD faults again.
120155 | 41927.5076 |  | LI20 | 41 | 4 | HARDWARE | HW Problem | FACET XTCAV W420 ion pump will not stay on. The pump can be restarted but will eventually shut down.
121101 | 41962.4702 |  | LI22 | 21 | 4 | HARDWARE | HW Problem | Station 22-2 is going haywire every several minutes. I see several vacuum errors looking at the klystron fault count. I occasionally see an 'MKS' fault flash on the klystron cud when this happens.
119683 | 41913.3226 | 1 | LI22 | 71 | 4 | HARDWARE | HW Problem | Klystron LI22-71 is tripped off on Mod Fault and will not reset remotely.
119689 | 41913.3475 | 1.5 | LI27 | 31 | 4 | HARDWARE | HW Problem | KLYS:LI27:31:AMPL shows a drop of amplitude at 05:46. Dropped from 52.8 down to 47.9.
119689 | 41913.3475 | 1 | LI27 | 31 | 4 | HARDWARE | HW Problem | KLYS:LI27:31:AMPL shows a drop of amplitude at 05:46. Dropped from 52.8 down to 47.9.
119809 | 41916.3456 | 7 | LI30 | 61 | 4 | HARDWARE | HW Problem | Klystron 30-6 mod fault.
119895 | 41919.6023 |  | LI02 | 61 | 4 | HARDWARE | HW Problem | 2-6 KMC fault.
119910 | 41920.24 | 1 | LI23 | 51 | 4 | HARDWARE | HW Problem | 23-5 is MOD faulted.
120436 | 41936.4969 |  | LI28 | 11 | 4 | HARDWARE | HW Problem | 28-1: RF AMPL Low OOT. BV and timing OK. Saturated klystron during the PAMM.
120600 | 41943.1588 | 2 | LI02 | 81 | 4 | HARDWARE | HW Problem | Klys 2-8 phase is bouncing around like mad. See: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=745947
120600 | 41943.1588 | 1.5 | LI02 | 81 | 4 | HARDWARE | HW Problem | Klys 2-8 phase is bouncing around like mad. See: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=745947
119739 | 41913.8968 | 0.5 | LI30 | 71 | 4 | HARDWARE | HW Problem | 30-7 beam voltage is erratic.
119799 | 41915.6887 | 2 | LI21 | 11 | 4 | HARDWARE | HW Problem | 21-1 supervisory alarm and MOD fault. PEM Security and FAC Alarm techs dispatched.
120028 | 41924.1676 |  | LI30 | 71 | 4 | HARDWARE | HW Problem | 30-7: WGV fault
120028 | 41924.1676 |  | LI30 | 71 | 4 | HARDWARE | HW Problem | 30-7: WGV fault
120073 | 41925.9386 | 1 | LI04 | 71 | 4 | HARDWARE | HW Problem | 4-7 is faulted on waveguide vacuum on purpose. Thyratron shorted reported.
119822 | 41917.3398 |  | LI23 | 51 | 4 | HARDWARE | HW Problem | Klystron 23-5 mod fault.
119938 | 41920.6469 |  | LI02 | 51 | 4 | HARDWARE | HW Problem | RFS 2-5 has a waveguide vacuum fault.
119940 | 41920.7503 |  | LI26 | 81 | 4 | HARDWARE | HW Problem | KLYS:LI26:81:WATRERRT / Flow Switch is tripping. 39 times since 1600
119941 | 41920.7657 |  | LI27 | 11 | 4 | HARDWARE | HW Problem | KLYS:LI27:21:WATRERRT / Flow Switch tripping off unit.
120015 | 41923.0922 | 0.5 | LI30 | 71 | 4 | HARDWARE | HW Problem | 30-7: WGV fault.
120038 | 41925.2369 |  | LI20 | 51 | 4 | HARDWARE | HW Problem | 20-5 TCAV0 tripped on focus current fault no reset.
120090 | 41926.4134 |  | LI06 | 21 | 4 | HARDWARE | HW Problem | 6-2 Holds HV but not triggering
120100 | 41926.4281 | 1 | LI18 | 41 | 4 | HARDWARE | HW Problem | 18-4 HV on but not triggering. Already raised Vres from 6-1 to 6-4 w. no luck
120100 | 41926.4281 | 0.8 | LI18 | 41 | 4 | HARDWARE | HW Problem | 18-4 HV on but not triggering. Already raised Vres from 6-1 to 6-4 w. no luck
120105 | 41926.4695 |  | LI04 | 41 | 4 | HARDWARE | HW Problem | 4-4: PAD Mixer Gain not within expected value.
119814 | 41916.6727 | 0.75 | LI21 | 71 | 4 | HARDWARE | HW Problem | 21-7: HVR status, no remote reset
120459 | 41937.4707 |  | LI03 | 21 | 4 | HARDWARE | HW Problem | 3-2 PAD phase near mixer null PAD phase readout not working Please add stub in PAD reference leg.
120721 | 41948.4662 |  | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 DQing current at 0A
120529 | 41941.0319 |  | LI27 | 81 | 4 | HARDWARE | HW Problem | Klys 27-8 keeps mod faulting. Comes back each time, but, grrr!
120530 | 41941.1068 | 0.5 | LI20 | 93 | 4 | HARDWARE | HW Problem | KLY LI 20 - 3c tripped on "WGV" and won't reset.
120557 | 41942.1126 | 1 | LI05 | 71 | 4 | HARDWARE | HW Problem | Kly. LI05 - 7 tripped on "MOD" fault and won't reset.
120604 | 41943.4144 |  | LI03 | 21 | 4 | HARDWARE | HW Problem | Time accounting CATER. FACET beam temporarily unusable due to (unreleased by EOIC) work on the 3-2 PAD.
120878 | 41952.8965 | 2 | LI23 | 71 | 4 | HARDWARE | HW Problem | Klys 23-7 has massive AMPL jitter.
120878 | 41952.8965 | 1 | LI23 | 71 | 4 | HARDWARE | HW Problem | Klys 23-7 has massive AMPL jitter.
120926 | 41955.2991 | 2 | LI29 | 31 | 4 | HARDWARE | HW Problem | 29-3 PHJ OOT
121047 | 41960.0185 |  | LI27 | 61 | 4 | HARDWARE | HW Problem | Klystron 27-6 constant cycling. Looks like water issues. WG1 faults.
121107 | 41962.5436 | 0.5 | LI11 | 21 | 4 | HARDWARE | HW Problem | 11-2: Replace T20 network resistor
121181 | 41964.5198 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X phase and amplitude jitter has been worse since the PAMM 11/20/14.
121425 | 41976.4315 |  | LI22 | 81 | 4 | HARDWARE | HW Problem | Klys 22-8 MOD fault.
121511 | 41980.0161 |  | LI02 | 41 | 4 | HARDWARE | HW Problem | Won't reset.
121520 | 41980.7414 | 0.75 | LI16 | 81 | 4 | HARDWARE | HW Problem | Station 16-8 has a WGV fault.
120928 | 41955.3093 | 0.2 | LI25 | 31 | 4 | HARDWARE | HW Problem | Klystron 25-3 mod fault. mod HV ready
120928 | 41955.3093 | 2 | LI25 | 31 | 4 | HARDWARE | HW Problem | Klystron 25-3 mod fault. mod HV ready
120951 | 41955.597 | 1 | LI04 | 31 | 4 | HARDWARE | HW Problem | 4-3 Thyratron compartment deQing cap has a burnt lug
120988 | 41956.8851 | 0.5 | LI02 | 41 | 4 | HARDWARE | HW Problem | 2-4 WGV
121013 | 41957.5743 | 0.5 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 DQ current at 1A
121039 | 41959.4577 | 0.75 | LI30 | 71 | 4 | HARDWARE | HW Problem | Klystron LI30-71 has many drop outs over the last 2 hrs.
121049 | 41960.172 |  | LI29 | 51 | 4 | HARDWARE | HW Problem | Klys 29-5 MOD fault, no reset. See 121269
121050 | 41960.251 | 0.5 | LI22 | 61 | 4 | HARDWARE | HW Problem | Klys 22-6 showing MKSU trigger enable fault.
121105 | 41962.5391 | 1 | LI09 | 81 | 4 | HARDWARE | HW Problem | 9-8: Replace EOLC stack
120843 | 41951.1796 | 0.5 | LI04 | 41 | 4 | HARDWARE | HW Problem | Klystron 4-4 (KLYS:LI04:41) has a waveguide vacuum fault.
120886 | 41953.3861 |  | LI04 | 81 | 4 | HARDWARE | HW Problem | LI04 DL_WG_08 temperature out of tolerance low since 4-Nov. Not regulating well since 7-Nov.
120896 | 41953.5031 | 3 | LI09 | 51 | 4 | HARDWARE | HW Problem | KLYS 9-5 is exhibiting strange bouts of large phase jitter (see attached FTP). This is not invasive to FACET operation, but it is contributing to general jitter in the machine.
120896 | 41953.5031 | 1 | LI09 | 51 | 4 | HARDWARE | HW Problem | KLYS 9-5 is exhibiting strange bouts of large phase jitter (see attached FTP). This is not invasive to FACET operation, but it is contributing to general jitter in the machine.
120953 | 41955.6006 | 1 | LI05 | 51 | 4 | HARDWARE | HW Problem | 5-5 Core bias variac needs to be replaced
120965 | 41956.0386 |  | LI20 | 41 | 4 | HARDWARE | HW Problem | Unable to rest. Mod fault lockout. Klystron LI20-4B Pad gain 0.0. as well.
120965 | 41956.0386 |  | LI20 | 41 | 4 | HARDWARE | HW Problem | Unable to rest. Mod fault lockout. Klystron LI20-4B Pad gain 0.0. as well.
121102 | 41962.4707 | 2 | LI11 | 11 | 4 | HARDWARE | HW Problem | 11-1: Replace HV rectifier stack & Thyratron blower
121102 | 41962.4707 | 2 | LI11 | 11 | 4 | HARDWARE | HW Problem | 11-1: Replace HV rectifier stack & Thyratron blower
120117 | 41926.6182 |  | LI30 | 61 | 4 | HARDWARE | HW Problem | 30-6 shows 461 faults since 8am today, we saw some KLT trips.
120117 | 41926.6182 | 1 | LI30 | 61 | 4 | HARDWARE | HW Problem | 30-6 shows 461 faults since 8am today, we saw some KLT trips.
120243 | 41929.4459 | 1 | LI02 | 51 | 4 | HARDWARE | HW Problem | Fire techs have replaced fire alarm monitoring modules for modulators 2-5 and 2-6. The resistors across the feedthroughs at the top of the modulators, (nside the little box that the red cable comes out) are 3.6 k ohms. This resistor value is not compatible with the new monitoring modules and therefore should be changed to 470 ohms (1/4 watt to 2 watt). This work should be coordinated with the fire tech because of the chance of causing a false fire alarm. Fire alarm monitoring for these two modulators should be bypassed (fire techs do this) while the resistor is being replaced. Mike Madayag (x4013) is the requester. Account # 15569-A8001
120243 | 41929.4459 | 2 | LI02 | 51 | 4 | HARDWARE | HW Problem | Fire techs have replaced fire alarm monitoring modules for modulators 2-5 and 2-6. The resistors across the feedthroughs at the top of the modulators, (nside the little box that the red cable comes out) are 3.6 k ohms. This resistor value is not compatible with the new monitoring modules and therefore should be changed to 470 ohms (1/4 watt to 2 watt). This work should be coordinated with the fire tech because of the chance of causing a false fire alarm. Fire alarm monitoring for these two modulators should be bypassed (fire techs do this) while the resistor is being replaced. Mike Madayag (x4013) is the requester. Account # 15569-A8001
120243 | 41929.4459 | 2 | LI02 | 51 | 4 | HARDWARE | HW Problem | Fire techs have replaced fire alarm monitoring modules for modulators 2-5 and 2-6. The resistors across the feedthroughs at the top of the modulators, (nside the little box that the red cable comes out) are 3.6 k ohms. This resistor value is not compatible with the new monitoring modules and therefore should be changed to 470 ohms (1/4 watt to 2 watt). This work should be coordinated with the fire tech because of the chance of causing a false fire alarm. Fire alarm monitoring for these two modulators should be bypassed (fire techs do this) while the resistor is being replaced. Mike Madayag (x4013) is the requester. Account # 15569-A8001
120245 | 41929.4676 | 15 | LI05 | 71 | 4 | HARDWARE | HW Problem | Klystron mod fault 5-7. Unable to remote reset.
120146 | 41927.3861 | 0.7 | LI25 | 71 | 4 | HARDWARE | HW Problem | The phase and amplitude readbacks for klystron LI25 71 are not available. Tried IPLing PIOP with no improvement. All other status looks good. Suspect bad PAD.
120316 | 41932.4065 | 0.5 | LI20 | 51 | 4 | HARDWARE | HW Problem | Klys 20-5 Mod fault - no remote reset
120502 | 41940.4509 | 0.5 | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 PEM and engineering installing filtering on dequing circuit could not find job on PAMM list.
119385 | 41903.8809 |  | LI29 | 71 | 4 | HARDWARE | HW Problem | Malfunctioning PIOP status. Klystron 29-7 and 25-1 report incorrect numbers in the value column. It flashes briefly but repeats often. PEM says klystrons appear to be function perfectly when checked locally.
120722 | 41948.4687 | 1 | LI29 | 51 | 4 | HARDWARE | HW Problem | 29-5 DQing current dropping when beam voltage raised above 340KV beam voltage.
119980 | 41921.7744 |  | LI09 | 81 | 4 | HARDWARE | HW Problem | With VVS 5 on, having klystron 9-8 doesn't make the red/green aisle lights flash red. With klystron 9-7 on, the red/green aisle lights correctly flash red.
120376 | 41934.4555 | 3 | LI05 | 41 | 4 | HARDWARE | HW Problem | 5-4 MOD fault lockout, won't reset. 25 TOC's on the error counter.
120376 | 41934.4555 | 1.5 | LI05 | 41 | 4 | HARDWARE | HW Problem | 5-4 MOD fault lockout, won't reset. 25 TOC's on the error counter.
121396 | 41975.2787 | 1 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 Beam Volts Jitter
119070 | 41887.7442 |  | LI14 | 61 | 4 | HARDWARE | HW Problem | Water leak from fitting in front of fiat rack for mod14-6.
119815 | 41916.702 | 0.75 | LI22 | 81 | 4 | HARDWARE | HW Problem | 22-8: MOD fault, no remote reset
119909 | 41920.1833 | 0.5 | LI28 | 81 | 4 | HARDWARE | HW Problem | 28-8 MOD fault. Won't reset.
120336 | 41933.372 | 1.5 | LI21 | 31 | 4 | HARDWARE | HW Problem | Kly Mod 21-3: at next opportunity (PAMM), intall tail clipper stack in 21-3 kly mod. Take snapshot of BV waveform prior to install and after install.
120336 | 41933.372 | 6 | LI21 | 31 | 4 | HARDWARE | HW Problem | Kly Mod 21-3: at next opportunity (PAMM), intall tail clipper stack in 21-3 kly mod. Take snapshot of BV waveform prior to install and after install.
120369 | 41934.2591 | 0.8 | LI26 | 81 | 4 | HARDWARE | HW Problem | 28-6 has 566 amplitude faults in the past 22 hours. Other stations in the sector are closer to 20 faults. It always resets fine via EPICS, but the modulator probably needs some attention.
120369 | 41934.2591 | 6 | LI26 | 81 | 4 | HARDWARE | HW Problem | 28-6 has 566 amplitude faults in the past 22 hours. Other stations in the sector are closer to 20 faults. It always resets fine via EPICS, but the modulator probably needs some attention.
120446 | 41936.6418 | 4 | LI28 | 61 | 4 | HARDWARE | HW Problem | 28-6 trips off to HVR excessively frequently (couple times per hour). We can reset it but it is a sever nuisance. Set to standby please address.
120680 | 41947.4076 | 0.2 | LI28 | 21 | 4 | HARDWARE | HW Problem | Kly. LI 28 - 2 klys is "MOD" fault. Won't reset.
120034 | 41925.1055 |  | LI22 | 51 | 4 | HARDWARE | HW Problem | 22-5 MKSU fault, won't reset.
121011 | 41957.5614 | 1 | LI21 | 81 | 4 | HARDWARE | HW Problem | 21-8: This klystron has MOD faulted and will not reset remotely.
120611 | 41943.6929 |  | LI14 | 21 | 4 | HARDWARE | HW Problem | 14-2 This Klystron MOD faulted and won't reset remotely. VVS breaker tripped.
121487 | 41978.564 | 1 | LI26 | 51 | 4 | HARDWARE | HW Problem | 26-5 showing beam volts jitter
121501 | 41979.3322 |  | LI16 | 51 | 4 | HARDWARE | HW Problem | Lots of reflected power faults from 16-5. duplicate see 117530
121637 | 41985.4776 | 0.25 | LI23 | 61 | 4 | HARDWARE | HW Problem | 23-6 MOD fault no reset
121643 | 41985.6018 | 1 | LI26 | 31 | 4 | HARDWARE | HW Problem | 26-3 MOD fault no reset. AC wiring to CB7 burnt.
121643 | 41985.6018 | 4 | LI26 | 31 | 4 | HARDWARE | HW Problem | 26-3 MOD fault no reset. AC wiring to CB7 burnt.
121645 | 41985.6319 | 0.2 | LI10 | 71 | 4 | HARDWARE | HW Problem | 10-7 MOD fault, won't remotely reset.
119829 | 41917.5369 | 0.5 | LI21 | 11 | 4 | HARDWARE | HW Problem | L1S 21-1 is having frequent beam volts dropouts.
120437 | 41936.5013 |  | LI30 | 81 | 4 | HARDWARE | HW Problem | 30-8: RF AMPL Low OOT. BV and timing OK. Did not check SAT.
120706 | 41947.6411 |  | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2: beam voltage is low due to radiation. Changes to the station might have reduced radiation sufficiently to allow higher voltage operation.
120871 | 41952.5136 |  | LI18 | 41 | 4 | HARDWARE | HW Problem | Klystron 18-4 has a CBL fault. It has been tripping off repeatedly in the last hour.
120903 | 41953.6192 | 1.5 | LI02 | 21 | 4 | HARDWARE | HW Problem | Station 2-2 is tripped on a klystron vacuum fault
121025 | 41958.1413 | 0.5 | LI23 | 21 | 4 | HARDWARE | HW Problem | Klys 23-2 AMPL is wandering. Not maintaining a centroid, moving around enough to affect the beam a bit.
121025 | 41958.1413 | 1.5 | LI23 | 21 | 4 | HARDWARE | HW Problem | Klys 23-2 AMPL is wandering. Not maintaining a centroid, moving around enough to affect the beam a bit.
121025 | 41958.1413 | 0.5 | LI23 | 21 | 4 | HARDWARE | HW Problem | Klys 23-2 AMPL is wandering. Not maintaining a centroid, moving around enough to affect the beam a bit.
121063 | 41960.4415 | 0.25 | LI22 | 81 | 4 | HARDWARE | HW Problem | PEM dispatched
121108 | 41962.5452 | 1 | LI12 | 61 | 4 | HARDWARE | HW Problem | 12-6: Replace PFN compartment fan
120580 | 41942.5588 | 1 | LI15 | 41 | 4 | HARDWARE | HW Problem | station 15-4 exhibiting hi Beam Volts jitter and AMPL jitter (for some reason BV jitter only shows up on FTP when it's assigned to a beam code).
120728 | 41948.5939 | 1 | LI00 | 2 | 4 | HARDWARE | HW Problem | K02 Thyratron compartment cooling fan is dead
121196 | 41965.1804 | 0.5 | LI29 | 51 | 4 | HARDWARE | HW Problem | 29-5 MOD fault. No remote reset
120140 | 41927.2923 | 2 | LI14 | 31 | 4 | HARDWARE | HW Problem | Rad levels at station 14-3 are >5mR/h. Lead will need to be added to Klystron.
120140 | 41927.2923 | 8 | LI14 | 31 | 4 | HARDWARE | HW Problem | Rad levels at station 14-3 are >5mR/h. Lead will need to be added to Klystron.
120652 | 41946.4223 | 0.5 | LI15 | 51 | 4 | HARDWARE | HW Problem | Beam current start 334.0 A.s to 340.0 A.s. Beam voltage move 311.0 kV to 313.0 kV
121344 | 41969.6538 | 0.5 | LI22 | 81 | 4 | HARDWARE | HW Problem | Klystron 22-8 mod fault.
121718 | 41990.3743 |  | LI17 | 51 | 4 | HARDWARE | HW Problem | PAD phase is - 45.9 / Expected = 0.0 Forward Energy Peak = 37.54 / expected 38.0 Low RF power"
119821 | 41917.2263 |  | LI01 | 51 | 4 | HARDWARE | HW Problem | 1-5 has a WGV fault
120103 | 41926.4587 | 1.5 | LI02 | 81 | 4 | HARDWARE | HW Problem | 2-8: PAD -15V out of tolerance.
120199 | 41928.5778 | 2 | LI03 | 21 | 4 | HARDWARE | HW Problem | 3-2 Several bad PAD readings on klystron PIOP display.
120442 | 41936.6052 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 L1X structure temperature OOT Fac Inst Shop dispatched.
120442 | 41936.6052 | 1 | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 L1X structure temperature OOT Fac Inst Shop dispatched.
120442 | 41936.6052 | 1 | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 L1X structure temperature OOT Fac Inst Shop dispatched.
120442 | 41936.6052 | 1 | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 L1X structure temperature OOT Fac Inst Shop dispatched.
120442 | 41936.6052 | 1 | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 L1X structure temperature OOT Fac Inst Shop dispatched.
120442 | 41936.6052 | 1 | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 L1X structure temperature OOT Fac Inst Shop dispatched.
120442 | 41936.6052 | 1 | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 L1X structure temperature OOT Fac Inst Shop dispatched.
120442 | 41936.6052 | 1 | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 L1X structure temperature OOT Fac Inst Shop dispatched.
120442 | 41936.6052 | 0.3 | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 L1X structure temperature OOT Fac Inst Shop dispatched.
120601 | 41943.3031 |  | LI06 | 71 | 4 | HARDWARE | HW Problem | Klys 6-7 is AMM; has low RF power
120631 | 41945.4897 | 2 | LI01 | 51 | 4 | HARDWARE | HW Problem | Station 1-5 faulted on Klystron Vacuum. Appears to be a real leak or electrical leakage across the klystron ion pump. 5.5e10-7. Stabilized for now. While have to wait for the PAMM on Tuesday November 11/4/14
120860 | 41951.8969 |  | LI14 | 51 | 4 | HARDWARE | HW Problem | 14-5 MOD fault, won't reset.
121001 | 41957.4276 | 0.25 | LI17 | 41 | 4 | HARDWARE | HW Problem | KLYS 17-4 is MOD faulted and won't reset remotely.
121136 | 41963.5936 | 0.5 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 Mod EOLC stack needs to be replaced.
121148 | 41963.7252 | 0.1 | LI16 | 81 | 4 | HARDWARE | HW Problem | Won't reset.
121151 | 41963.7316 | 0.75 | LI09 | 21 | 4 | HARDWARE | HW Problem | Won't reset.
121185 | 41964.5728 |  | LI11 | 81 | 4 | HARDWARE | HW Problem | Tried doing phase scans on this station and noticed that the phase would not move beyond about 15 degrees. Also tried changing the PDES and trimming, but it would not trim to the new value.
121210 | 41965.9908 |  | LI17 | 61 | 4 | HARDWARE | HW Problem | Tried to reset but it failed to come back on.
119937 | 41920.6461 |  | LI02 | 31 | 4 | HARDWARE | HW Problem | RFS 2-3 has a klystron magnet breaker fault.
121251 | 41967.6659 | 0.75 | LI09 | 11 | 4 | HARDWARE | HW Problem | Klystron 9-1 mod fault. Unable to reset.
121289 | 41968.8397 |  | LI30 | 51 | 4 | HARDWARE | HW Problem | LI30 - 5 PIOP status on PAD is purple and has no read back on the panel. Tried to IPL the PIOP but doesn't fix it.
121510 | 41980.0137 | 0 | LI16 | 51 | 4 | HARDWARE | HW Problem | Modulator 16-5 is faulting on reflected power pretty frequently, it might need some attention.
121729 | 41990.7065 |  | LI20 | 93 | 4 | HARDWARE | HW Problem | Klystron LI20-3C is tripping on reflected energy fault. Klystron has been tripped off for more than 8 hrs. and started cycling on RE faults after resetting.
120435 | 41936.4493 |  | DR02 | 61 | 4 | HARDWARE | HW Problem | SDR RF tripping repeatedly on over current.
120449 | 41936.7666 | 0.5 | LI04 | 51 | 4 | HARDWARE | HW Problem | 4-5 Had greater than 40 re faults today
120458 | 41937.3716 | 1 | LI25 | 31 | 4 | HARDWARE | HW Problem | 25-3 found HVR reset okay back to HVR within two minutes.
120471 | 41938.626 | 1 | LI19 | 41 | 4 | HARDWARE | HW Problem | 19-4 Vacuum gauge controller failed. Tried activating the filament. no help.
120477 | 41939.3364 | 2 | LI21 | 11 | 4 | HARDWARE | HW Problem | 21-1 the PLC code should be updated to regulate the thyratron reservoir on voltage as was demonstrated in 29-1
120182 | 41928.2313 |  | LI05 | 71 | 4 | HARDWARE | HW Problem | 5-7 MOD fault no reset
120209 | 41928.6847 | 0.1 | LI17 | 21 | 4 | HARDWARE | HW Problem | 17-2 Bad high power attenuator. Replace IPA when spares are available.
119935 | 41920.6419 | 1 | LI10 | 21 | 4 | HARDWARE | HW Problem | 10-2 shows EVOC fault.
119970 | 41921.6007 | 1 | LI02 | 51 | 4 | HARDWARE | HW Problem | 2-5 is faulted on waveguide vacuum.
120479 | 41939.3791 | 1.5 | LI29 | 21 | 4 | HARDWARE | HW Problem | 29-2 has low beam voltage for excessive radiation. Several changes have been made to the station since it was lowered. The beam voltage should be increased and the station re-surveyed.
120527 | 41940.917 | 2 | LI30 | 81 | 4 | HARDWARE | HW Problem | This klystron has been in amplitude mean status since the start of swing shift. However, the beam volts are ok. Is this just a software limit issue?
120561 | 41942.3339 | 1.5 | LI04 | 71 | 4 | HARDWARE | HW Problem | 4-7: Beam voltage is low, raise beam voltage, saturate station, and green piop
119682 | 41913.3133 | 1 | LI00 | 11 | 4 | HARDWARE | HW Problem | 01-1A MKSU water summary shows ACC1 flashing.
119714 | 41913.5594 | 0.5 | LI30 | 71 | 4 | HARDWARE | HW Problem | Kly. 30 - 7 tripped off, won't come back on.Errors are 2 vac. faults, 1 drive fault, and 2 TOC faults.
119733 | 41913.7267 | 1 | LI23 | 21 | 4 | HARDWARE | HW Problem | 23-2: MOD fault, no remote reset
119765 | 41914.8046 | 0.5 | LI21 | 11 | 4 | HARDWARE | HW Problem | 21-1 L1S PIOP no IPL
119918 | 41920.3489 |  | LI21 | 51 | 4 | HARDWARE | HW Problem | 21-5 K26 contacts need to be replaced. T20 input snubber assem wire badlt burnt Cable between despiking coil and thyratron needs to be replaced.
119986 | 41921.8216 | 1 | LI11 | 71 | 4 | HARDWARE | HW Problem | 11-7: K38 relay hangs up with CB7 on.
119986 | 41921.8216 | 1 | LI11 | 71 | 4 | HARDWARE | HW Problem | 11-7: K38 relay hangs up with CB7 on.
120029 | 41924.6239 | 0.2 | LI26 | 11 | 4 | HARDWARE | HW Problem | 26-1: KMC fault.
120127 | 41926.8198 | 0.5 | LI06 | 71 | 4 | HARDWARE | HW Problem | The HV/RESET light stays on.
120268 | 41930.313 | 0.5 | LI09 | 71 | 4 | HARDWARE | HW Problem | Lots of OV and EOLC faults.
120374 | 41934.4183 |  | LI26 | 41 | 4 | HARDWARE | HW Problem | 26-4: MKSU fault, no remote reset
120358 | 41933.6203 | 0.5 | LI28 | 61 | 4 | HARDWARE | HW Problem | 28-6 repeated MOD faults, won't stay on.
120434 | 41936.4407 | 0.5 | LI02 | 61 | 4 | HARDWARE | HW Problem | 2-6 WGV.
120438 | 41936.5069 | 0.5 | LI02 | 61 | 4 | HARDWARE | HW Problem | 2-6 VGC faulted
120460 | 41937.8662 | 1 | LI02 | 61 | 4 | HARDWARE | HW Problem | WGV on station 2-6.
121046 | 41960.0174 | 1 | LI24 | 21 | 4 | HARDWARE | HW Problem | Klys 24-2 beam volts are kind of ratty
121179 | 41964.4773 | 2 | LI28 | 61 | 4 | HARDWARE | HW Problem | 28-6 is tripping off repeatedly, took off the beam.
121845 | 42009.7824 | 0.5 | LI28 | 51 | 4 | HARDWARE | HW Problem | Additions to previous group CATER #121830
120615 | 41943.8645 | 1 | LI21 | 11 | 4 | HARDWARE | HW Problem | Station 21-1 (L1S) tripped off for the second time this shift. The first recovery took ~1hr.
120657 | 41946.4789 |  | LI20 | 93 | 4 | HARDWARE | HW Problem | Amplit. DES. should be set to the proper set point.
120696 | 41947.5978 |  | LI29 | 81 | 4 | HARDWARE | HW Problem | Beam voltage and current are steady.
120700 | 41947.6152 | 0.75 | LI23 | 81 | 4 | HARDWARE | HW Problem | Beam voltage and current are ok.
120717 | 41948.3613 | 1 | LI24 | 21 | 4 | HARDWARE | HW Problem | The station has tripped to HVR state three times in the last hour. It will reset then trip again soon after. Error counter shows water flow faults. See PIOP Messages on EPICS.
120873 | 41952.5931 |  | LI20 | 93 | 4 | HARDWARE | HW Problem | 20-3C not putting out any power. LI20 CR09 died - contains the SBI for klys 20-3C
120879 | 41953.0762 | 3 | LI30 | 41 | 4 | HARDWARE | HW Problem | Took 30-4 off the beam because it is very unstable; seems to cycle every few minutes.
120879 | 41953.0762 | 2 | LI30 | 41 | 4 | HARDWARE | HW Problem | Took 30-4 off the beam because it is very unstable; seems to cycle every few minutes.
120879 | 41953.0762 | 1 | LI30 | 41 | 4 | HARDWARE | HW Problem | Took 30-4 off the beam because it is very unstable; seems to cycle every few minutes.
120879 | 41953.0762 | 4 | LI30 | 41 | 4 | HARDWARE | HW Problem | Took 30-4 off the beam because it is very unstable; seems to cycle every few minutes.
120879 | 41953.0762 | 3 | LI30 | 41 | 4 | HARDWARE | HW Problem | Took 30-4 off the beam because it is very unstable; seems to cycle every few minutes.
119013 | 41880.5282 |  | LI30 | 31 | 4 | HARDWARE | HW Problem | 30-3: Waveguied #2 fault on the MKSU. Valves are LOTOed.
119797 | 41915.6688 | 0.75 | LI23 | 81 | 4 | HARDWARE | HW Problem | Klys 23-8 MOD fault and will not reset remotely.
119813 | 41916.6635 | 2 | LI23 | 51 | 4 | HARDWARE | HW Problem | 23-5: MOD fault, no remote reset
120655 | 41946.4522 |  | LI09 | 31 | 4 | HARDWARE | HW Problem | 9-3: Beam volts is too high, RF AMPL is too low and there is lots of beam volts jitter.
120656 | 41946.4585 | 0.5 | LI15 | 61 | 4 | HARDWARE | HW Problem | Beam current started at 375.2A.s and climbed to 378.4 A.s Beam voltage started at 311.0 kV and climbed to 313.5 kV.
120662 | 41946.5812 | 2 | LI26 | 71 | 4 | HARDWARE | HW Problem | AIP move from Build 15 to Sec 26-7 and connect all connection
119846 | 41918.4192 |  | LI26 | 51 | 4 | HARDWARE | HW Problem | 26-5 is tripped off, can't reset remotely. See 118725
119999 | 41922.4931 |  | LI22 | 81 | 4 | HARDWARE | HW Problem | 22-8: "AMM" low due to beam volts low.
120251 | 41929.5776 | 1 | LI28 | 21 | 4 | HARDWARE | HW Problem | Unable to reset fault on station 28-2.
120996 | 41957.306 | 1.5 | LI00 | 12 | 4 | HARDWARE | HW Problem | 1-1B has door light on with MOD running.
120996 | 41957.306 |  | LI00 | 12 | 4 | HARDWARE | HW Problem | 1-1B has door light on with MOD running.
121098 | 41962.3986 | 3 | LI29 | 31 | 4 | HARDWARE | HW Problem | 29-3 tripped off. Unable to reset remotely.
120890 | 41953.4058 | 0 | LI23 | 71 | 4 | HARDWARE | HW Problem | Kly. LI23 - 7 tripped on "MOD" fault, Thyratron U/C. The meter is pegged high.
120913 | 41954.6099 | 0.5 | LI21 | 11 | 4 | HARDWARE | HW Problem | 21-1 (L1S) is MOD faulted and won't reset remotely.
121007 | 41957.4927 | 2 | LI30 | 51 | 4 | HARDWARE | HW Problem | Kly Mod 30-5: keeps cycling on external faults.
121007 | 41957.4927 | 0.5 | LI30 | 51 | 4 | HARDWARE | HW Problem | Kly Mod 30-5: keeps cycling on external faults.
121041 | 41959.5847 | 0.2 | LI29 | 51 | 4 | HARDWARE | HW Problem | Dispatched PEM
119763 | 41914.7198 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 L1X BV dropouts.
119991 | 41922.1961 |  | LI29 | 51 | 4 | HARDWARE | HW Problem | 29-5: MOD fault, no remote reset.
119991 | 41922.1961 | 0.5 | LI29 | 51 | 4 | HARDWARE | HW Problem | 29-5: MOD fault, no remote reset.
120286 | 41931.1566 |  | LI00 | 13 | 4 | HARDWARE | HW Problem | LI0-3 MOD fault, won't reset
120699 | 41947.6064 |  | LI21 | 41 | 4 | HARDWARE | HW Problem | + / - 5 (see pic.s in MCC log.)
120732 | 41948.6012 | 2 | LI02 | 41 | 4 | HARDWARE | HW Problem | 2-4 The cable between the T20 and HV rect stack ("B") is showing signs of overheating and needs to be replaced.
120737 | 41948.7146 | 0.75 | LI23 | 31 | 4 | HARDWARE | HW Problem | Seeing a lot of jitter on klystron 23-3. Please range the thyratron.
120774 | 41949.4689 | 1 | LI22 | 81 | 4 | HARDWARE | HW Problem | Was able to reset but it tripped a gain and won't reset.
119746 | 41914.3096 | 1 | LI17 | 81 | 4 | HARDWARE | HW Problem | LI17 Q801 fails to trim
119767 | 41915.0977 | 0.5 | LI29 | 51 | 4 | HARDWARE | HW Problem | Klystron LI29-51 is tripped off on Mod Fault and will not reset remotely.
120223 | 41929.1344 | 0.5 | LI21 | 51 | 4 | HARDWARE | HW Problem | 21-5 MOD fault no reset
120290 | 41931.3292 | 0 | LI20 | 93 | 4 | HARDWARE | HW Problem | 20-3C WGV fault. Won't reset.
121027 | 41958.2545 |  | LI09 | 11 | 4 | HARDWARE | HW Problem | Klys 9-1 and 9-2 keep losing phase after summer shutdown. The readback was solid last run, but this run the phases seem to get lost frequently and for extended periods of time. This makes phase scans difficult.
121075 | 41961.2243 | 0.5 | LI12 | 51 | 4 | HARDWARE | HW Problem | Klys 12-5 WGV, due to actual poor vacuum. VP30 doesn't seem to be pumping down.
121075 | 41961.2243 | 1 | LI12 | 51 | 4 | HARDWARE | HW Problem | Klys 12-5 WGV, due to actual poor vacuum. VP30 doesn't seem to be pumping down.
121147 | 41963.7219 | 0.75 | LI19 | 11 | 4 | HARDWARE | HW Problem | Tried to reset but won't reset.
121155 | 41963.7447 | 0.75 | LI03 | 61 | 4 | HARDWARE | HW Problem | Won't reset.
121162 | 41963.8133 | 0.5 | LI14 | 21 | 4 | HARDWARE | HW Problem | The feeder breaker tripped and was reset earlier this month of Nov 2014
121190 | 41964.7066 | 1 | LI00 | 11 | 4 | HARDWARE | HW Problem | 01-1A - we've had to reset it at least a few times per shift. RE faults
121190 | 41964.7066 |  | LI00 | 11 | 4 | HARDWARE | HW Problem | 01-1A - we've had to reset it at least a few times per shift. RE faults
121197 | 41965.2839 | 0.5 | LI17 | 71 | 4 | HARDWARE | HW Problem | 17-7 FLK. Won't reset.
119684 | 41913.3282 |  | LI15 | 81 | 4 | HARDWARE | HW Problem | 15-8 HVOC light on all the time
120792 | 41949.8536 |  | LI20 | 93 | 4 | HARDWARE | HW Problem | Klys 20-3C (LI20 klys 93), tripping on reflected energy.
120852 | 41951.5533 |  | LI30 | 71 | 4 | HARDWARE | HW Problem | Klys 30-7 has a MOD fault, but reads full BVLTs. Bad: PAD phase and amplitude, klystron drive, forward energy peak, peak micro perveance. Same symptoms as 30-5 yesterday, as chronicled in CATER #120812. Solution to that CATER included below: MKSU showing drive fault. Low klystron current indication on the PIOP made me measure the current out of the klystron tank BNC. Readings were 343.3KV and 371.8Amps. Reconnecting BNC cables corrected error annunciators on the MKSU. System is back to normal. Maybe false status.
120859 | 41951.8127 | 1 | LI14 | 61 | 4 | HARDWARE | HW Problem | Kly Mod 14:6: Unable to reset mod from mod fault.
120899 | 41953.5163 | 0.5 | LI14 | 61 | 4 | HARDWARE | HW Problem | KLYS 14-6 is MOD faulted and won't reset remotely.
120952 | 41955.5984 | 1 | LI04 | 41 | 4 | HARDWARE | HW Problem | 4-4 PFN compartment cooling fan is dead
119819 | 41916.7979 | 1 | LI26 | 71 | 4 | HARDWARE | HW Problem | 26-7 VGC chassis removed for use in 30-3.
119819 | 41916.7979 | 0.5 | LI26 | 71 | 4 | HARDWARE | HW Problem | 26-7 VGC chassis removed for use in 30-3.
119901 | 41919.8009 | 1 | LI27 | 41 | 4 | HARDWARE | HW Problem | stations 25-4,27-4, 28-2, 28-5, 28-8, 29-5, 30-7 all MOD FAULTED
119990 | 41922.1714 |  | LI19 | 11 | 4 | HARDWARE | HW Problem | 19-1 Trips on Klystron Vacuum. Lowered the BV to get the modulator running for PPS test. During test the Klys Vac Meter dropped on the ION pump P/S indicator controller. Cater to AMRF
120261 | 41930.193 | 1 | LI24 | 41 | 4 | HARDWARE | HW Problem | 24-4 has 477 AMPL errors so far today, 10 times as many as the other klystrons in LI04, and needs to be reset remotely at least once an hour.
121249 | 41967.6202 | 0.1 | LI30 | 71 | 4 | HARDWARE | HW Problem | Klystron 30-7 repeating modulator water fault.
121336 | 41969.5485 | 1 | LI02 | 81 | 4 | HARDWARE | HW Problem | 2-8: RF AMPL low due to Beam Volts low
121343 | 41969.6495 | 1 | LI29 | 61 | 4 | HARDWARE | HW Problem | Klystron 29-6 repeating mod fault. On the beam right now. Request that PEM take a look into Amp stability.
119614 | 41911.3335 | 1 | LI29 | 51 | 4 | HARDWARE | HW Problem | Klystron 29-5 MOD faulted and will not remote reset.
119907 | 41919.9081 | 0.5 | LI25 | 41 | 4 | HARDWARE | HW Problem | KLYS:LI25:41:WATRERRT shows over 8052 fault today. Looks like the Flow Switch is to close to it's trip point or it might be getting flaky.
120013 | 41922.9054 | 0.2 | LI09 | 81 | 4 | HARDWARE | HW Problem | 9-8 PAD is CBL
120062 | 41925.6219 | 0.01 | LI26 | 11 | 4 | HARDWARE | HW Problem | Klystron 26- electro-magnet current.
121029 | 41958.323 | 0.5 | LI17 | 81 | 4 | HARDWARE | HW Problem | 17-8 FLK will not reset. RE and MKOC reported?
121029 | 41958.323 | 1.5 | LI17 | 81 | 4 | HARDWARE | HW Problem | 17-8 FLK will not reset. RE and MKOC reported?
121106 | 41962.5404 | 1 | LI10 | 61 | 4 | HARDWARE | HW Problem | 10-6: Replace Thyratron Grid Trigger chassis
121106 | 41962.5404 | 1 | LI10 | 61 | 4 | HARDWARE | HW Problem | 10-6: Replace Thyratron Grid Trigger chassis
121154 | 41963.7423 | 0.75 | LI04 | 41 | 4 | HARDWARE | HW Problem | Won't reset.
121166 | 41963.9346 | 0.5 | LI21 | 81 | 4 | HARDWARE | HW Problem | 21-8 trips on FLK
121167 | 41964.0316 | 1 | LI22 | 31 | 4 | HARDWARE | HW Problem | 22-3 Phase and amplitude jitter (mostly phase) causing beam stability issues.
121167 | 41964.0316 | 1 | LI22 | 31 | 4 | HARDWARE | HW Problem | 22-3 Phase and amplitude jitter (mostly phase) causing beam stability issues.
121183 | 41964.5525 | 1 | LI09 | 21 | 4 | HARDWARE | HW Problem | This klystron MOD faulted and will not reset remotely
119915 | 41920.3487 | 1.5 | LI21 | 51 | 4 | HARDWARE | HW Problem | 21-5 K26 contacts need to be replaced. T20 input snubber assem wire badlt burnt Cable between despiking coil and thyratron needs to be replaced.
119915 | 41920.3487 | 1 | LI21 | 51 | 4 | HARDWARE | HW Problem | 21-5 K26 contacts need to be replaced. T20 input snubber assem wire badlt burnt Cable between despiking coil and thyratron needs to be replaced.
119950 | 41921.0561 | 0.5 | LI23 | 51 | 4 | HARDWARE | HW Problem | 23-5 is MOD faulted.
121392 | 41974.6015 | 1 | LI17 | 71 | 4 | HARDWARE | HW Problem | 17-7 HV meter cap is blown.
121438 | 41976.6119 | 2 | LI20 | 81 | 4 | HARDWARE | HW Problem | L0B HV timing sees random jumps of about 8.4 ns, it is nearly a two-state.
121443 | 41976.7447 | 0.8 | LI20 | 71 | 4 | HARDWARE | HW Problem | L0A DQing current is jumping between two states, 20 and 12 A: http://physics-elog.slac.stanford.edu/lclselog/show.jsp?dir=/2014/49/03.12&pos=2014-12-03T16:53:48 So we have also some 1% lower HV pulses which trip the beam.
119753 | 41914.5249 | 1 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4 electro-magnet current; focus supply looks like it isn't working.
119811 | 41916.4663 | 0 | LI27 | 61 | 4 | HARDWARE | HW Problem | Klystron 27-6 mod fault lockout.
121189 | 41964.6984 | 0.1 | LI11 | 81 | 4 | HARDWARE | HW Problem | station 11-8 amplitude dropped slightly more than a day ago. Beam volt and beam current histories show no change.
121203 | 41965.559 | 3 | DR13 | 1 | 4 | HARDWARE | HW Problem | DR13 klys 1, intermittent BV jitter
121209 | 41965.9022 |  | LI21 | 61 | 4 | HARDWARE | HW Problem | Electro magnet current supply down. Focus supply ground faulted due to water spray from cracked fitting on klystron tank return line. Checked with lead tech; inspection. /phill
121258 | 41967.7101 | 0.5 | LI20 | 51 | 4 | HARDWARE | HW Problem | Mod fault on 20-5 will not reset remotely.
120174 | 41927.9247 | 1 | LI18 | 41 | 4 | HARDWARE | HW Problem | 18-4 PAD has False DAC status. Tried cycling power and IPLing the PIOP
120764 | 41949.4328 | 0.25 | LI28 | 81 | 4 | HARDWARE | HW Problem | Kly. LI 28 - 8 tripped on Klystron Electro-Magnet Breaker.
121366 | 41972.8718 |  | LI28 | 41 | 4 | HARDWARE | HW Problem | Klystron 28-4 has had many AMM, MOD, and HVR faults this shift. We are able to reset them remotely, but they are a nuisance and affect beam stability. The Klystron Fault Count panel shows this station has had over 200 AMPL faults since this morning.
121496 | 41978.948 | 1 | LI29 | 51 | 4 | HARDWARE | HW Problem | Klystron 29-5 (KLYS:LI29:51) MOD faulted and won't reset
121514 | 41980.5506 | 0.2 | LI23 | 61 | 4 | HARDWARE | HW Problem | This klystron mod faulted and won't reset remotely
121515 | 41980.572 |  | LI18 | 21 | 4 | HARDWARE | HW Problem | 18-4: CB7 tripped. Indicates T20 leaking oil. Need to verify before turning station back on.
121545 | 41981.7127 | 0.1 | LI29 | 51 | 4 | HARDWARE | HW Problem | Klys 29-5 (KLYS:LI29:51) mod faulted and won't reset
119904 | 41919.9038 |  | LI27 | 11 | 4 | HARDWARE | HW Problem | KLYS:LI27:11:WATRERRT shows over 4359 fault today. Looks like the Flow Switch is to close to it's trip point or it might be getting flaky.
120302 | 41931.6661 | 0.5 | LI18 | 41 | 4 | HARDWARE | HW Problem | Klystron 18-4 mod fault.
120384 | 41934.6173 | 3 | LI20 | 41 | 4 | HARDWARE | HW Problem | FACET XTCAV: I see no output no drive on diagnostic PAD - is PAC ok?
120603 | 41943.3858 |  | LI29 | 11 | 4 | HARDWARE | HW Problem | LLRF PH4 diagnostic PAD has NO RF waveform signals. So no L2 phase history visible hear, also Li29 phase and PCAV4 are affected.
121338 | 41969.5496 | 0.25 | LI02 | 81 | 4 | HARDWARE | HW Problem | 2-8: RF AMPL low due to Beam Volts low
121340 | 41969.553 | 0.75 | LI11 | 61 | 4 | HARDWARE | HW Problem | 11-6: MOD fault, won't reset remotely
121375 | 41973.9903 |  | LI20 | 93 | 4 | HARDWARE | HW Problem | Klystron 20-3C is tripped on MOD fault lockout and won't reset remotely. This is holding off the positron beam, and thus preventing experimenters from taking data.
121390 | 41974.5989 | 1 | LI16 | 31 | 4 | HARDWARE | HW Problem | 16-3 Despiking coil resistor is broken
121391 | 41974.6003 | 1 | LI16 | 81 | 4 | HARDWARE | HW Problem | 16-8 PFN compartment cooling fan is dead
119591 | 41909.327 | 2 | LI21 | 31 | 4 | HARDWARE | HW Problem | Klys 21-3 is cycling/tripping far more than any other stations. Not sure if this has to do with a change made to the station over the down.
120091 | 41926.4135 | 0.8 | LI06 | 21 | 4 | HARDWARE | HW Problem | 6-2 Holds HV but not triggering
120091 | 41926.4135 | 1 | LI06 | 21 | 4 | HARDWARE | HW Problem | 6-2 Holds HV but not triggering
120122 | 41926.7065 | 1 | LI05 | 41 | 4 | HARDWARE | HW Problem | 5-4 Found HV breaker tripped and the kly fil current was 1/2 of expected.kly fil volt was normal. Kly Fil Socket was very hot to touch.
120122 | 41926.7065 | 4 | LI05 | 41 | 4 | HARDWARE | HW Problem | 5-4 Found HV breaker tripped and the kly fil current was 1/2 of expected.kly fil volt was normal. Kly Fil Socket was very hot to touch.
121567 | 41982.5405 | 0.75 | LI27 | 31 | 4 | HARDWARE | HW Problem | 27-3: Seeing Beam Volts Jitter
121591 | 41983.5488 | 0.05 | LI06 | 41 | 4 | HARDWARE | HW Problem | 6-4: Bad PAD Parameters on PIOP Data
121631 | 41985.3968 | 0.5 | LI25 | 71 | 4 | HARDWARE | HW Problem | 25-7: Beam Volts Jitter
121640 | 41985.5508 | 0.5 | LI09 | 61 | 4 | HARDWARE | HW Problem | 9-6: MKSU "Online" LED is so dim it appears to be out.
121659 | 41987.4744 | 0.5 | LI30 | 41 | 4 | HARDWARE | HW Problem | Klystron 30-4 tripped off and will not reset remotely
121458 | 41977.5739 | 2 | LI29 | 61 | 4 | HARDWARE | HW Problem | 29-6 MOD: Touch panel com port not working.
121475 | 41978.0709 | 1 | LI05 | 21 | 4 | HARDWARE | HW Problem | Reset but it tripped again.
119430 | 41905.5954 | 3.5 | LI21 |  | 4 | HARDWARE | HW Problem | LCLS Kly Mods: manufacture and install 4 new tail clipper stacks in next 4 upstream AIP modulators from L1S.
119430 | 41905.5954 | 4.5 | LI21 |  | 4 | HARDWARE | HW Problem | LCLS Kly Mods: manufacture and install 4 new tail clipper stacks in next 4 upstream AIP modulators from L1S.
119430 | 41905.5954 | 3.5 | LI21 |  | 4 | HARDWARE | HW Problem | LCLS Kly Mods: manufacture and install 4 new tail clipper stacks in next 4 upstream AIP modulators from L1S.
119430 | 41905.5954 | 5 | LI21 |  | 4 | HARDWARE | HW Problem | LCLS Kly Mods: manufacture and install 4 new tail clipper stacks in next 4 upstream AIP modulators from L1S.
119828 | 41917.4491 | 0.5 | LI23 | 51 | 4 | HARDWARE | HW Problem | Klystron 23-5, looks like a real mod fault now. Ampl 0.0. Unable to reset remotely.
119916 | 41920.3487 |  | LI21 | 51 | 4 | HARDWARE | HW Problem | 21-5 K26 contacts need to be replaced. T20 input snubber assem wire badlt burnt Cable between despiking coil and thyratron needs to be replaced.
119985 | 41921.8187 |  | LI09 | 81 | 4 | HARDWARE | HW Problem | 9-8: Fails PPS Flashing Red Light test.
120030 | 41924.6808 |  | LI28 | 11 | 4 | HARDWARE | HW Problem | DLWG Temp on 28-1 and 28-7 are both above the high alarm limit.
120924 | 41955.2946 |  | LI30 | 61 | 4 | HARDWARE | HW Problem | 30-6 Phase jitter is out of tol
121448 | 41977.1937 |  | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3 MOD fault no reset
121572 | 41982.5971 | 1 | LI19 | 31 | 4 | HARDWARE | HW Problem | 19-3 PFN compartment fan is dead
119020 | 41880.658 |  | LI19 | 11 | 4 | HARDWARE | HW Problem | 19-1 Evaluate elevated klystron vacuum. Pressure is 1.0E-7. Kly has leaky window and prior leak that was vac sealed.
120704 | 41947.6372 |  | LI28 | 51 | 4 | HARDWARE | HW Problem | 28-5: beam voltage is low due to radiation. Changes to the station might have reduced radiation sufficiently to allow higher voltage operation.
120829 | 41950.7091 |  | LI30 | 41 | 4 | HARDWARE | HW Problem | 30-4 has core bias problem. The fuse was blown out twice. Needs TS & R.
119650 | 41912.3195 |  | LI00 | 11 | 4 | HARDWARE | HW Problem | 1-1A VGC gauge fail interlock trips too early
119735 | 41913.7372 | 0.5 | LI29 | 51 | 4 | HARDWARE | HW Problem | 29-5: MOD fault, no remote reset
119761 | 41914.6666 | 0.5 | LI23 | 21 | 4 | HARDWARE | HW Problem | 23-2 tripped, can't reset remotely.
119762 | 41914.6671 | 0.5 | LI22 | 41 | 4 | HARDWARE | HW Problem | 22-4 WGV.
120734 | 41948.6068 | 1 | LI23 | 11 | 4 | HARDWARE | HW Problem | 23-1 KMC no reset.
120768 | 41949.4497 |  | LI21 | 81 | 4 | HARDWARE | HW Problem | Kly. LI 21 - 8 phase jitter for the last 3 days shows jumps of +3.5 deg.s This unit last did this is 9 -19 -14
120846 | 41951.4349 | 0.3 | LI16 | 81 | 4 | HARDWARE | HW Problem | Klystron 16-8 (KLYS:LI16:81) has a waveguide vacuum fault
121580 | 41983.1717 | 1 | LI23 | 81 | 4 | HARDWARE | HW Problem | Klystron 23-8 tripped off and will not reset remotely
119612 | 41911.1752 |  | LI26 | 51 | 4 | HARDWARE | HW Problem | Klys 26-5 is tripping off frequently causing some consternation.
119686 | 41913.332 |  | LI18 | 81 | 4 | HARDWARE | HW Problem | 18-8 VGC failed NIRP test, Gauge Fail trip point is out of adjustment
119823 | 41917.3601 | 1 | LI23 | 51 | 4 | HARDWARE | HW Problem | 23-5 has a false MOD fault indication on EPICS
120795 | 41950.033 | 0.5 | LI23 | 11 | 4 | HARDWARE | HW Problem | 23-1 is tripped off on KMB
120795 | 41950.033 |  | LI23 | 11 | 4 | HARDWARE | HW Problem | 23-1 is tripped off on KMB
120858 | 41951.7779 |  | LI14 | 61 | 4 | HARDWARE | HW Problem | 14-6 MOD fault, won't reset
120895 | 41953.5014 | 1 | LI02 | 31 | 4 | HARDWARE | HW Problem | KLYS 2-3 is exhibiting large phase jitter (see attached FTP). This is not invasive to FACET, but it does contribute to general jitter in the machine.
119845 | 41918.4079 |  | LI26 | 51 | 4 | HARDWARE | HW Problem | PEM requests PIOP swap for 26-5.
120020 | 41923.2197 | 1 | LI23 | 61 | 4 | HARDWARE | HW Problem | 23-6: MOD fault, no remote reset
120021 | 41923.2332 | 0.5 | LI22 | 41 | 4 | HARDWARE | HW Problem | 22-4 WGV fault.
120027 | 41923.7768 | 0.5 | LI20 | 51 | 4 | HARDWARE | HW Problem | 20-5 has klystron magnet current fault.
120097 | 41926.42 | 0.5 | LI09 | 81 | 4 | HARDWARE | HW Problem | 9-8 No turn on with green light & EXT fault light flashing
120190 | 41928.417 | 2 | LI20 | 81 | 4 | HARDWARE | HW Problem | Klystron mod fault 20-8 (L0B) again. Will replace the Thyratron at MCC discretion or on the next scheduled POMM. /phill
120589 | 41942.6837 | 0.5 | LI30 | 81 | 4 | HARDWARE | HW Problem | Klys 30-8 has bad AMPL jitter.
120589 | 41942.6837 | 0.75 | LI30 | 81 | 4 | HARDWARE | HW Problem | Klys 30-8 has bad AMPL jitter.
120592 | 41942.8062 | 0.7 | LI03 | 21 | 4 | HARDWARE | HW Problem | Klys 3-2 is tripping on reflected energy.
120735 | 41948.6357 | 0.5 | LI22 | 41 | 4 | HARDWARE | HW Problem | 22-4 WGV no reset
120738 | 41948.7157 | 0.75 | LI26 | 81 | 4 | HARDWARE | HW Problem | Seeing a lot of jitter on klystron 26-8 Please range the thyratron.
120749 | 41948.9059 | 0.8 | LI21 | 21 | 4 | HARDWARE | HW Problem | Klys 21-2 (L1X) experiencing bad AMPL jitter.
121035 | 41958.6712 | 2 | LI20 | 93 | 4 | HARDWARE | HW Problem | Kly Mod 20-3C has no RF amplitude. Traced to 20-9 CAMAC crate problem.
121035 | 41958.6712 |  | LI20 | 93 | 4 | HARDWARE | HW Problem | Kly Mod 20-3C has no RF amplitude. Traced to 20-9 CAMAC crate problem.
119587 | 41909.2766 |  | LI22 | 41 | 4 | HARDWARE | HW Problem | Klys 22-4 has been having a number of reflected energy faults (tripping repeatedly during shift)
120703 | 41947.6351 |  | LI29 | 51 | 4 | HARDWARE | HW Problem | 29-5 beam voltage is low due to radiation. Changes to the station might have reduced radiation sufficiently to allow higher voltage operation.
120854 | 41951.617 | 0.5 | LI12 | 81 | 4 | HARDWARE | HW Problem | 12-8 MKT fault.
120882 | 41953.2693 |  | LI30 | 41 | 4 | HARDWARE | HW Problem | 30-4 mod fault again. Station has been troublesome tonight. Unable to reset this time.
121034 | 41958.5051 | 1 | LI04 | 51 | 4 | HARDWARE | HW Problem | 4-5: RF AMPL low OOT due to beam volts low
121059 | 41960.4178 | 1 | LI13 | 51 | 4 | HARDWARE | HW Problem | Klys. 13-5 has a Water Waveguide fault and will not reset remotely.
121231 | 41967.3136 | 0.5 | LI11 | 61 | 4 | HARDWARE | HW Problem | 11-6 MOD fault
121248 | 41967.6173 | 0.1 | LI22 | 81 | 4 | HARDWARE | HW Problem | Klystron 22-8 mod fault.
121339 | 41969.5514 | 1 | LI19 | 41 | 4 | HARDWARE | HW Problem | 19-4: RF AMPL low due to Beam Volts low
121410 | 41975.5424 | 0.75 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3: DQing current shows 0A
121414 | 41975.5957 | 0.5 | LI22 | 71 | 4 | HARDWARE | HW Problem | 22-7 MOD fault no remote reset
119041 | 41885.556 |  | LI04 | 81 | 4 | HARDWARE | HW Problem | 4-8: Water dripping from overhead piping onto the floor between the klystron and the modulator cabinet.
119688 | 41913.341 |  | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3 is FLK
119704 | 41913.4372 | 1 | LI28 | 61 | 4 | HARDWARE | HW Problem | 28-6 keeps MOD or HVR faulting. It resets remotely, but each time it trips off the FEL is interrupted.
119806 | 41916.0263 | 1.5 | LI20 | 51 | 4 | HARDWARE | HW Problem | Klystron 20-5 TCAV0 mod fault lockout.
119806 | 41916.0263 | 1.5 | LI20 | 51 | 4 | HARDWARE | HW Problem | Klystron 20-5 TCAV0 mod fault lockout.
119840 | 41918.3753 | 0.5 | LI26 | 61 | 4 | HARDWARE | HW Problem | Klystron 26-6 MOD faulting frequently. It resets remotely for about a minute or two and then MOD faults again.
119893 | 41919.598 | 1 | LI22 | 21 | 4 | HARDWARE | HW Problem | 22-2 T20 Network resistors are Burnt/broken . (2 ea)
120043 | 41925.3231 | 2 | LI23 | 51 | 4 | HARDWARE | HW Problem | 23-5: MOD EOLC fault
120065 | 41925.6922 | 0.5 | LI23 | 11 | 4 | HARDWARE | HW Problem | Klys 23-1 has tripped on KMC, won't reset.
120189 | 41928.3809 | 0.5 | LI20 | 81 | 4 | HARDWARE | HW Problem | kly 20-8 (L0B) mod fault.
120949 | 41955.5951 | 1 | LI16 | 81 | 4 | HARDWARE | HW Problem | 16-8 T20 compartment cooling fan is dead
121346 | 41969.915 | 0.75 | LI17 | 61 | 4 | HARDWARE | HW Problem | KLY. LI 17 - 6 faulted on "WGV" and won't reset.
121621 | 41984.5849 |  | LI03 | 21 | 4 | HARDWARE | HW Problem | Since this station is so far upstream, each RE causes the machine to trip.
121621 | 41984.5849 |  | LI03 | 21 | 4 | HARDWARE | HW Problem | Since this station is so far upstream, each RE causes the machine to trip.
121632 | 41985.4313 | 0.5 | LI23 | 81 | 4 | HARDWARE | HW Problem | Beam volts FTP shows a jitter sigma of 0.28
121632 | 41985.4313 | 1 | LI23 | 81 | 4 | HARDWARE | HW Problem | Beam volts FTP shows a jitter sigma of 0.28
119882 | 41919.0737 |  | LI21 | 31 | 4 | HARDWARE | HW Problem | 21-3 is faulting on over-voltage.
119892 | 41919.5952 | 2 | LI21 | 11 | 4 | HARDWARE | HW Problem | 21-1 T20 Input resisters and wiring shows signs of overheating and needs to be replaced.
119903 | 41919.9002 | 0.75 | LI28 | 61 | 4 | HARDWARE | HW Problem | KLYS:LI28:61:AMPL and BVLT came back low from before the PAMM. before after AMPL = 406.3 383.2 BVLT = 349.8 336.4
119911 | 41920.3063 |  | LI21 | 31 | 4 | HARDWARE | HW Problem | 21-3 reports of arcing sound near CCG, goes away VGC is temporarily turned off
119912 | 41920.3065 |  | LI12 | 51 | 4 | HARDWARE | HW Problem | 12-5 fails NIRP test:
120011 | 41922.7486 |  | LI09 | 81 | 4 | HARDWARE | HW Problem | With Modulator 9-8 on, the Gallery Red/Green lights in sector 9 do not flash. (They do flash with 9-7 on).
120066 | 41925.7201 |  | LI00 | 96 | 4 | HARDWARE | HW Problem | Kly. LI00 - 96 (CIDMAS) PAD RF Pedestal" ( volts ) should read 0.50 but is reading 0.12
120144 | 41927.3497 | 0.5 | LI18 | 21 | 4 | HARDWARE | HW Problem | Mod fault on klystron 18-2.
120686 | 41947.4666 |  | LI25 | 11 | 4 | HARDWARE | HW Problem | Suspect the negative T.C. is a bad read back The 137 temp. might be real. KLYS:LI25:11:WNDW1_TEMP Klystron window 1 reads 135 DegF, this is real temp. KLYS:LI25:11:WNDW2_TEMP
120693 | 41947.5718 | 0.1 | LI09 | 31 | 4 | HARDWARE | HW Problem | 9-3 low amplitude, changing drive request in SCP does nothing. https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=746796
120709 | 41947.7307 | 0.75 | LI29 | 81 | 4 | HARDWARE | HW Problem | Klys 29-8 has some excessive beam volts jitter.
120855 | 41951.6305 | 0.75 | LI26 | 71 | 4 | HARDWARE | HW Problem | 26-7 is showing VVS fault. The rest of the sector looks good.
120863 | 41952.0557 |  | LI20 | 93 | 4 | HARDWARE | HW Problem | 20-3c back on, but it quickly trips off again.
120927 | 41955.3031 |  | LI29 | 81 | 4 | HARDWARE | HW Problem | 29-8 beam volts jitter OOT
119607 | 41911.0026 |  | LI27 | 41 | 4 | HARDWARE | HW Problem | 27-4 has a 1 drip/2 seconds water leak at the klystron LCW return hose coupling on the klystron
119812 | 41916.644 | 0.25 | LI30 | 31 | 4 | HARDWARE | HW Problem | 30-3: WGV fault, no remote reset
119919 | 41920.3489 |  | LI21 | 51 | 4 | HARDWARE | HW Problem | 21-5 K26 contacts need to be replaced. T20 input snubber assem wire badlt burnt Cable between despiking coil and thyratron needs to be replaced.
119981 | 41921.7763 | 1 | LI22 | 21 | 4 | HARDWARE | HW Problem | AMRF says station 22-2 needs a new MKSU.
121435 | 41976.5496 |  | LI22 | 81 | 4 | HARDWARE | HW Problem | 22-8: RF AMPL low OOT due to Beam Volts low
121639 | 41985.5488 | 1 | LI12 | 61 | 4 | HARDWARE | HW Problem | 12-6: Modulator HV ON lamp not lit when MOD running. Found damaged wiring close to lamp socket, lamp itself is OK.
121657 | 41987.2261 |  | LI22 | 81 | 4 | HARDWARE | HW Problem | klystron 22-8 tripped and will not reset remotely
121713 | 41989.8046 | 2 | LI16 | 41 | 4 | HARDWARE | HW Problem | Klystron 16-4 mod fault.
121749 | 41991.8491 | 0.75 | LI02 | 11 | 4 | HARDWARE | HW Problem | Klystron 2-4 waveguide vacuum fault.
119687 | 41913.3327 |  | LI15 | 31 | 4 | HARDWARE | HW Problem | 15-3 VGC failed NIRP test, Over Pressure trip point is out of adjustment
119808 | 41916.272 | 0.5 | LI30 | 31 | 4 | HARDWARE | HW Problem | 30-3 is WGV faulted.
119930 | 41920.5838 | 1 | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7 Mod needs to have 1ea Carborundum resistor replaced.
120193 | 41928.4559 |  | LI04 | 81 | 4 | HARDWARE | HW Problem | KLYS 4-8 DLWG analog history shows temperature fluctuating between 113 and 108 degF over last day or two. See plot: http://physics-elog.slac.stanford.edu/facetelog/show.jsp?dir=/2014/42/16.10&pos=2014-10-16T10:53:10
120193 | 41928.4559 |  | LI04 | 81 | 4 | HARDWARE | HW Problem | KLYS 4-8 DLWG analog history shows temperature fluctuating between 113 and 108 degF over last day or two. See plot: http://physics-elog.slac.stanford.edu/facetelog/show.jsp?dir=/2014/42/16.10&pos=2014-10-16T10:53:10
120193 | 41928.4559 |  | LI04 | 81 | 4 | HARDWARE | HW Problem | KLYS 4-8 DLWG analog history shows temperature fluctuating between 113 and 108 degF over last day or two. See plot: http://physics-elog.slac.stanford.edu/facetelog/show.jsp?dir=/2014/42/16.10&pos=2014-10-16T10:53:10
120654 | 41946.445 | 0.5 | LI15 | 41 | 4 | HARDWARE | HW Problem | Beam voltage started at 393.0A.s and climbed to 397.4 A.s Beam current started at 344.8 kV and climbed to 349.7 kV.
120672 | 41947.1643 | 6 | LI06 | 81 | 4 | HARDWARE | HW Problem | 6-8 is MOD faulted and won't reset. VVS breaker reported tripped. Please reset. klystron TBR
120672 | 41947.1643 | 2 | LI06 | 81 | 4 | HARDWARE | HW Problem | 6-8 is MOD faulted and won't reset. VVS breaker reported tripped. Please reset. klystron TBR
120672 | 41947.1643 | 0.5 | LI06 | 81 | 4 | HARDWARE | HW Problem | 6-8 is MOD faulted and won't reset. VVS breaker reported tripped. Please reset. klystron TBR
120672 | 41947.1643 | 0.5 | LI06 | 81 | 4 | HARDWARE | HW Problem | 6-8 is MOD faulted and won't reset. VVS breaker reported tripped. Please reset. klystron TBR
120786 | 41949.6796 | 1 | LI29 | 81 | 4 | HARDWARE | HW Problem | Klys 29-8 tripped. Was showing tons of beam volts jitter.
121546 | 41981.7835 |  | LI20 | 81 | 4 | HARDWARE | HW Problem | On recovery from a PPS trip 12/8/14, L0B Acc #2 flow switch was tripped. Was able to clear it by adjusting Temperature setpoint, but it seemed to trip again when the flow got too high. See plot in MCC elog. There shouldn't be a high trip. It took multiple iterations of temperature setpoint adjustment to get it back on. These flow switches are in the Injector Vault.
119455 | 41906.2193 |  | LI26 | 51 | 4 | HARDWARE | HW Problem | Klystron 26-5 MOD faults frequently.
119729 | 41913.6593 | 1.5 | LI24 | 31 | 4 | HARDWARE | HW Problem | TBR 24-3 tripped on "OC". Now "MOD" fault lock out.
119965 | 41921.5135 | 1 | LI17 | 61 | 4 | HARDWARE | HW Problem | 17-6 has a KMC fault.
121798 | 42009.3022 |  | LI25 | 11 | 4 | HARDWARE | HW Problem | 25-1 Has a noisy Thyratron compartment fan.
121802 | 42009.3055 | 2 | LI29 | 41 | 4 | HARDWARE | HW Problem | 29-4 klystron filament indicates 6 amps at 1.5 volts even with 110vac to the ac controller turned off. Measured 6 volts out of the current transducer for the kly filament current with the kly filament power breaker off.
119249 | 41898.4764 |  | LI02 | 51 | 4 | HARDWARE | HW Problem | MOD 2-5: WGVGC has a broken meter glass face.
119890 | 41919.5465 | 1 | LI11 | 81 | 4 | HARDWARE | HW Problem | 11-8: KMC fault on MKSU.
119905 | 41919.9057 | 0.5 | LI26 | 21 | 4 | HARDWARE | HW Problem | KLYS:LI26:21:WATRERRT shows over 2416 fault today. Looks like the Flow Switch is to close to it's trip point or it might be getting flaky.
120712 | 41948.2548 | 2 | LI25 | 21 | 4 | HARDWARE | HW Problem | This klystron mod faulted and cannot be reset remotely.
120867 | 41952.3439 |  | LI02 | 71 | 4 | HARDWARE | HW Problem | 2-7 high voltage meter doesn't indicate any voltage when on
121383 | 41974.5464 | 0.25 | LI25 | 61 | 4 | HARDWARE | HW Problem | 25-6 MOD fault. Won't reset.
121388 | 41974.5954 | 1 | LI15 | 41 | 4 | HARDWARE | HW Problem | 15-4 Mod EOLC stack and 1ea Carborundum resistor needs to be replaced.
121460 | 41977.608 | 2 | LI22 | 71 | 4 | HARDWARE | HW Problem | 22-7 EOLC has many loose connections needs to be replaced before it overheats
121499 | 41979.2726 | 0.5 | LI28 | 61 | 4 | HARDWARE | HW Problem | 28-6 is tripping off frequently, and has 206 AMPL faults in the past 24 hours. It resets fine every time, but trips off too frequently to use. No tripping at this time.
119040 | 41885.5543 |  | LI10 | 31 | 4 | HARDWARE | HW Problem | 10-3: Water dripping from overhead piping onto the fiat rack.
119449 | 41905.7385 |  | LI29 | 71 | 4 | HARDWARE | HW Problem | 29-7 has tripped several times with fault log showing HVOC, HVOV, and low voltage PS fault (Reg Bd). Regulator board may be failing.
119633 | 41911.7609 | 0.5 | LI27 | 21 | 4 | HARDWARE | HW Problem | 27-2: Klystron is reporting terrible phase and amplitude jitter. FTP Images: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=739270
119835 | 41917.9731 | 0.5 | LI27 | 61 | 4 | HARDWARE | HW Problem | station 27-6 reports 375 water faults since 08:00 this morning. Waveguide 1 water.
119931 | 41920.585 | 1 | LI24 | 81 | 4 | HARDWARE | HW Problem | 24-8 Mod T20 has a burnt/broken Network resistor.
119936 | 41920.6441 |  | LI15 | 31 | 4 | HARDWARE | HW Problem | 15-3 has a waveguide vacuum fault.
120751 | 41949.0042 | 1.5 | LI20 |  | 4 | HARDWARE | HW Problem | Klystron 20-3C (LI20 Klys 93) amplitude has dropped to zero. Ran well for almost an hour after turn on, then suddenly dropped out.
121037 | 41959.2255 |  | LI20 | 41 | 4 | HARDWARE | HW Problem | Klystron 20-4B. Facet XTCAV mod fault. Reflected power.
121114 | 41962.8674 | 1.5 | LI13 | 21 | 4 | HARDWARE | HW Problem | 13-2 HV meter not reading
121831 | 42009.6082 | 1 | LI27 | 11 | 4 | HARDWARE | HW Problem | 27-1: MAIN contactor was not closed
119751 | 41914.4826 |  | LI21 | 51 | 4 | HARDWARE | HW Problem | 21-5 tripped, can't reset remotely.
120363 | 41933.7707 | 0.5 | LI09 | 81 | 4 | HARDWARE | HW Problem | 9-8 is tripping CB7, HV current is unstable as is PFN waveform. Won't trim phase. 36 deg off.
120363 | 41933.7707 | 1 | LI09 | 81 | 4 | HARDWARE | HW Problem | 9-8 is tripping CB7, HV current is unstable as is PFN waveform. Won't trim phase. 36 deg off.
120363 | 41933.7707 | 1 | LI09 | 81 | 4 | HARDWARE | HW Problem | 9-8 is tripping CB7, HV current is unstable as is PFN waveform. Won't trim phase. 36 deg off.
120363 | 41933.7707 | 2 | LI09 | 81 | 4 | HARDWARE | HW Problem | 9-8 is tripping CB7, HV current is unstable as is PFN waveform. Won't trim phase. 36 deg off.
121961 | 42012.4335 |  | LI14 | 21 | 4 | HARDWARE | HW Problem | W/G strainer at station 14-2 is leaking.
122127 | 42018.539 | 1 | LI06 | 61 | 4 | HARDWARE | HW Problem | 6-6: T20: fan touching the side cover. 6-6: Primary capacitor C59, and C60 =25uF have a leak. 6-6: T20: Secondary network resistor at R1 is open. It should be measured = 20K-ohm.
122170 | 42020.2641 |  | LI25 | 11 | 4 | HARDWARE | HW Problem | 25-1 has faulted twice in the last ten minutes. RF amplitude drops by 75% & takes a minute or two to recover.
122170 | 42020.2641 | 1 | LI25 | 11 | 4 | HARDWARE | HW Problem | 25-1 has faulted twice in the last ten minutes. RF amplitude drops by 75% & takes a minute or two to recover.
122284 | 42025.8788 | 1 | LI27 | 61 | 4 | HARDWARE | HW Problem | Klys 27-6 has a mod fault and won't reset.
122284 | 42025.8788 | 0.8 | LI27 | 61 | 4 | HARDWARE | HW Problem | Klys 27-6 has a mod fault and won't reset.
122284 | 42025.8788 | 0.8 | LI27 | 61 | 4 | HARDWARE | HW Problem | Klys 27-6 has a mod fault and won't reset.
122285 | 42026.3078 | 2 | LI29 | 41 | 4 | HARDWARE | HW Problem | 29-4 MOD fault no reset.
122285 | 42026.3078 |  | LI29 | 41 | 4 | HARDWARE | HW Problem | 29-4 MOD fault no reset.
122313 | 42027.2328 | 0.5 | LI28 | 51 | 4 | HARDWARE | HW Problem | 28-5 MOD faulted and won't reset remotely
122335 | 42028.2069 | 1 | LI02 | 51 | 4 | HARDWARE | HW Problem | 2-5 VGC HV reset keeps fault tripping
123103 | 42061.2181 | 0.5 | LI06 | 81 | 4 | HARDWARE | HW Problem | Tried resetting but not coming back,
123155 | 42062.6887 | 3 | LI05 | 21 | 4 | HARDWARE | HW Problem | klys 5-2 MOD faulted and won't reset remotely
123299 | 42069.5987 | 0.75 | LI30 | 11 | 4 | HARDWARE | HW Problem | Puts it at the top of the BV Jitter Report
123471 | 42076.4364 | 1 | LI23 | 21 | 4 | HARDWARE | HW Problem | Looks like it's been doing it for at least the last 5 days
123471 | 42076.4364 | 0.8 | LI23 | 21 | 4 | HARDWARE | HW Problem | Looks like it's been doing it for at least the last 5 days
123505 | 42077.687 | 0.75 | LI00 | 1 | 4 | HARDWARE | HW Problem | 0-1 WGV fault.
123736 | 42087.5836 | 1.5 | LI00 | 1 | 4 | HARDWARE | HW Problem | K27 center contacts need to be replaced. T-20 Jumper cable needs to be replaced PFN Strap nut for the cap is missing (Upper, RHS, #2 )
122185 | 42020.5406 | 1.5 | LI17 | 21 | 4 | HARDWARE | HW Problem | 17-2: Here is the list below: 1. K26 needs to replace all three contactors (L1-L3). 2. K27 need to replace two contactors L1 and L2. 3. Primary resistor network: R262=10-ohm needs to replace. 4. Loose straps: upper (Lhs #4, Rhs #3); lower (Lhs #2,3,4). 5. EOLC stack has a broken cap.
122191 | 42020.7433 | 1 | LI22 | 21 | 4 | HARDWARE | HW Problem | Klys. 22-2 (KLYS:LI22:21) jitter causing many drop outs in FEL.
122302 | 42026.6247 | 1 | LI17 | 31 | 4 | HARDWARE | HW Problem | Kly Mod 17-3: use dipstick to check T20 xformer oil level. Add oil as necessary.
122331 | 42027.6478 | 1.5 | LI26 | 11 | 4 | HARDWARE | HW Problem | 26-1 MOD faulted
122511 | 42034.6215 | 1 | LI25 | 21 | 4 | HARDWARE | HW Problem | Amplt. jitter is 1.8 sigma.
123414 | 42074.4836 | 1.25 | LI03 | 31 | 4 | HARDWARE | HW Problem | 3-3 Thy compartment D Q ing cap has a badly burnt connection K26 has 1 pair of contacts that need to be replaced.
123692 | 42084.3744 | 1.5 | LI24 | 61 | 4 | HARDWARE | HW Problem | Duplicate See 124329(TEC) Station began tripping about once per minute as soon as we put it on the beam.
123692 | 42084.3744 | 1 | LI24 | 61 | 4 | HARDWARE | HW Problem | Duplicate See 124329(TEC) Station began tripping about once per minute as soon as we put it on the beam.
123692 | 42084.3744 | 1 | LI24 | 61 | 4 | HARDWARE | HW Problem | Duplicate See 124329(TEC) Station began tripping about once per minute as soon as we put it on the beam.
123692 | 42084.3744 | 1.5 | LI24 | 61 | 4 | HARDWARE | HW Problem | Duplicate See 124329(TEC) Station began tripping about once per minute as soon as we put it on the beam.
123692 | 42084.3744 | 6 | LI24 | 61 | 4 | HARDWARE | HW Problem | Duplicate See 124329(TEC) Station began tripping about once per minute as soon as we put it on the beam.
123692 | 42084.3744 |  | LI24 | 61 | 4 | HARDWARE | HW Problem | Duplicate See 124329(TEC) Station began tripping about once per minute as soon as we put it on the beam.
123697 | 42084.7631 | 0.75 | LI22 | 81 | 4 | HARDWARE | HW Problem | station 22-8 Mod faulted
123839 | 42090.6564 | 7 | LI25 | 71 | 4 | HARDWARE | HW Problem | Klys 25-7 has VVS MOD fault and can't be reset remotely. Found CB7 input wiring overheated.
123849 | 42091.1539 | 1 | LI23 | 21 | 4 | HARDWARE | HW Problem | This Klystron has tripped 6 times the last 3 hours but can still be reset remotely. The 23-2 plot of the beam volts on the fast time plot also looks super jittery.
123849 | 42091.1539 | 2 | LI23 | 21 | 4 | HARDWARE | HW Problem | This Klystron has tripped 6 times the last 3 hours but can still be reset remotely. The 23-2 plot of the beam volts on the fast time plot also looks super jittery.
123849 | 42091.1539 | 1 | LI23 | 21 | 4 | HARDWARE | HW Problem | This Klystron has tripped 6 times the last 3 hours but can still be reset remotely. The 23-2 plot of the beam volts on the fast time plot also looks super jittery.
122628 | 42039.6076 | 0.75 | LI30 | 71 | 4 | HARDWARE | HW Problem | 30-7: Thyratron replacement. Vres at 9.70V
122674 | 42040.8446 | 0.5 | LI22 | 81 | 4 | HARDWARE | HW Problem | 22-8 MOD fault; not reset remotely
122882 | 42051.2705 | 0.5 | LI12 | 21 | 4 | HARDWARE | HW Problem | Klystron 12-2 will not come back on remotely
122957 | 42054.5752 | 1 | LI17 | 21 | 4 | HARDWARE | HW Problem | 17-2 "MOD bit 4 " error on PIOP. CASSELL box needs to be replaced.
122959 | 42054.6035 | 0.5 | LI02 | 11 | 4 | HARDWARE | HW Problem | LI02-11 tripped off and will not reset remotely.
122982 | 42055.6189 | 0.5 | LI03 | 81 | 4 | HARDWARE | HW Problem | Klystron mod fault 3-8
122988 | 42056.9428 |  | LI21 | 31 | 4 | HARDWARE | HW Problem | Many over-current faults from 21-3. 10 since 08:00 this morning.
122988 | 42056.9428 |  | LI21 | 31 | 4 | HARDWARE | HW Problem | Many over-current faults from 21-3. 10 since 08:00 this morning.
122990 | 42057.1168 | 0.8 | LI15 | 51 | 4 | HARDWARE | HW Problem | 15-5 won't stay on, many TOCs, thyratron is bad
122990 | 42057.1168 | 0.75 | LI15 | 51 | 4 | HARDWARE | HW Problem | 15-5 won't stay on, many TOCs, thyratron is bad
123016 | 42058.427 | 0.5 | LI12 | 81 | 4 | HARDWARE | HW Problem | Klystron 12-8 mod fault. Heater delay (1HR)
121664 | 41987.9273 | 0.8 | LI30 | 41 | 4 | HARDWARE | HW Problem | Klystron 30-4 mod fault
121664 | 41987.9273 | 1 | LI30 | 41 | 4 | HARDWARE | HW Problem | Klystron 30-4 mod fault
121888 | 42010.6048 |  | LI21 | 41 | 4 | HARDWARE | HW Problem | 21-4 Need to adjust beam volts to spec (might need to change T20 taps)
121992 | 42013.1796 | 0.5 | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2 Thyr Klix flt will not clear.
122967 | 42054.8301 | 0.5 | LI17 | 61 | 4 | HARDWARE | HW Problem | Tried resetting three times but no change.
122967 | 42054.8301 | 1 | LI17 | 61 | 4 | HARDWARE | HW Problem | Tried resetting three times but no change.
123739 | 42087.5901 | 1.2 | LI01 | 41 | 4 | HARDWARE | HW Problem | 1-4 Kly Fil Variac is showing signs of overheating and is hot to touch...Needs to be replaced
123743 | 42087.5963 | 0.5 | LI06 | 81 | 4 | HARDWARE | HW Problem | 6-8 Mod HV meter cap is blown
124064 | 42098.4626 |  | LI14 | 81 | 4 | HARDWARE | HW Problem | Was able to reset it a few times
124187 | 42102.5753 | 0.75 | LI14 | 61 | 4 | HARDWARE | HW Problem | 14-6 Kly trim P/S needs to be replaced for degraded cooling Fan.
124197 | 42102.6503 | 1 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-5: PFN Tune to reduce jitter.
123326 | 42070.6559 | 0.5 | LI17 | 61 | 4 | HARDWARE | HW Problem | Klys 17-6 MOD fault.
123411 | 42074.4805 | 1 | LI02 | 61 | 4 | HARDWARE | HW Problem | 2-6 T20 has a badly burnt/broken network resistor
123499 | 42077.5559 | 2 | LI22 | 51 | 4 | HARDWARE | HW Problem | Kly Mod 22-5: station tripped off on klystron delta temperature fault. TSR station.
123499 | 42077.5559 |  | LI22 | 51 | 4 | HARDWARE | HW Problem | Kly Mod 22-5: station tripped off on klystron delta temperature fault. TSR station.
123702 | 42085.562 | 0.5 | LI03 | 21 | 4 | HARDWARE | HW Problem | 3-2 FLK, won't reset.
123816 | 42089.8477 |  | LI21 | 71 | 4 | HARDWARE | HW Problem | 21-7 klystron uk is drooping. Kly Installed 18-JUL-2013. S/N: 606A.
123817 | 42089.8671 |  | LI20 | 51 | 4 | HARDWARE | HW Problem | 20-5 MKSU reports Klystron focus is low and unstable at 13.5 Amps. expected is 14.5. LI20 focus magnet epics display indicates 14.5A.
122124 | 42018.4472 | 1 | LI25 | 51 | 4 | HARDWARE | HW Problem | 25-5 T20 compartment cooling fan is dead
122339 | 42028.4127 | 0.75 | LI10 | 41 | 4 | HARDWARE | HW Problem | 10-4: "Doors" fault on modulator.
122697 | 42041.9754 | 2 | LI09 | 41 | 4 | HARDWARE | HW Problem | 9-4 mod fault. Thought this was in "MNT" for T20?
122979 | 42055.5164 |  | LI27 | 21 | 4 | HARDWARE | HW Problem | 27-2 PHJ/AMJ not appearing to be a function of high voltage (BV).
123095 | 42060.8004 | 1 | LI26 | 71 | 4 | HARDWARE | HW Problem | This station cycled MOD fault four times in about 5 minutes. Took it off the beam and I noticed it has continued to cycle. Now has settled on a FLK error.
123095 | 42060.8004 | 2 | LI26 | 71 | 4 | HARDWARE | HW Problem | This station cycled MOD fault four times in about 5 minutes. Took it off the beam and I noticed it has continued to cycle. Now has settled on a FLK error.
123101 | 42061.1505 | 1 | LI09 | 71 | 4 | HARDWARE | HW Problem | Tried to reset 3 times but it didn't come back.
123185 | 42065.3845 | 0.5 | LI02 | 31 | 4 | HARDWARE | HW Problem | Klystron 2-3 has a waveguide Vacuum fault
123215 | 42066.4703 |  | LI30 | 61 | 4 | HARDWARE | HW Problem | 30-6 14.5A entered on kly enter panel for the focus magnet. Strip chart shows 15A and unstable read back. PIOP reflects this discrepancy.
123258 | 42068.3059 | 1.5 | LI24 | 11 | 4 | HARDWARE | HW Problem | 24-1: Pulse WIDTH = 0.000 uS; Slope = 1.1; Ripl = 0.2
123259 | 42068.3085 | 2.5 | LI24 | 41 | 4 | HARDWARE | HW Problem | 24-4: Pulse WIDTH = 2.85; Slope = .2; Ripl = 0.3
121715 | 41989.9876 | 0.75 | LI01 | 41 | 4 | HARDWARE | HW Problem | Klystron 1-4 mod fault.
122107 | 42017.5781 | 3 | LI05 | 81 | 4 | HARDWARE | HW Problem | 5-8: K26 L1, L2 contactors are showing burnt marks. 5-8: K27 L1 and L3 contactors are showing burnt marks. 5-8: Charging Diode Assembly stack has a broken cap lead. 5-8: PFN strap faulty (Lower, lhs, #1)
122333 | 42027.8005 | 3 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6: Upgraded Modulator PFN Pulse Width Very Short for OPS. Try to widen. Try for 3.4 uS.
122706 | 42043.2254 | 1 | LI20 | 41 | 4 | HARDWARE | HW Problem | 20-4B high voltage meter is inoperative
122993 | 42057.3805 | 2 | LI10 | 21 | 4 | HARDWARE | HW Problem | station 10-2 has Mod EOLC fault
122993 | 42057.3805 | 0.5 | LI10 | 21 | 4 | HARDWARE | HW Problem | station 10-2 has Mod EOLC fault
123134 | 42062.278 | 0.5 | LI22 | 71 | 4 | HARDWARE | HW Problem | Tried to reset 3 times but it didn't come back.
123197 | 42065.6005 |  | LI24 | 41 | 4 | HARDWARE | HW Problem | 24-4 Jitters have returned.
122208 | 42023.3814 | 2 | LI17 | 41 | 4 | HARDWARE | HW Problem | 17-4: Here is the list below: 1. T20 - AC input wire phase C loose. 2. EOLC stack need to replace. It have burnt caps. 3. Loose strap: upper (Lhs #2).
122391 | 42031.2343 | 0.5 | LI00 | 2 | 4 | HARDWARE | HW Problem | KO-2 WGV flt on scp
122803 | 42047.5561 | 1 | LI11 | 71 | 4 | HARDWARE | HW Problem | 11-7: Seeing some beam volts jitter.
122832 | 42048.5545 |  | LI30 | 71 | 4 | HARDWARE | HW Problem | 30-7: Reflected energy faults. May be beam volts related.
122849 | 42048.8264 | 1 | LI13 | 81 | 4 | HARDWARE | HW Problem | station 13-8 has a Waveguide Vacuum Fault
122891 | 42052.2805 | 1 | LI15 | 71 | 4 | HARDWARE | HW Problem | 15-7 MOD fault, won't reset
122891 | 42052.2805 | 1 | LI15 | 71 | 4 | HARDWARE | HW Problem | 15-7 MOD fault, won't reset
123230 | 42067.3304 | 3 | LI29 | 61 | 4 | HARDWARE | HW Problem | 29-6 has narrrow WIDTH = 2.79 uS, Slope = 0.5 and Ripl = 0.4
123286 | 42069.2676 | 3 | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3: pulse WIDTH= 2.97 uS, Slope=0.6, Ripl=0.2
123288 | 42069.3518 | 3 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 needs PFN tune. PW = 2.95uS; slope = 0.1; ripple = 0.6
123294 | 42069.4662 | 0.8 | LI26 | 51 | 4 | HARDWARE | HW Problem | 26-5: DQing Under Current bubbling up to LCLSHOME. DQing is 3 Amps.
123307 | 42069.8329 | 1 | LI02 | 31 | 4 | HARDWARE | HW Problem | LI02-3 MOD fault not reset
123391 | 42073.486 | 1 | LI01 | 41 | 4 | HARDWARE | HW Problem | 1-4 has a broken despiking coil resistor
123403 | 42074.3641 | 0.75 | LI21 | 51 | 4 | HARDWARE | HW Problem | Like to tweak it down.
123102 | 42061.2106 | 2 | LI09 | 81 | 4 | HARDWARE | HW Problem | 9-8 Needs 3 shorty PFN Caps replace due to arcing
123102 | 42061.2106 | 2 | LI09 | 81 | 4 | HARDWARE | HW Problem | 9-8 Needs 3 shorty PFN Caps replace due to arcing
123845 | 42090.7737 | 1 | LI02 | 61 | 4 | HARDWARE | HW Problem | 2-6: RF AMPL low OOT. BV and timing OK. Did not check saturation.
123847 | 42091.0242 | 1 | LI20 | 41 | 4 | HARDWARE | HW Problem | Facet XTCAV continually MOD faulting.
123860 | 42092.5206 | 0.5 | LI14 | 11 | 4 | HARDWARE | HW Problem | 14-1: Tried to reset twice but it failed.
123919 | 42094.4739 |  | LI24 | 61 | 4 | HARDWARE | HW Problem | Klystron LI24-61 is tripped and will not reset remotely. (Recently returned to service)
122418 | 42031.6403 | 1 | LI19 | 21 | 4 | HARDWARE | HW Problem | Kly Mod 19-2: check T20 oil level. Add oil as necessary to make T20 safe for HV ops.
122418 | 42031.6403 | 1 | LI19 | 21 | 4 | HARDWARE | HW Problem | Kly Mod 19-2: check T20 oil level. Add oil as necessary to make T20 safe for HV ops.
122562 | 42037.551 | 1 | LI26 | 11 | 4 | HARDWARE | HW Problem | 26-1: Beam volts jitter (spikes)
123386 | 42073.4802 | 0.75 | LI15 | 71 | 4 | HARDWARE | HW Problem | 15-7 Plexiglass cover needs to be replaced also taped klaxon wires need to be reconnected.
123851 | 42091.5693 | 1 | DR13 | 1 | 4 | HARDWARE | HW Problem | Kly Mod 1-8, DR 13 Compressor: tripped with no reset.
122160 | 42019.5824 | 1.5 | LI16 | 41 | 4 | HARDWARE | HW Problem | 16-4: K26 needs to replace contactor L3. 16-4: T20 fan touching the side cover. 16-4: Primary capacitor network C58 and C60, the wires/jumpers connected between them are burnt/melt. 16-4: Primary cap network C59 has a loose screw. 16-4: Primary resistor network R260=10ohm needs to replace. 16-4: Lug burnt at Thyratron tube connections.
122535 | 42036.3345 | 1 | LI19 | 11 | 4 | HARDWARE | HW Problem | 19-1 tripping on Klystron Vacuum when ramping up
123111 | 42061.4002 | 1 | LI13 | 21 | 4 | HARDWARE | HW Problem | 13-2 tripping repeatedly.
123122 | 42061.7059 | 1 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3 MOD fault no reset
123480 | 42076.6197 | 0.5 | LI05 | 11 | 4 | HARDWARE | HW Problem | 5-1 MOD faulted, won't reset.
123480 | 42076.6197 | 3 | LI05 | 11 | 4 | HARDWARE | HW Problem | 5-1 MOD faulted, won't reset.
123480 | 42076.6197 | 5 | LI05 | 11 | 4 | HARDWARE | HW Problem | 5-1 MOD faulted, won't reset.
123500 | 42077.5696 | 0.5 | LI06 | 61 | 4 | HARDWARE | HW Problem | 6-6 WGV fault.
123540 | 42079.4533 | 0.75 | LI13 | 61 | 4 | HARDWARE | HW Problem | Klystron 13-6 is MOD faulted and doesn't come back with a remote reset.
123567 | 42080.2641 | 1 | LI26 | 21 | 4 | HARDWARE | HW Problem | 26-2: HVPS Regulator Board needs modification
121858 | 42010.3346 | 0.8 | LI30 | 41 | 4 | HARDWARE | HW Problem | 30-4 is tripping on thyratron filament O/C, resets and runs for a while then trips again.
121858 | 42010.3346 | 1 | LI30 | 41 | 4 | HARDWARE | HW Problem | 30-4 is tripping on thyratron filament O/C, resets and runs for a while then trips again.
122095 | 42017.2726 | 2 | LI18 | 31 | 4 | HARDWARE | HW Problem | 18-3 KMB box has 2 badly burnt terminal board connections
122310 | 42026.923 | 0.5 | LI02 | 81 | 4 | HARDWARE | HW Problem | 2-8 at KHD
122611 | 42039.1277 | 1 | LI05 | 51 | 4 | HARDWARE | HW Problem | 5-5 Needs PFN tuning
123265 | 42068.4674 | 0.75 | LI14 | 51 | 4 | HARDWARE | HW Problem | station 14-5 mod faulted -won't reset from scp
123638 | 42082.5409 | 1 | LI09 | 81 | 4 | HARDWARE | HW Problem | klystron 9-8 mod faulted and won't reset
123785 | 42089.3046 | 1 | LI21 | 21 | 4 | HARDWARE | HW Problem | Downtime: LINAC EAST 21-2 PM Klystron Solenoid PS - PM (4 of 5)
123785 | 42089.3046 | 12 | LI21 | 21 | 4 | HARDWARE | HW Problem | Downtime: LINAC EAST 21-2 PM Klystron Solenoid PS - PM (4 of 5)
122930 | 42053.7048 | 1 | LI24 | 51 | 4 | HARDWARE | HW Problem | Tried three times
123092 | 42060.7456 | 1.5 | LI26 | 41 | 4 | HARDWARE | HW Problem | 26-4 MOD fault, no remote reset
123338 | 42071.5624 | 1 | LI14 | 31 | 4 | HARDWARE | HW Problem | Klystron 14-3 TBR
123338 | 42071.5624 | 0.5 | LI14 | 31 | 4 | HARDWARE | HW Problem | Klystron 14-3 TBR
123921 | 42094.537 | 1 | LI04 | 11 | 4 | HARDWARE | HW Problem | Kly 4-1 amp dropped from 52 to 33 at around 5:00. Forward energy dropped as well. Modulator HV was fine. Found Drive set to 100%, then tried Saturation FTP to see if it was overdriven. It was slightly overdriven, but not enough to explain the >20% drop. Accepted peak drive, and Ampl returned to normal. Suspect there was something wrong with the attenuator with it set to 100%.
123937 | 42094.7774 | 1 | LI15 | 51 | 4 | HARDWARE | HW Problem | 15-5 - Wont reset. HV on, thyratron not firing.
123960 | 42095.451 | 0.8 | LI02 | 61 | 4 | HARDWARE | HW Problem | Klystron LI02-61 has very high phase and amplitude jitter.
122798 | 42047.4683 | 2 | LI25 | 31 | 4 | HARDWARE | HW Problem | 25-3: Modulator fault, no remote reset
123078 | 42060.4601 |  | LI20 | 51 | 4 | HARDWARE | HW Problem | TCav0 tripping when we try and bring it on.
122794 | 42047.437 | 0 | LI28 | 61 | 4 | HARDWARE | HW Problem | It goes to HVR, and resets every time, but doesn't stay on for more than an hour.
122802 | 42047.554 | 1 | LI05 | 61 | 4 | HARDWARE | HW Problem | 5-6: Seeing some beam volts jitter.
122835 | 42048.5905 | 0.75 | LI29 | 11 | 4 | HARDWARE | HW Problem | 29-1: BV low. Needs to adjust to spec.
122867 | 42050.5344 | 0.75 | LI21 | 81 | 4 | HARDWARE | HW Problem | 21-8: RF AMPL. is dropping out.:(see pic.s in MCC logbook)
123177 | 42064.3539 | 1 | LI30 | 51 | 4 | HARDWARE | HW Problem | Noticed Kly 30-3 beam volts and RF ampl. jitter is bad. Plots are attached to the elog entry: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=767677
123605 | 42081.44 | 2 | LI27 | 21 | 4 | HARDWARE | HW Problem | 27-2 trips MOD fault resets then trips on run up.
123621 | 42081.9004 | 0.75 | LI06 | 81 | 4 | HARDWARE | HW Problem | 6-8 MOD fault, won't reset.
123621 | 42081.9004 | 0.5 | LI06 | 81 | 4 | HARDWARE | HW Problem | 6-8 MOD fault, won't reset.
123625 | 42082.0709 | 1 | LI26 | 61 | 4 | HARDWARE | HW Problem | Station 26-6 BVLT has dropped down 90% or more 8 times in last 4 hours.
123625 | 42082.0709 | 0.25 | LI26 | 61 | 4 | HARDWARE | HW Problem | Station 26-6 BVLT has dropped down 90% or more 8 times in last 4 hours.
122108 | 42017.5812 | 1 | LI15 | 21 | 4 | HARDWARE | HW Problem | 15-2: T20 HV (R3) bushing leaking (sealed top).
122288 | 42026.3344 | 0.5 | LI18 | 41 | 4 | HARDWARE | HW Problem | 18-4 VGC high voltage is tripped and won't reset
122861 | 42049.5505 | 2 | LI30 | 41 | 4 | HARDWARE | HW Problem | Klys 30-4 just went KHD for no apparent reason. Filament fuse open.
122894 | 42052.4064 | 0.75 | LI26 | 51 | 4 | HARDWARE | HW Problem | 26-5 has large Beam Volts fluctuations, 1-2 kV single pulse drops. DeQuing current also shows large fluctuations compared to other LI26 stations (archiver data).
123132 | 42062.1446 | 1.5 | LI29 | 21 | 4 | HARDWARE | HW Problem | Also tripped 239 times on "AMPLT."
123132 | 42062.1446 | 0.75 | LI29 | 21 | 4 | HARDWARE | HW Problem | Also tripped 239 times on "AMPLT."
123254 | 42067.9915 | 1 | LI05 | 11 | 4 | HARDWARE | HW Problem | Station 5-1 has a cable fault that won't clear.
123349 | 42072.3913 |  | LI29 | 11 | 4 | HARDWARE | HW Problem | LI29-1, Tried to reset the PIOP but it still is doing it.
123419 | 42074.5712 | 0.1 | LI21 | 31 | 4 | HARDWARE | HW Problem | 2 volt faults 57 current faults 1 reflected fault 1 link fault 1 drive fault 1 temp. fault 777 amplt. fault.
123429 | 42074.8583 | 2 | LI17 | 51 | 4 | HARDWARE | HW Problem | Klystron 17-5 mod fault.
123460 | 42076.2091 | 0.5 | LI20 | 51 | 4 | HARDWARE | HW Problem | Klys 20-5 (T-Cav0) has MOD faulted and will not reset remotely.
122513 | 42034.6661 |  | LI26 | 41 | 4 | HARDWARE | HW Problem | 26-4 Mod fault, VVS breaker tripped.
122755 | 42046.3322 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 intermittent DQ under current, modification to HV regulator board may fix the problem
122836 | 42048.5912 | 0.75 | LI29 | 51 | 4 | HARDWARE | HW Problem | 29-5: BV low. Needs to adjust to spec.
122927 | 42053.6948 |  | LI27 | 21 | 4 | HARDWARE | HW Problem | (See pic. in MCC log)
122927 | 42053.6948 |  | LI27 | 21 | 4 | HARDWARE | HW Problem | (See pic. in MCC log)
122927 | 42053.6948 | 0.5 | LI27 | 21 | 4 | HARDWARE | HW Problem | (See pic. in MCC log)
123231 | 42067.3636 | 1.5 | LI21 | 81 | 4 | HARDWARE | HW Problem | 21-8 Modulator, R & R Thyratron.
123458 | 42076.0328 | 1 | LI10 | 41 | 4 | HARDWARE | HW Problem | Station 10-4 fault lockout, can't reset remotely.
123479 | 42076.5992 | 0.25 | LI26 | 71 | 4 | HARDWARE | HW Problem | Should get it turn back on.
123576 | 42080.5211 | 2 | LI07 | 11 | 4 | HARDWARE | HW Problem | 7-1 Modulator, Remove T20 and Install in 5-3
123576 | 42080.5211 | 3 | LI07 | 11 | 4 | HARDWARE | HW Problem | 7-1 Modulator, Remove T20 and Install in 5-3
123576 | 42080.5211 | 6 | LI07 | 11 | 4 | HARDWARE | HW Problem | 7-1 Modulator, Remove T20 and Install in 5-3
122064 | 42016.4548 | 1 | LI26 | 51 | 4 | HARDWARE | HW Problem | Modulator AIP- PLC shows DeQuing Under Current alarm on 'Interlocks' list, even though DeQuing analog readback of 1.5 Amps is NOT in alarm and station is running. This makes a red alarm on LCLS Home - LI26 RF. 1.5 Amps does sounds low, but not clear if alarm is false, or if this is just a warning that station needs tuning?
122064 | 42016.4548 | 1 | LI26 | 51 | 4 | HARDWARE | HW Problem | Modulator AIP- PLC shows DeQuing Under Current alarm on 'Interlocks' list, even though DeQuing analog readback of 1.5 Amps is NOT in alarm and station is running. This makes a red alarm on LCLS Home - LI26 RF. 1.5 Amps does sounds low, but not clear if alarm is false, or if this is just a warning that station needs tuning?
122125 | 42018.4492 | 1 | LI20 | 41 | 4 | HARDWARE | HW Problem | 20-4 Has a burnt PFN strap ( Lower, RHS. #3) Cables hanging loose over door #3 need to be secured
122126 | 42018.5318 | 1 | LI06 | 51 | 4 | HARDWARE | HW Problem | 6-5: K26 need to replace three contactors. K27 needs to replace contactor at L3. 6-5: Primary resistor network R261 installed wrong 20-ohm resistor value. It is 10-ohm. 6-5: Primary capacitor C60=25uF has a leak. 6-5: Loose Straps: upper (LHS #1) ; lower: (LHS #1,2,3). 6-5: Rack KF06-5B: Klystron Magnet Breaker have bad AC wires. The wire insulation is cracked.
122126 | 42018.5318 | 1 | LI06 | 51 | 4 | HARDWARE | HW Problem | 6-5: K26 need to replace three contactors. K27 needs to replace contactor at L3. 6-5: Primary resistor network R261 installed wrong 20-ohm resistor value. It is 10-ohm. 6-5: Primary capacitor C60=25uF has a leak. 6-5: Loose Straps: upper (LHS #1) ; lower: (LHS #1,2,3). 6-5: Rack KF06-5B: Klystron Magnet Breaker have bad AC wires. The wire insulation is cracked.
122139 | 42018.6458 | 1 | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X (LI21 KLYS 21) Klystron Diagnostic PAD Beam Volts channel has no signal. Went to zero at around 07:48 Sunday 1/11/15.
122139 | 42018.6458 | 0.5 | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X (LI21 KLYS 21) Klystron Diagnostic PAD Beam Volts channel has no signal. Went to zero at around 07:48 Sunday 1/11/15.
122496 | 42033.7113 | 0.5 | LI21 | 21 | 4 | HARDWARE | HW Problem | Calling PEM to have a look
123191 | 42065.5322 | 0.5 | LI02 | 61 | 4 | HARDWARE | HW Problem | Klystron 2-6 has tripped off and will not reset remotely
121997 | 42013.2865 | 1 | LI20 | 71 | 4 | HARDWARE | HW Problem | L0A (klys 20-7) dropping beam volts. At certain times every 10 seconds or so.
122163 | 42019.6008 | 1.5 | LI16 | 71 | 4 | HARDWARE | HW Problem | 16-7: T20: two resistors of HV divider stack need to replace. 16-7: K26 needs to replace the contactor L3. 16-7: K27 needs to replace the contactor L3. 16-7: T20 at R3 terminal has a oil leak and around the T20. 16-7: Charging diode stack has a broken cap, the third column from the right.
122831 | 42048.5505 |  | LI30 | 71 | 4 | HARDWARE | HW Problem | This is independent of the fiducial faults that have been happening (unless ghosts are causing both). Currently FLK with beam volts at 14.5.
123043 | 42059.5204 |  | LI26 | 51 | 4 | HARDWARE | HW Problem | 26-5: Check the WG2 flow switch. Went straight to WG1 when pump off. Al other stations had WG2.
123057 | 42059.7178 |  | LI21 | 31 | 4 | HARDWARE | HW Problem | This klystron becomes FLK and we are able to reset it, but it immediately trips off after a few seconds.
123057 | 42059.7178 | 0.8 | LI21 | 31 | 4 | HARDWARE | HW Problem | This klystron becomes FLK and we are able to reset it, but it immediately trips off after a few seconds.
122419 | 42031.6414 | 1 | LI19 | 41 | 4 | HARDWARE | HW Problem | Kly Mod 19-4: check T20 oil level. Add oil as necessary for T20 HV ops.
121920 | 42011.4517 | 0.1 | LI24 | 51 | 4 | HARDWARE | HW Problem | Drive request is at 88%, Drive FTP shows no signal. Klys Ampl FTP and Forward Energy are also 0. Edit by Zimmer: putting the station on/off the beam shows a change in the energy, so this station is doing something.
122086 | 42016.8632 | 0.5 | LI29 | 51 | 4 | HARDWARE | HW Problem | 29-5 WGV
122123 | 42018.446 | 1 | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7 T20 compartment cooling fan is dead
123173 | 42064.2425 | 2 | LI28 | 61 | 4 | HARDWARE | HW Problem | 28-6 faulting on over-voltage. Six times this shift.
122576 | 42037.7895 | 0.1 | LI27 | 41 | 4 | HARDWARE | HW Problem | 27-4 HVR does not reset
123653 | 42083.2889 | 3 | LI25 | 71 | 4 | HARDWARE | HW Problem | 25-7 repeating mod faults. Ramps it AMPL back up mod faults. Its on an endless cycle. Does not seem to enter the lockout state.
121658 | 41987.3485 | 0.5 | LI04 | 41 | 4 | HARDWARE | HW Problem | Klystron 4-4 MOD fault, won't reset
121734 | 41991.3821 | 0.25 | LI26 | 11 | 4 | HARDWARE | HW Problem | the klystron display shows that it's on and happy.
121903 | 42011.1125 |  | LI20 | 81 | 4 | HARDWARE | HW Problem | Klys 20-8 (L0B) AMPL seeing dropouts. Not seeing anything on beam volts.
122059 | 42016.4112 | 2 | LI06 | 41 | 4 | HARDWARE | HW Problem | 6-4: K27 L2=B is showing burnt marks. 6-4: T20 primary capacitor C60=25uF leak. 6-4: Loose strap: upper (RHS #2). 6-4: Kly Magnet breaker (KMB) has a burnt lug.
122117 | 42017.8986 | 1 | LI24 | 41 | 4 | HARDWARE | HW Problem | 24-4's modulator is tripping off pretty frequently. It shows 120 AMPL faults in the last 13 hours.
122117 | 42017.8986 | 1 | LI24 | 41 | 4 | HARDWARE | HW Problem | 24-4's modulator is tripping off pretty frequently. It shows 120 AMPL faults in the last 13 hours.
122154 | 42019.5507 | 2 | LI09 | 51 | 4 | HARDWARE | HW Problem | 9-5: K26 and K27 needs to replace the contactors. 9-5: HV rectifier stack is showing burnt mark. 9-5: Primary Capacitor network C60=25uF has a leak. 9-5: T20: AC input terminals: phase A, B, C are showing oil leak. 9-5: T20: shunt trip board wires need to secure by tie-wrap. 9-5: Loose straps: upper (Lhs #3); lower (Rhs #1).
122289 | 42026.3359 | 0.75 | LI18 | 61 | 4 | HARDWARE | HW Problem | 18-6 has a KMC fault that won't clear, focus supply is putting out 15 amps
122361 | 42030.2652 | 1 | LI15 | 31 | 4 | HARDWARE | HW Problem | 15-3 has a cable fault, swapping PIOP cables with 15-2 didn't clear the problem but using the PAD substitution box did clear the fault - appears to be a bad PAD. I did try cycling its power and IPLing the PIOP with no help.
122361 | 42030.2652 | 1 | LI15 | 31 | 4 | HARDWARE | HW Problem | 15-3 has a cable fault, swapping PIOP cables with 15-2 didn't clear the problem but using the PAD substitution box did clear the fault - appears to be a bad PAD. I did try cycling its power and IPLing the PIOP with no help.
122392 | 42031.2406 | 1 | LI10 | 31 | 4 | HARDWARE | HW Problem | 10-3 WGV flt O.P light will not clear
122810 | 42047.7374 | 1 | LI21 | 31 | 4 | HARDWARE | HW Problem | TBR - 21-3 FLK and not resetting
122810 | 42047.7374 |  | LI21 | 31 | 4 | HARDWARE | HW Problem | TBR - 21-3 FLK and not resetting
123312 | 42070.0969 | 1 | LI26 | 61 | 4 | HARDWARE | HW Problem | Station 26-6 has dropped BVLT down to zero 5 times in last 4 hours.
123319 | 42070.3444 | 0.8 | LI26 | 71 | 4 | HARDWARE | HW Problem | Station 26-7 BVLT has dropped down 90% or more 6 times in last 24 hours. CB7 reported tripping.
123320 | 42070.3529 | 0.75 | LI27 | 61 | 4 | HARDWARE | HW Problem | Station 27-6 BVLT has dropped down 33% or more 12 times in last 24 hours.
123445 | 42075.5834 | 0.75 | LI11 | 51 | 4 | HARDWARE | HW Problem | Tried to reset but it didn't clear.
123506 | 42077.7381 | 0.75 | LI15 | 71 | 4 | HARDWARE | HW Problem | 15-7 MOD fault, won't reset.
121928 | 42011.572 | 1 | DR01 | 1 | 4 | HARDWARE | HW Problem | 1-7 KMB test point resistor is burnt/broken (1/4watt 10K )
122188 | 42020.6047 |  | LI28 | 61 | 4 | HARDWARE | HW Problem | Klys 28-6 is tripping off frequently. 286 times since 08:00 by 14:30 same day.
123684 | 42083.9182 | 0.75 | LI20 | 71 | 4 | HARDWARE | HW Problem | LO-A mod faulted and now reads TOC-Thryatron over current. This cannot be reset remotely.
123715 | 42086.3522 | 0.3 | LI13 | 61 | 4 | HARDWARE | HW Problem | 13-6 is tripped on mod fault. Can't reset remotely.
123715 | 42086.3522 | 0.3 | LI13 | 61 | 4 | HARDWARE | HW Problem | 13-6 is tripped on mod fault. Can't reset remotely.
121927 | 42011.5691 | 1 | LI01 | 31 | 4 | HARDWARE | HW Problem | 1-3 large number of heat sinks on stack are loose and it needs to be replaced.
121934 | 42011.6003 | 1 | LI22 | 21 | 4 | HARDWARE | HW Problem | 22-2 Thyratron compartment cooling fan is dead
122016 | 42013.4943 |  | LI02 | 21 | 4 | HARDWARE | HW Problem | 2-2 K17 Relay is broken K27 contacts need to be replaced
122173 | 42020.3076 | 2 | LI16 | 81 | 4 | HARDWARE | HW Problem | 16-8: K26 needs to replace the contactor L3. 16-8: K27 needs to replace the contactor L2. 16-8: Primary capacitor network C58 and C60, the wires/jumper insulation are burnt/melt. 16-8: All AC input wires of T20 loose. 16-8: T20 at R1 and R2's terminal connections have oil leak. 16-8: Charging diode stack has a broken cap of third column. 16-8: Loose strap at upper (Rhs #3)
122197 | 42022.1666 | 1 | LI29 | 11 | 4 | HARDWARE | HW Problem | Klys 29-1 has had 13 over voltage errors in the last 20 hours. It will stay on 10 - 60 minutes before tripping again.
122206 | 42023.0539 | 1 | LI22 | 61 | 4 | HARDWARE | HW Problem | Runs 0.10 then jumps up to 0.50 Should be checked out. Possible thermal problem?
122206 | 42023.0539 | 1 | LI22 | 61 | 4 | HARDWARE | HW Problem | Runs 0.10 then jumps up to 0.50 Should be checked out. Possible thermal problem?
122206 | 42023.0539 | 0.75 | LI22 | 61 | 4 | HARDWARE | HW Problem | Runs 0.10 then jumps up to 0.50 Should be checked out. Possible thermal problem?
122767 | 42046.593 | 1 | LI00 | 11 | 4 | HARDWARE | HW Problem | 01-1A Has a dead thyratron compartment cooling fan
123619 | 42081.7713 |  | LI20 | 41 | 4 | HARDWARE | HW Problem | 20-4B (FXTCAV): Information and warning regarding lowering beam volts.
121668 | 41988.2277 | 0.75 | LI17 | 21 | 4 | HARDWARE | HW Problem | 17-2 waveguide vac faulted
121925 | 42011.5657 | 0.5 | LI26 | 71 | 4 | HARDWARE | HW Problem | Klystron 26-7 VVS mod fault.
121952 | 42012.3058 | 1.5 | LI30 | 71 | 4 | HARDWARE | HW Problem | 30-7 Thy comp de-spiking coil has a broken resistor
121964 | 42012.4548 | 1.5 | DR01 | 1 | 4 | HARDWARE | HW Problem | 1-7B showing CBL fault. Fault clears with PAD sub box, PAD +5 LED out.
122275 | 42025.5319 |  | LI26 |  | 4 | HARDWARE | HW Problem | 26-2 and 26-3 started taking turns tripping off this morning
122277 | 42025.5363 | 1 | LI29 | 41 | 4 | HARDWARE | HW Problem | 29-4: Seeing some beam volts jitter.
123174 | 42064.2462 | 1 | LI28 | 81 | 4 | HARDWARE | HW Problem | 28-8 MOD faulted several times. Reset remotely each time, but eventually became too frequent to tolerate and it was taken off the beam.
123218 | 42066.5836 | 2.5 | LI28 | 81 | 4 | HARDWARE | HW Problem | 28-8 have narrow Width = 2.91uS and Slope = -0.2
123782 | 42089.264 | 0.01 | LI04 | 51 | 4 | HARDWARE | HW Problem | 4-5 phase and amplitude readback mean nothing
123784 | 42089.2974 |  | LI21 | 61 | 4 | HARDWARE | HW Problem | 21-6: high phase jitter, no klystron beam voltage jitter
121733 | 41991.3534 | 1 | LI28 | 11 | 4 | HARDWARE | HW Problem | Voltage come back after a couple of seconds, then repeats cycle.
122240 | 42024.5053 | 0.5 | LI20 | 71 | 4 | HARDWARE | HW Problem | Klystron 20-7 (L0A) is MOD faulted and won't reset remotely. This is causing downtime for the MD program.
123158 | 42062.9171 |  | LI24 | 31 | 4 | HARDWARE | HW Problem | Would expect triggering PAC at standby 50uS after accelerate time would not effect Accelerate RF. See MCC ELOG entry https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=767524 Please schedule POMM job to look at this situation locally with scope to understand what is going on when we get in this condition.
123158 | 42062.9171 | 1.5 | LI24 | 31 | 4 | HARDWARE | HW Problem | Would expect triggering PAC at standby 50uS after accelerate time would not effect Accelerate RF. See MCC ELOG entry https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=767524 Please schedule POMM job to look at this situation locally with scope to understand what is going on when we get in this condition.
123408 | 42074.476 | 1 | LI17 | 21 | 4 | HARDWARE | HW Problem | 17-2 the F6 fuse holder is broken needs to be replaced PFN strap Needs to be replaced (top, LHS # 3 ) PFN strap very loose may have to be replaced (Lower, LHS, # 3)
123107 | 42061.3511 | 0.2 | LI26 | 41 | 4 | HARDWARE | HW Problem | 26-4 CAM fault
123979 | 42095.5884 | 0.5 | LI24 | 81 | 4 | HARDWARE | HW Problem | 24-8 Modulator: SCR chassis needs to replace the snubber board
123980 | 42095.6063 | 1 | LI27 | 11 | 4 | HARDWARE | HW Problem | Beam volts jitter on 27-1 is very high, station is not usable in this condition. Update: Ops put in MAINT
123980 | 42095.6063 | 2.5 | LI27 | 11 | 4 | HARDWARE | HW Problem | Beam volts jitter on 27-1 is very high, station is not usable in this condition. Update: Ops put in MAINT
123980 | 42095.6063 | 1 | LI27 | 11 | 4 | HARDWARE | HW Problem | Beam volts jitter on 27-1 is very high, station is not usable in this condition. Update: Ops put in MAINT
123988 | 42095.8211 | 2 | LI25 | 81 | 4 | HARDWARE | HW Problem | Klys 25-8 is tripping repeatedly.
121854 | 42010.3138 | 3 | LI26 | 61 | 4 | HARDWARE | HW Problem | 26-6 PFN is unstable and is tripping on HVOC
122056 | 42016.3862 | 1 | LI27 | 41 | 4 | HARDWARE | HW Problem | KLYS:LI27:41:DL_WG_TEMP has been climbing slowly, is now out of tolerance at 114 F.
122153 | 42019.5452 | 1 | LI09 | 41 | 4 | HARDWARE | HW Problem | 9-4: K26 needs to replace L3. 9-4: T20 have oil leak at R1 and R2 terminal connections. 9-4: Thyratron compartment: two capacitors near DeQing box#2 have oil leak.
122178 | 42020.4276 | 0.5 | LI30 | 51 | 4 | HARDWARE | HW Problem | 30-5 is cycling incessantly
122201 | 42022.5344 | 1.5 | LI25 | 11 | 4 | HARDWARE | HW Problem | Klystron 25-1 cycling through faults (TRG, CBL, MKT, MOD); tried to reset remotely a few times but it keeps faulting.
122201 | 42022.5344 | 0.5 | LI25 | 11 | 4 | HARDWARE | HW Problem | Klystron 25-1 cycling through faults (TRG, CBL, MKT, MOD); tried to reset remotely a few times but it keeps faulting.
123811 | 42089.6785 | 1 | LI20 | 41 | 4 | HARDWARE | HW Problem | station 20-4b (aka XTCAV) mod faulted
123812 | 42089.726 | 1 | LI20 | 61 | 4 | HARDWARE | HW Problem | Temp. is going from 64.0 deg.s to 75.5 deg.F
122190 | 42020.6544 | 1 | LI22 | 81 | 4 | HARDWARE | HW Problem | 22-8 cannot be reset remotely
123740 | 42087.5915 |  | LI02 | 11 | 4 | HARDWARE | HW Problem | 2-1 K26 Contacts need to be replaced
123755 | 42088.5451 | 0.75 | LI19 | 31 | 4 | HARDWARE | HW Problem | 19-3 Core bias Variac needs to be replaced.
123757 | 42088.5485 | 0.5 | LI18 | 21 | 4 | HARDWARE | HW Problem | 18-2 Mod core bias variac needs to be replaced 1 pair of K26 contacts need to be replaced
123758 | 42088.5501 | 0.75 | LI17 | 11 | 4 | HARDWARE | HW Problem | 17-1 Mod fire alarm latching box system ok LED is burnt out
123764 | 42088.6228 | 0.5 | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 L1X filament timeout Need to start the filamnents
122996 | 42057.6433 |  | LI11 | 21 | 4 | HARDWARE | HW Problem | 11-2: Mod faulted due to VVS breaker trip.
123600 | 42081.221 | 0.5 | LI09 | 51 | 4 | HARDWARE | HW Problem | Klystron 9-5 mod fault.
123989 | 42095.9279 | 0.75 | LI13 | 11 | 4 | HARDWARE | HW Problem | Won't reset.
121966 | 42012.4563 | 1 | LI14 | 81 | 4 | HARDWARE | HW Problem | 14-8 showing CBL fault
122152 | 42019.5436 | 1 | LI28 | 11 | 4 | HARDWARE | HW Problem | 28-1: False "Electro-Magnet Breaker" Fault. Trim PS AMP connector has a pushed in pin.
122195 | 42020.9256 | 0.1 | LI21 | 81 | 4 | HARDWARE | HW Problem | 21-8 mod faulted and won't reset remotely
122286 | 42026.3114 | 1 | LI04 | 21 | 4 | HARDWARE | HW Problem | 4-2 PFN Cap needs replace LHS 2nd from the buttom. No core bias reading at meter veriac looks burnt.
122425 | 42032.1885 | 2 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 Mod faulted on KHD.
122501 | 42034.0146 | 0.5 | LI30 | 31 | 4 | HARDWARE | HW Problem | Dropping out by 3-4 kV. Thyratron probably needs to be ranged.
122501 | 42034.0146 | 0.5 | LI30 | 31 | 4 | HARDWARE | HW Problem | Dropping out by 3-4 kV. Thyratron probably needs to be ranged.
122501 | 42034.0146 | 0.5 | LI30 | 31 | 4 | HARDWARE | HW Problem | Dropping out by 3-4 kV. Thyratron probably needs to be ranged.
122501 | 42034.0146 | 0.5 | LI30 | 31 | 4 | HARDWARE | HW Problem | Dropping out by 3-4 kV. Thyratron probably needs to be ranged.
122652 | 42040.4803 | 2.5 | LI03 | 31 | 4 | HARDWARE | HW Problem | 3-3: Seeing lots of beam volts jitter.
123883 | 42093.5929 | 2 | LI27 | 11 | 4 | HARDWARE | HW Problem | 27-1: Needs to replace 1. 3 burnt snubber boards. 2. missing 4 AC filters. 3. R3 on Enerpro board has burnt mark.
123903 | 42094.2319 | 1 | LI21 | 61 | 4 | HARDWARE | HW Problem | 21-6 very jittery
123917 | 42094.4567 | 2 | LI20 | 51 | 4 | HARDWARE | HW Problem | 20-5 (TCAV0): HV CBT. Lowered Vres by 0.2 and tried to run. Tripped again. TS&R.
123920 | 42094.4887 | 0.2 | LI02 | 71 | 4 | HARDWARE | HW Problem | 2-7: RF AMPL huge diurnal swing taking it in and out of TOLS.
123926 | 42094.5815 | 1 | LI30 | 61 | 4 | HARDWARE | HW Problem | Klystron LI30-61 has accumulated 38 water faults today.
123953 | 42095.415 |  | LI23 | 21 | 4 | HARDWARE | HW Problem | Klystron 23-2 mod fault. Frequent mod faults.
122341 | 42028.4823 | 0.5 | LI18 | 51 | 4 | HARDWARE | HW Problem | Kly Mod 18-5: check T20 oil level. Add oil as necessary.
122500 | 42033.9707 |  | LI26 | 41 | 4 | HARDWARE | HW Problem | 26-4 has VVS breaker tripped.
122545 | 42037.3597 | 3 | LI23 | 61 | 4 | HARDWARE | HW Problem | 23-6: Needs PFN tune to get BV in spec.
123787 | 42089.3077 | 8 | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2 install L3 50 kV thyratron for tests with DC power supplies used for filament and reservoir current sources.
123787 | 42089.3077 | 6 | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2 install L3 50 kV thyratron for tests with DC power supplies used for filament and reservoir current sources.
123787 | 42089.3077 | 8 | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2 install L3 50 kV thyratron for tests with DC power supplies used for filament and reservoir current sources.
123787 | 42089.3077 | 8 | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2 install L3 50 kV thyratron for tests with DC power supplies used for filament and reservoir current sources.
123798 | 42089.4544 | 1.5 | LI24 | 81 | 4 | HARDWARE | HW Problem | The Touch Panel of Modulator 24-8 does not connect to Ethernet network
123803 | 42089.5129 | 1 | LI29 | 81 | 4 | HARDWARE | HW Problem | Large phase jitter on station 29-8; could clearly see a correlation in the FEL performance. Took off the beam and FEL jitter is much better.
123345 | 42072.3643 | 1 | LI23 | 21 | 4 | HARDWARE | HW Problem | 23-2 beam voltage jitter 1 sigma = 0.2 kV
123362 | 42072.6339 | 0.75 | LI05 | 71 | 4 | HARDWARE | HW Problem | 5-7: MOD fault, won't remote reset.
123680 | 42083.6161 | 0.5 | LI26 | 71 | 4 | HARDWARE | HW Problem | 26-7 MOD VVS voltage fault
124037 | 42097.4452 | 0.2 | LI22 | 81 | 4 | HARDWARE | HW Problem | 22-8: Beam volts low at 338 KV (345 is spec).
121970 | 42012.4758 | 1 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 EOLC stack needs to be replaced 1 Carborundum resistor is burnt/broken
122054 | 42016.3046 | 1 | LI14 | 81 | 4 | HARDWARE | HW Problem | 14-8: Needs NIRP test when the FACET turn back on.
123633 | 42082.4941 | 0.5 | LI10 | 41 | 4 | HARDWARE | HW Problem | Klystron 10-4 MOD faulted and won't reset remotely
123985 | 42095.6521 |  | LI13 | 21 | 4 | HARDWARE | HW Problem | 13-2 151 over-voltage faults, 19 thyratron over-current faults.
124026 | 42097.2316 |  | LI17 | 61 | 4 | HARDWARE | HW Problem | 17-6 Thyratron will not fire. Raised the Eres in tenths increments and when it reached 5.0 it fires thru and tripped CB7. This happened 2 times.
124051 | 42097.6951 |  | LI23 | 21 | 4 | HARDWARE | HW Problem | (see history in MCC log.)
124055 | 42097.9224 |  | LI23 | 11 | 4 | HARDWARE | HW Problem | Klys 23-1 stopped getting triggers suddenly. Deact and React brought them back, but it started losing triggers again shortly after.
124074 | 42099.6181 | 4 | LI22 | 11 | 4 | HARDWARE | HW Problem | Klystron 22-1 (KLYS:LI22:11) cycling violently on MOD and over current faults. TBR due to HV arcing in klystron tank, no klystron vacuum activity.
121947 | 42011.8014 | 1 | LI28 | 61 | 4 | HARDWARE | HW Problem | 28-6 fault on HVOC. Ranged thyratron from 6.0 to 8.1. Worked for a short time, faulting again. Needs new thyratron.
122336 | 42028.3085 | 0.5 | LI17 | 61 | 4 | HARDWARE | HW Problem | 17-6 At yellow light no green when key on
123760 | 42088.5806 | 1 | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2: Modulator, modify Enerpro Board
123477 | 42076.5148 | 1 | LI29 | 41 | 4 | HARDWARE | HW Problem | 29-4: Seeing some beam volts jitter
123729 | 42087.3259 | 1 | LI00 | 12 | 4 | HARDWARE | HW Problem | Downtime, FACET;00-1B, cold tune PFN.
124045 | 42097.5543 |  | LI25 | 71 | 4 | HARDWARE | HW Problem | FTP amp jitter is large. FTP beam volts is oscillating at the peak. Unable to put it on the beam at this time. From another CATER(124062(TEC)) Re-energized 25-7 after heater time ins. Station on and stable. Ramped slowly up to 22.2 KV. It appears the DQing firing is moving a little back and forth in time. Raised DQ current from 21.5 to 22.5 amps. This lowered BV by 3 K, but still in tols. Gave station back to MCC. Hope the regulation issues are solved. Passed on issues to swings for monitoring. It appears lowering HV set point by even small amounts stops the motorboating" and wild BV swings we are seeing.
124048 | 42097.6386 | 1 | LI28 | 31 | 4 | HARDWARE | HW Problem | I was having trouble phasing 28-3 - looked like a good scan, but would come back with some crazy result when I would re-scan. Something was definitely being moved relative to the beam (looking at energy BPMs). Tried doing a hand-phasing: no luck with that either. Tried homing the FOX phase shifter...it failed with a BAD FOX" status on the EDM panel, all but the first and last times. Went from 306 -> 274 -> 306 with those scans.
124072 | 42099.4656 | 0.5 | LI15 | 71 | 4 | HARDWARE | HW Problem | Klystron LI15-71 is tripped off and will not reset remotely.
124072 | 42099.4656 | 1.5 | LI15 | 71 | 4 | HARDWARE | HW Problem | Klystron LI15-71 is tripped off and will not reset remotely.
124077 | 42100.1687 |  | LI13 | 51 | 4 | HARDWARE | HW Problem | 13-5 vacuum fault does not reset. KLV fault code.
124079 | 42100.3453 |  | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 klystron beam current readback drops 300 Amps every morning.
122350 | 42028.6691 | 0.5 | LI26 | 71 | 4 | HARDWARE | HW Problem | 26-7 VVS fault
122691 | 42041.7303 | 1.5 | LI09 | 81 | 4 | HARDWARE | HW Problem | 9-8 MOD faulting repeatedly.
122699 | 42042.0654 | 1 | LI02 | 61 | 4 | HARDWARE | HW Problem | Klystron LI02-61 is tripped off on Fault Lockout and will not reset remotely.
122748 | 42045.505 | 1 | LI29 | 81 | 4 | HARDWARE | HW Problem | 29-8: Seeing some beam volts jitter.
122830 | 42048.5458 | 1 | LI30 | 81 | 4 | HARDWARE | HW Problem | 30-8 has terrible beam volts jitter.
122830 | 42048.5458 | 1 | LI30 | 81 | 4 | HARDWARE | HW Problem | 30-8 has terrible beam volts jitter.
122637 | 42040.0352 | 0.5 | LI05 | 81 | 4 | HARDWARE | HW Problem | Klystron 5-8 mod fault.
121994 | 42013.182 | 0.5 | LI25 | 31 | 4 | HARDWARE | HW Problem | 25-3 VVS Brk tripped but not CB7. Dropped Eres from 3.95 to 3.73
122351 | 42028.7319 | 0.75 | LI22 | 81 | 4 | HARDWARE | HW Problem | Klystron 22-8 mod fault. Unable to reset.
123810 | 42089.6526 | 0.75 | LI17 | 21 | 4 | HARDWARE | HW Problem | Klys 17-2 WGV, can't be reset
122869 | 42050.5714 | 0.5 | LI14 | 51 | 4 | HARDWARE | HW Problem | 14-5: Electromagnet current fault, modulator still on.
123128 | 42062.1214 | 1 | LI17 | 61 | 4 | HARDWARE | HW Problem | 17-6 WGV flt
123650 | 42083.15 | 1.5 | LI27 | 21 | 4 | HARDWARE | HW Problem | Klystron 27-2 mod fault
123650 | 42083.15 | 3 | LI27 | 21 | 4 | HARDWARE | HW Problem | Klystron 27-2 mod fault
124094 | 42100.6073 |  | LI15 | 81 | 4 | HARDWARE | HW Problem | 15-8 The history of the focus I show a sudden drop of 3A around 05:00 on 4/6/15. It dropped from a stable 15.5A to a very unstable 12.5V. Beam current, FEPK, focus desired, and delta temp all remain stable suggesting that this is a read back problem.
124099 | 42100.7433 | 4 | LI02 | 81 | 4 | HARDWARE | HW Problem | 2-8: This klystron mod faults and then when it is reset, immediately trips again.
124107 | 42101.3744 |  | LI29 | 21 | 4 | HARDWARE | HW Problem | 29-2: HV CBT (CB7). Did not reset.
121877 | 42010.5543 | 0.75 | LI26 | 31 | 4 | HARDWARE | HW Problem | 26-3 AMM low beam voltage
121890 | 42010.6144 | 2 | LI26 | 41 | 4 | HARDWARE | HW Problem | 26-4 Need to adjust beam volts to spec( If we go higher, DQing will go down)
121890 | 42010.6144 |  | LI26 | 41 | 4 | HARDWARE | HW Problem | 26-4 Need to adjust beam volts to spec( If we go higher, DQing will go down)
121990 | 42013.177 |  | LI27 | 41 | 4 | HARDWARE | HW Problem | 27-4 Tripping on intermittent WGV. Dropped BV a bit then started tripping on OV. Kept dropping BV to 253KV were there was no further drops
122043 | 42015.2863 | 1.5 | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X has had multiple klystron vacuum faults, no extended pump downs
122251 | 42024.8557 |  | LI21 | 31 | 4 | HARDWARE | HW Problem | 21-3: Left at 310 KV after AIP Upgrade, was running at 350 KV. TS&R.
122251 | 42024.8557 | 3.5 | LI21 | 31 | 4 | HARDWARE | HW Problem | 21-3: Left at 310 KV after AIP Upgrade, was running at 350 KV. TS&R.
122251 | 42024.8557 | 2.5 | LI21 | 31 | 4 | HARDWARE | HW Problem | 21-3: Left at 310 KV after AIP Upgrade, was running at 350 KV. TS&R.
122251 | 42024.8557 | 2 | LI21 | 31 | 4 | HARDWARE | HW Problem | 21-3: Left at 310 KV after AIP Upgrade, was running at 350 KV. TS&R.
122251 | 42024.8557 | 2 | LI21 | 31 | 4 | HARDWARE | HW Problem | 21-3: Left at 310 KV after AIP Upgrade, was running at 350 KV. TS&R.
122251 | 42024.8557 | 2.5 | LI21 | 31 | 4 | HARDWARE | HW Problem | 21-3: Left at 310 KV after AIP Upgrade, was running at 350 KV. TS&R.
122251 | 42024.8557 | 1.5 | LI21 | 31 | 4 | HARDWARE | HW Problem | 21-3: Left at 310 KV after AIP Upgrade, was running at 350 KV. TS&R.
122417 | 42031.639 | 1 | LI19 | 11 | 4 | HARDWARE | HW Problem | Kly Mod 19-1: check T20 oil level. Add oil as necessary to make T20 safe for HV ops.
122845 | 42048.6862 | 1 | LI21 | 81 | 4 | HARDWARE | HW Problem | Beam voltage is 0.5 % sigma also.
122859 | 42049.4424 | 1 | DR13 | 1 | 4 | HARDWARE | HW Problem | Kly Mod 1-8, NRTL DR 13 Compressor: Mod faulted, no reset.
122884 | 42051.5292 | 2 | LI12 | 81 | 4 | HARDWARE | HW Problem | Klystron 12-8 mod fault. MKSU Trigger enable fault.
122886 | 42051.6367 | 1.5 | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X It comes back with a reset each time, but it's worrisome that it keeps dying.
122886 | 42051.6367 | 2 | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X It comes back with a reset each time, but it's worrisome that it keeps dying.
122886 | 42051.6367 | 1 | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X It comes back with a reset each time, but it's worrisome that it keeps dying.
122463 | 42033.287 |  | LI28 | 11 | 4 | HARDWARE | HW Problem | LI28 LCW system was raised by .5 deg F. Please check rotometer flow.
122577 | 42037.8155 | 0.5 | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 faulting often.
123142 | 42062.4847 | 2.5 | LI26 | 81 | 4 | HARDWARE | HW Problem | 26-8: TBR - EOLC/MKOC to FLK very quickly. Klystron TBR.
124126 | 42101.5541 |  | LI17 | 11 | 4 | HARDWARE | HW Problem | 17-1 Mod Fire alarm latch System ok LED is burnt out.
124127 | 42101.5562 | 0.75 | LI15 | 11 | 4 | HARDWARE | HW Problem | 15-1 Mod Fire alarm latch System ok LED is burnt out.
124128 | 42101.5573 | 0.75 | LI14 | 41 | 4 | HARDWARE | HW Problem | 14-4 Mod Fire alarm latch System ok LED is burnt out.
124130 | 42101.5592 | 0.75 | LI03 | 71 | 4 | HARDWARE | HW Problem | 3-7 Mod Fire alarm latch System ok LED is burnt out.
124132 | 42101.5619 | 0.25 | DR01 | 1 | 4 | HARDWARE | HW Problem | 1-7 B Kly trim P/S Panel Meter is inop
124137 | 42101.5669 | 0.75 | LI02 | 71 | 4 | HARDWARE | HW Problem | 2-7 Mod has a pair of K26 contacts that need to be replaced-
122116 | 42017.8757 | 1 | LI30 | 11 | 4 | HARDWARE | HW Problem | 30-1 has high beam volt jitter.
122168 | 42020.1834 | 1 | LI25 | 51 | 4 | HARDWARE | HW Problem | 25-5 MOD fault
122168 | 42020.1834 | 0.5 | LI25 | 51 | 4 | HARDWARE | HW Problem | 25-5 MOD fault
122249 | 42024.6189 | 1 | LI20 | 51 | 4 | HARDWARE | HW Problem | station 20-5 mod faulted-won't reset remotely
122359 | 42030.1236 | 0.8 | LI25 | 41 | 4 | HARDWARE | HW Problem | 25-4 has a lot of beamvolts jitter. The thyratron probably needs to be ranged.
122359 | 42030.1236 | 0.25 | LI25 | 41 | 4 | HARDWARE | HW Problem | 25-4 has a lot of beamvolts jitter. The thyratron probably needs to be ranged.
122359 | 42030.1236 | 2 | LI25 | 41 | 4 | HARDWARE | HW Problem | 25-4 has a lot of beamvolts jitter. The thyratron probably needs to be ranged.
122359 | 42030.1236 | 0.8 | LI25 | 41 | 4 | HARDWARE | HW Problem | 25-4 has a lot of beamvolts jitter. The thyratron probably needs to be ranged.
123028 | 42058.857 | 5 | LI03 | 51 | 4 | HARDWARE | HW Problem | Robert from security called about smoke in sec 3. Had MCC do a check of any equipment in area acting weird. Nothing but mod 3-5 which tripped at the same time the Vesda picked up the smoke. CB7 was tripped. Opened for inspection. Everything was cool to the touch, no oil anywhere, T20 coils cool, core still warm, PFN caps all cool. Fireman's thermal scanner showed nothing (cool, gotta et me one of those). Nothing burned, nothing out of place. ????? Met fire marshal Curran there. We all decided that whatever happened was over. Leaving this one for dayshift when we have 2 people. LOTO'd CB7 and 120 wall breaker. MT
123028 | 42058.857 | 1 | LI03 | 51 | 4 | HARDWARE | HW Problem | Robert from security called about smoke in sec 3. Had MCC do a check of any equipment in area acting weird. Nothing but mod 3-5 which tripped at the same time the Vesda picked up the smoke. CB7 was tripped. Opened for inspection. Everything was cool to the touch, no oil anywhere, T20 coils cool, core still warm, PFN caps all cool. Fireman's thermal scanner showed nothing (cool, gotta et me one of those). Nothing burned, nothing out of place. ????? Met fire marshal Curran there. We all decided that whatever happened was over. Leaving this one for dayshift when we have 2 people. LOTO'd CB7 and 120 wall breaker. MT
123028 | 42058.857 | 2 | LI03 | 51 | 4 | HARDWARE | HW Problem | Robert from security called about smoke in sec 3. Had MCC do a check of any equipment in area acting weird. Nothing but mod 3-5 which tripped at the same time the Vesda picked up the smoke. CB7 was tripped. Opened for inspection. Everything was cool to the touch, no oil anywhere, T20 coils cool, core still warm, PFN caps all cool. Fireman's thermal scanner showed nothing (cool, gotta et me one of those). Nothing burned, nothing out of place. ????? Met fire marshal Curran there. We all decided that whatever happened was over. Leaving this one for dayshift when we have 2 people. LOTO'd CB7 and 120 wall breaker. MT
123028 | 42058.857 | 2 | LI03 | 51 | 4 | HARDWARE | HW Problem | Robert from security called about smoke in sec 3. Had MCC do a check of any equipment in area acting weird. Nothing but mod 3-5 which tripped at the same time the Vesda picked up the smoke. CB7 was tripped. Opened for inspection. Everything was cool to the touch, no oil anywhere, T20 coils cool, core still warm, PFN caps all cool. Fireman's thermal scanner showed nothing (cool, gotta et me one of those). Nothing burned, nothing out of place. ????? Met fire marshal Curran there. We all decided that whatever happened was over. Leaving this one for dayshift when we have 2 people. LOTO'd CB7 and 120 wall breaker. MT
123028 | 42058.857 | 2 | LI03 | 51 | 4 | HARDWARE | HW Problem | Robert from security called about smoke in sec 3. Had MCC do a check of any equipment in area acting weird. Nothing but mod 3-5 which tripped at the same time the Vesda picked up the smoke. CB7 was tripped. Opened for inspection. Everything was cool to the touch, no oil anywhere, T20 coils cool, core still warm, PFN caps all cool. Fireman's thermal scanner showed nothing (cool, gotta et me one of those). Nothing burned, nothing out of place. ????? Met fire marshal Curran there. We all decided that whatever happened was over. Leaving this one for dayshift when we have 2 people. LOTO'd CB7 and 120 wall breaker. MT
123273 | 42068.5592 | 1 | LI28 | 61 | 4 | HARDWARE | HW Problem | HVR faults
123342 | 42072.1186 | 1 | LI17 | 61 | 4 | HARDWARE | HW Problem | Kly 17-6 MOD fault won't reset
123342 | 42072.1186 | 1 | LI17 | 61 | 4 | HARDWARE | HW Problem | Kly 17-6 MOD fault won't reset
121698 | 41989.4626 |  | LI12 | 21 | 4 | HARDWARE | HW Problem | 12-2: PAD phase mode in "Track" intermittently.
121719 | 41990.3845 | 0.5 | LI09 | 51 | 4 | HARDWARE | HW Problem | Won't reset.
122010 | 42013.4459 | 0.5 | LI24 | 11 | 4 | HARDWARE | HW Problem | Klys 24-1 has no RF drive. Didn't recover after PAMM.
122053 | 42016.2711 | 0.25 | LI22 | 31 | 4 | HARDWARE | HW Problem | 22-3 is WGV faulted
122298 | 42026.5464 | 1 | LI06 | 11 | 4 | HARDWARE | HW Problem | 6-1 SCRTD has an internal fault flashing and needs to be replaced.
122347 | 42028.6499 |  | LI19 | 51 | 4 | HARDWARE | HW Problem | Kly Mod 19-5: check T20 oil level. Add oil as necessary.
122348 | 42028.651 | 1 | LI20 | 31 | 4 | HARDWARE | HW Problem | Kly Mod 20-3: check T20 oil level. Add oil as necessary.
122374 | 42030.5204 | 0.75 | LI30 | 31 | 4 | HARDWARE | HW Problem | 30-3: Seeing some beam volts jitter (Spikes)
122382 | 42030.8187 | 0.7 | LI05 | 11 | 4 | HARDWARE | HW Problem | Remove PAD, replace fan and re-install PAD.
122581 | 42037.8395 | 0.1 | LI02 | 21 | 4 | HARDWARE | HW Problem | 2-2 MOD fault, won't reset.
122453 | 42032.8126 | 3 | LI26 | 61 | 4 | HARDWARE | HW Problem | Klystron 26-6 mod fault.
122673 | 42040.7346 | 0.75 | LI17 | 61 | 4 | HARDWARE | HW Problem | 17-6 MOD fault won't reset remotely
122858 | 42049.4348 | 0.5 | LI12 | 31 | 4 | HARDWARE | HW Problem | They always clear but we have seen CBL, RE, TOC ADC etc. errors all for this klystron. We have reset this Klystron remotely but the flashes of statuses remain. Ops swapped PIOP on 2/14, now showing MK trigger fault.
123069 | 42060.3089 | 1 | LI21 | 11 | 4 | HARDWARE | HW Problem | 21-1 T20 Input snubber resistors show heat damage. The wire Between the Mod filter caps and HV rectifier shows heat damage, insulation is brittle/broken, from the snubber resistors.
123337 | 42071.5129 | 0.5 | LI15 | 21 | 4 | HARDWARE | HW Problem | Klystron 15-2 mod fault.
121849 | 42010.2688 | 1 | LI26 | 81 | 4 | HARDWARE | HW Problem | Klys 26-8 showing beam volts dropouts
122385 | 42030.8238 | 1 | LI10 | 51 | 4 | HARDWARE | HW Problem | Replace PAD and Run Poly when there is RF
122685 | 42041.5102 | 0.75 | LI30 | 41 | 4 | HARDWARE | HW Problem | 30-4: RF AMPL low OOT due to BV low.
122812 | 42048.2007 | 1.5 | LI30 | 41 | 4 | HARDWARE | HW Problem | 30-4 MOD fault, won't reset
122843 | 42048.6557 | 1.5 | LI21 | 31 | 4 | HARDWARE | HW Problem | Kly Mod 21-3: lock on T20 not working. Tang is broken. Replace lock with lock from shop stock.
123330 | 42071.1314 | 0.5 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3 mod faulted and won't reset remotely
123331 | 42071.1403 | 0.5 | LI18 | 11 | 4 | HARDWARE | HW Problem | 18-1 FLK no reset
123608 | 42081.5558 | 2 | LI27 | 71 | 4 | HARDWARE | HW Problem | 27-2: Seeing some beam volts jitter.
123618 | 42081.6293 | 0.3 | LI26 | 71 | 4 | HARDWARE | HW Problem | 26-7 MOD VVS fault and won't reset remotely
123651 | 42083.2227 | 1 | LI26 | 21 | 4 | HARDWARE | HW Problem | 26-2 has noticeable phase and amplitude jitter.
123051 | 42059.6166 | 0.5 | LI20 | 61 | 4 | HARDWARE | HW Problem | 20-6 MOD fault, won't reset.
123108 | 42061.3524 |  | LI29 | 11 | 4 | HARDWARE | HW Problem | 29-1 CAM fault, IPL PIOP did not fix the problem.
121746 | 41991.6753 | 1 | LI26 | 11 | 4 | HARDWARE | HW Problem | 26-1: no EOLC signal present on Thy driver or on J9 bnc jack. Find and fix problem. This was the old 24-7 with many missing parts. Verify T23, its load R785, and cabling are OK.
121833 | 42009.6434 | 0.5 | LI13 | 61 | 4 | HARDWARE | HW Problem | KLV fault.
122683 | 42041.4667 | 2 | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 (aka L1X) thyratron should be replaced on PAMM. It has phase jitter of 0.18 deg, compared to 0.15 deg in July 2014, or compared to 0.11 for XTCAV. It sees a timing" like jitter, which we believe is a thyratron grid break-down effecting directly the trigger system, so the "timing" is moving. If it can't be fixed by ranging, we are to close to tripping the modulator, we should replace the thyratron.
122987 | 42056.7188 | 0.75 | LI21 | 11 | 4 | HARDWARE | HW Problem | L1S struggling to make up its amplitude. Possibly related to reduction in beam voltage during the 2/11 PAMM? See this MCC elog entry: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=766157
123205 | 42065.8563 |  | LI02 | 61 | 4 | HARDWARE | HW Problem | 2-6: RF AMPL low OOT, BV/timing OK, see HSTB.
123217 | 42066.5814 | 2.5 | LI29 | 21 | 4 | HARDWARE | HW Problem | 29-2 have narrow width = 2.88uS and slope = 0.7
123224 | 42066.9127 | 1 | LI00 | 11 | 4 | HARDWARE | HW Problem | 1-1A is tripped off, remote resets are unsuccessful.
121732 | 41991.1287 | 0.5 | LI18 | 31 | 4 | HARDWARE | HW Problem | 18-3 (KLYS:LI18:31) WGV
122303 | 42026.6276 | 1 | LI17 | 41 | 4 | HARDWARE | HW Problem | Kly Mod 17-4: use dipstick to check xformer oil level and add oil as needed.
122383 | 42030.8201 | 0.7 | LI05 | 31 | 4 | HARDWARE | HW Problem | Remove PAD, replace fan and re-install PAD.
122384 | 42030.8219 | 0.7 | LI06 | 41 | 4 | HARDWARE | HW Problem | Remove PAD, replace fan and re-install PAD.
122393 | 42031.2491 |  | LI00 | 11 | 4 | HARDWARE | HW Problem | 01-1A ACC-1 Flashing on the MKSU
122595 | 42038.5369 | 1 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4: Seeing some beam volts jitter
123160 | 42062.9501 | 0.75 | LI06 | 61 | 4 | HARDWARE | HW Problem | 6-6 WGV fault
122096 | 42017.2737 | 1 | LI18 | 31 | 4 | HARDWARE | HW Problem | 18-3 EOLC stack needs to be replaced.
122200 | 42022.329 | 1 | LI27 | 81 | 4 | HARDWARE | HW Problem | Klys 27-8 tripping frequently
122387 | 42030.8272 | 0.7 | LI15 | 71 | 4 | HARDWARE | HW Problem | Remove PAD, replace fan and re-install PAD.
122388 | 42030.8286 | 0.7 | LI18 | 21 | 4 | HARDWARE | HW Problem | Remove PAD, replace fan and re-install PAD.
122443 | 42032.5963 | 3 | LI22 | 51 | 4 | HARDWARE | HW Problem | 22-5 PFN strap needs to be replaced (lower, RHS, no! 2 ) Charging choke cable is loose.
122443 | 42032.5963 | 1 | LI22 | 51 | 4 | HARDWARE | HW Problem | 22-5 PFN strap needs to be replaced (lower, RHS, no! 2 ) Charging choke cable is loose.
122868 | 42050.5527 | 0.5 | LI21 | 61 | 4 | HARDWARE | HW Problem | (see pic. in MCC log book)
123250 | 42067.6738 | 3 | LI28 | 81 | 4 | HARDWARE | HW Problem | 28-8: Replace PFN coils 5&6 for better tuning.
123261 | 42068.3884 |  | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2 PHJ - PAD stuck in track mode. 360 scan shows phase at zero crossing.
123264 | 42068.4529 | 0.75 | LI01 | 31 | 4 | HARDWARE | HW Problem | 1-3: Seeing a lot of beam volts jitter.
123290 | 42069.393 | 0.5 | LI17 | 61 | 4 | HARDWARE | HW Problem | 17-6: Tried resetting but not clearing.
123293 | 42069.455 | 1 | LI13 | 41 | 4 | HARDWARE | HW Problem | 13-4: Seeing some beam volts jitter
123318 | 42070.3391 | 0.25 | LI26 | 61 | 4 | HARDWARE | HW Problem | Station 26-6 BVLT has dropped down 90% or more 12 times in last 24 hours.
123322 | 42070.5405 | 1 | LI13 | 41 | 4 | HARDWARE | HW Problem | Klys 13-4 has ever growing AMPL jitter.
123322 | 42070.5405 | 0.8 | LI13 | 41 | 4 | HARDWARE | HW Problem | Klys 13-4 has ever growing AMPL jitter.
123322 | 42070.5405 | 1 | LI13 | 41 | 4 | HARDWARE | HW Problem | Klys 13-4 has ever growing AMPL jitter.
123322 | 42070.5405 | 1 | LI13 | 41 | 4 | HARDWARE | HW Problem | Klys 13-4 has ever growing AMPL jitter.
123324 | 42070.5525 | 1 | LI23 | 11 | 4 | HARDWARE | HW Problem | Klys 23-1 having bad AMPL jitter.
121991 | 42013.1786 | 0.5 | LI27 | 51 | 4 | HARDWARE | HW Problem | 27-5 Phase Loss flt will not clear
122753 | 42046.2538 | 1 | LI04 | 41 | 4 | HARDWARE | HW Problem | 9-4 needs PFN Cap replaced due to oil leak and arc marks
123165 | 42063.4984 | 1 | LI15 | 31 | 4 | HARDWARE | HW Problem | looks like it started on the 25 th around 2000 then jumped to 2.0 deg.s on the 27 th around 1130
123166 | 42063.6226 | 2 | LI15 | 31 | 4 | HARDWARE | HW Problem | Kly Mod 15-3: DMM shows res volts unstable, varying by .050 volts consistently. Sprayed variac wiper and coil: stabilized res volts temporarily. Variac should be replaced next FACET PAMM.
123219 | 42066.5839 |  | LI27 | 41 | 4 | HARDWARE | HW Problem | 27-4: ON/OFF function in REMOTE mode is not working Update from ops, 5/15: Klys 27-4 is MOD and won't reset remotely
123219 | 42066.5839 | 4 | LI27 | 41 | 4 | HARDWARE | HW Problem | 27-4: ON/OFF function in REMOTE mode is not working Update from ops, 5/15: Klys 27-4 is MOD and won't reset remotely
123219 | 42066.5839 | 5 | LI27 | 41 | 4 | HARDWARE | HW Problem | 27-4: ON/OFF function in REMOTE mode is not working Update from ops, 5/15: Klys 27-4 is MOD and won't reset remotely
123219 | 42066.5839 | 2 | LI27 | 41 | 4 | HARDWARE | HW Problem | 27-4: ON/OFF function in REMOTE mode is not working Update from ops, 5/15: Klys 27-4 is MOD and won't reset remotely
123219 | 42066.5839 | 1 | LI27 | 41 | 4 | HARDWARE | HW Problem | 27-4: ON/OFF function in REMOTE mode is not working Update from ops, 5/15: Klys 27-4 is MOD and won't reset remotely
123222 | 42066.6948 | 0.75 | LI02 | 31 | 4 | HARDWARE | HW Problem | LI2-3 MOD fault not reset remotely
122340 | 42028.4804 | 1 | LI18 | 31 | 4 | HARDWARE | HW Problem | Kly Mod 18-3: Check T20 oil level. Add oil as needed.
122844 | 42048.6596 | 2 | LI21 | 31 | 4 | HARDWARE | HW Problem | Kly Mod 21-3: T20 door now has a unique lock that requires second interior door switch key to unlock. This is temp fix only to get mod safe for ops. Mod requires all 5 locks unlocked by interior interlocked door switch key replaced and keyed to one key only that fits the interlock switch inside mod thyratron door.
122853 | 42049.0821 | 1 | DR13 | 1 | 4 | HARDWARE | HW Problem | DR13 compressor klystron phase changing spontaneously. Trimming doesn't actually set it to the desired phase, but somewhere within about 1 degrees of the desired phase. This is significant enough to change the energy and compression scheme.
123044 | 42059.5225 | 0.3 | LI26 | 41 | 4 | HARDWARE | HW Problem | 26-4 MOD locked off after HV work, just needs PEM to bring it on.
123368 | 42073.2009 | 0.3 | LI26 | 61 | 4 | HARDWARE | HW Problem | This klystron MOD faults and then goes to HVR. We are able to reset it, however it trips again about 10 minutes later.
123368 | 42073.2009 |  | LI26 | 61 | 4 | HARDWARE | HW Problem | This klystron MOD faults and then goes to HVR. We are able to reset it, however it trips again about 10 minutes later.
123384 | 42073.4781 | 1 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 Modulator, Field mod for drop outs
123413 | 42074.4817 | 1 | LI02 | 71 | 4 | HARDWARE | HW Problem | 2-7 Mod T20 compartment cooling fan is dead
123436 | 42075.3743 | 1 | LI02 | 31 | 4 | HARDWARE | HW Problem | 2-3: Kly has been jittery at lest from 4 weeks ago. see 127690.
123436 | 42075.3743 | 0.5 | LI02 | 31 | 4 | HARDWARE | HW Problem | 2-3: Kly has been jittery at lest from 4 weeks ago. see 127690.
123436 | 42075.3743 | 1 | LI02 | 31 | 4 | HARDWARE | HW Problem | 2-3: Kly has been jittery at lest from 4 weeks ago. see 127690.
123451 | 42075.7097 | 0.75 | LI15 | 71 | 4 | HARDWARE | HW Problem | Klystron LI15-71 is tripped off on "Mod Fault" and will not reset remotely.
123453 | 42075.8539 | 0.5 | LI22 | 51 | 4 | HARDWARE | HW Problem | station 22-5 Waveguide Vacuum Fault
121850 | 42010.2692 | 0 | LI30 | 71 | 4 | HARDWARE | HW Problem | Klys 30-7 showing beam volts jitter
121850 | 42010.2692 | 1 | LI30 | 71 | 4 | HARDWARE | HW Problem | Klys 30-7 showing beam volts jitter
122311 | 42026.9242 | 1 | LI04 | 51 | 4 | HARDWARE | HW Problem | 4-5 SCRTD appears dead
122404 | 42031.4262 | 1 | LI25 | 41 | 4 | HARDWARE | HW Problem | 25-4: Bad straps: upper (Lhs #3); lower (Lhs #3, #4).
122640 | 42040.2069 | 1 | LI14 | 61 | 4 | HARDWARE | HW Problem | 14-6 Mod flt will not remote clear.
123266 | 42068.4692 |  | LI11 | 51 | 4 | HARDWARE | HW Problem | 11-5: MOD fault, won't reset.
123271 | 42068.5402 | 0.75 | LI26 | 61 | 4 | HARDWARE | HW Problem | 26-6: Beam Volts is very low OOT
123289 | 42069.3535 | 2.5 | LI21 | 41 | 4 | HARDWARE | HW Problem | 21-4 needs PFN tune. PW = 2.92uS; slope = 0.4; ripple = 0.3
123301 | 42069.6348 | 1 | LI09 | 81 | 4 | HARDWARE | HW Problem | Klys 9-8 cycling on TOC frequently.
123309 | 42070.0534 |  | LI26 | 31 | 4 | HARDWARE | HW Problem | 26-3 shows a phase of -70 degrees. Every time we press the trim button, the phase shifter moves by 70 degrees, but the phase readback doesn't change. We've hit the button a bunch of times, so this station is severely mis-phased. Maybe a PAD problem? Put the station into CKP, so that it doesn't mistakenly get put on the beam.
123315 | 42070.3172 | 0.5 | LI25 | 71 | 4 | HARDWARE | HW Problem | Station 25-7 BVLT has dropped down 25% or more 15 times in last 24 hours.
123315 | 42070.3172 | 2 | LI25 | 71 | 4 | HARDWARE | HW Problem | Station 25-7 BVLT has dropped down 25% or more 15 times in last 24 hours.
123316 | 42070.3229 | 0.5 | LI25 | 81 | 4 | HARDWARE | HW Problem | Station 25-8 BVLT has dropped down 25% or more 14 times in last 24 hours.
123316 | 42070.3229 |  | LI25 | 81 | 4 | HARDWARE | HW Problem | Station 25-8 BVLT has dropped down 25% or more 14 times in last 24 hours.
123323 | 42070.5416 |  | LI09 | 51 | 4 | HARDWARE | HW Problem | Klys 9-5 has large phase jitter during day. Quiets at night. Not seeing on beam volts.
121951 | 42012.304 | 1 | LI30 | 81 | 4 | HARDWARE | HW Problem | 30-8 Grid trigger chassis is broken. thy comp light bulb burnt out
122148 | 42019.438 | 0.5 | LI22 | 81 | 4 | HARDWARE | HW Problem | 22-8 MOD fault
122304 | 42026.6311 | 1.5 | LI17 | 51 | 4 | HARDWARE | HW Problem | Kly Mod 17-5: check oil level with dipstick and add oil as necessary.
122464 | 42033.287 |  | LI28 | 11 | 4 | HARDWARE | HW Problem | LI28 LCW system was raised by .5 deg F. Please check rotometer flow.
123475 | 42076.5047 | 1 | LI11 | 51 | 4 | HARDWARE | HW Problem | 11-5 dropped to "Green Light". ERRLOG says "ACC1" faults.
123476 | 42076.5122 | 0.8 | LI12 | 21 | 4 | HARDWARE | HW Problem | 12-2: Seeing some beam volts jitter. VRES at 4.61, high limit at 4.75. Unable to range out BVJ
123476 | 42076.5122 | 1 | LI12 | 21 | 4 | HARDWARE | HW Problem | 12-2: Seeing some beam volts jitter. VRES at 4.61, high limit at 4.75. Unable to range out BVJ
123496 | 42077.3475 | 1 | LI13 | 21 | 4 | HARDWARE | HW Problem | 13-2 is suffering from frequent over-voltage faults
123501 | 42077.5696 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 (L1X) has an MKSU trigger enable fault
123501 | 42077.5696 | 2 | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 (L1X) has an MKSU trigger enable fault
123501 | 42077.5696 | 2 | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 (L1X) has an MKSU trigger enable fault
123507 | 42077.7519 | 0.5 | LI23 | 51 | 4 | HARDWARE | HW Problem | Klystron LI23-51 is tripped off on Mod Fault and will not reset remotely.
123515 | 42078.4993 | 1.5 | LI15 | 31 | 4 | HARDWARE | HW Problem | Klystron LI15-31 has no PAD phase and zero PAD mixer gain readings. IPLing PIOP has no effect. Suspect bad PAD.
121887 | 42010.6034 | 5 | LI25 | 71 | 4 | HARDWARE | HW Problem | 25-7 VVS tripped
122130 | 42018.5544 | 2 | LI06 | 81 | 4 | HARDWARE | HW Problem | 6-8: K26 need to replace three contactors. 6-8: K27 needs to replace one contactor at L3. 6-8: Inside the Thyratron compartment: the third carborundum resistor count from the left is bad. 6-8: The door #2's lock is keep turning around.
122261 | 42025.2724 | 7 | LI20 | 81 | 4 | HARDWARE | HW Problem | L0B (20-8) BV Jitter FTP is not bad, but looking at the regular FTP, the HV turn-on time is jumping around. PEM investigated and replaced the thyratron.
122381 | 42030.817 | 0.7 | LI04 | 51 | 4 | HARDWARE | HW Problem | Remove PAD, replace fan and re-install PAD.
122465 | 42033.2893 |  | LI28 | 21 | 4 | HARDWARE | HW Problem | Klystron 28-2 has a MOD fault and we can't recover it
122690 | 42041.7039 | 0.75 | LI11 | 81 | 4 | HARDWARE | HW Problem | 11-8 MOD fault
123335 | 42071.2998 | 0.5 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3 (KLYS:LI23:31) mod faulted and won't reset remotely
123375 | 42073.4242 |  | LI26 | 31 | 4 | HARDWARE | HW Problem | Klys li26 31 has 'Fox Home error'. Run the Fox Home test, and the result comes back 'Home Position was OK. Old = 106.2, New = 106.2". But the STAT bit for 'Fox Phase Home Error' never gets cleared (STAT = 80 hex).
123376 | 42073.4683 | 0.5 | LI13 | 41 | 4 | HARDWARE | HW Problem | 13-4 despiking coil has a broken resistor
123380 | 42073.4732 | 0.75 | LI14 | 41 | 4 | HARDWARE | HW Problem | 14-4 Mod EOLC stack needs to be replaced
123381 | 42073.4745 | 0.75 | LI14 | 51 | 4 | HARDWARE | HW Problem | 14-5 Mod despiking coil coil has a broken resistor
123383 | 42073.4769 | 0.75 | LI14 | 81 | 4 | HARDWARE | HW Problem | 14-8 Mod has a broken De-spiking coil resistor
122209 | 42023.3862 | 2 | LI17 | 51 | 4 | HARDWARE | HW Problem | 17-5: Here is the list below: 1. Resistor R258-R259-.5-ohm have burnt insulation. 2. K26 need to replace the contactors L1-L3. 3. K27 need to replace the contactors: L1 and L3. 4. T20 terminal at R2 has oil leak. 5. Loose strap: lower (Lhs #2).
123070 | 42060.3661 | 0.75 | LI10 | 51 | 4 | HARDWARE | HW Problem | 10-5 Charging diode stack caps are blown, Needs R/R
123157 | 42062.8259 | 0.5 | LI06 | 61 | 4 | HARDWARE | HW Problem | WGV fault on 6-6
124249 | 42104.5266 | 0.75 | LI30 | 31 | 4 | HARDWARE | HW Problem | 30-3: Seeing some beam volts jitter.
124250 | 42104.5274 | 0.75 | LI27 | 81 | 4 | HARDWARE | HW Problem | 27-8: Seeing some beam volts jitter.
124257 | 42104.8302 | 0.75 | LI11 | 61 | 4 | HARDWARE | HW Problem | 11-6 mod faulted and won't reset remotely
124260 | 42105.4702 | 0.5 | LI17 | 61 | 4 | HARDWARE | HW Problem | Tried three times but failed turn on.
124271 | 42106.0369 |  | LI29 | 51 | 4 | HARDWARE | HW Problem | klys 29-5 showing beam volts jitter.
124271 | 42106.0369 | 2 | LI29 | 51 | 4 | HARDWARE | HW Problem | klys 29-5 showing beam volts jitter.
124271 | 42106.0369 | 2 | LI29 | 51 | 4 | HARDWARE | HW Problem | klys 29-5 showing beam volts jitter.
124272 | 42106.0769 |  | LI29 | 51 | 4 | HARDWARE | HW Problem | (see pic. in MCC log.)
124278 | 42106.5266 | 1 | LI26 | 71 | 4 | HARDWARE | HW Problem | 26-7 VVS fault. CB4 tripping.
124278 | 42106.5266 | 0 | LI26 | 71 | 4 | HARDWARE | HW Problem | 26-7 VVS fault. CB4 tripping.
121654 | 41986.3891 | 1 | LI22 | 81 | 4 | HARDWARE | HW Problem | 22-8 MOD fault no reset
122155 | 42019.5549 | 1.5 | LI09 | 61 | 4 | HARDWARE | HW Problem | 9-6: K26 need to replace the contactors. 9-6: T20: secondary resistor network at R2 needs to replace. 9-6: Loose straps: upper (Lhs #1, #3).
122159 | 42019.5717 | 1 | LI16 | 31 | 4 | HARDWARE | HW Problem | 16-3: K26 need to replace the contactors. 16-3: Loose straps: upper (Rhs #2, #4); lower (Lhs #1).
122769 | 42046.5973 | 2 | LI29 | 51 | 4 | HARDWARE | HW Problem | 29-5 T20 Input snubber wires are badly burnt
123387 | 42073.4802 | 1 | LI23 | 61 | 4 | HARDWARE | HW Problem | 23-6 Modulator, Field mod for drop outs. upgrade hardware and software as needed.
123388 | 42073.4821 | 0.75 | LI00 | 1 | 4 | HARDWARE | HW Problem | K01 Fault stepper is showing a error message that will not reset.
123394 | 42073.5686 | 1 | LI26 | 31 | 4 | HARDWARE | HW Problem | 26-3 Fox phase home error
123396 | 42073.6438 | 2 | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2 new thyratron needs independent grid monitor for grid 1 trigger monitor (vs. grid 2)
123409 | 42074.4776 | 0.75 | LI17 | 41 | 4 | HARDWARE | HW Problem | 17-4 Core bias Variac needs to be replaced K24 relay has a burnt lug that needs r/r
123442 | 42075.4646 | 1 | LI03 | 41 | 4 | HARDWARE | HW Problem | Downtime; 3-(4, 6 and 8) has phase jitter of 0.4 to 0.8 sigma. jitter comes and goes Jitter is sigma= 0.5 to 0.8 (see pic.s in mcc log)
123442 | 42075.4646 | 1 | LI03 | 41 | 4 | HARDWARE | HW Problem | Downtime; 3-(4, 6 and 8) has phase jitter of 0.4 to 0.8 sigma. jitter comes and goes Jitter is sigma= 0.5 to 0.8 (see pic.s in mcc log)
123442 | 42075.4646 | 1 | LI03 | 41 | 4 | HARDWARE | HW Problem | Downtime; 3-(4, 6 and 8) has phase jitter of 0.4 to 0.8 sigma. jitter comes and goes Jitter is sigma= 0.5 to 0.8 (see pic.s in mcc log)
122458 | 42032.9512 | 1.5 | LI20 | 41 | 4 | HARDWARE | HW Problem | Check that heaters are timed out, focus magnets are on, TWT is on, etc.
122610 | 42038.9413 | 1 | LI04 | 21 | 4 | HARDWARE | HW Problem | 4-2 needs PFN tuning
122708 | 42043.5724 | 0.8 | LI05 | 31 | 4 | HARDWARE | HW Problem | station 5-3 mod faulted-not resetable from SCP
122708 | 42043.5724 |  | LI05 | 31 | 4 | HARDWARE | HW Problem | station 5-3 mod faulted-not resetable from SCP
122787 | 42047.3255 | 1 | LI30 | 51 | 4 | HARDWARE | HW Problem | 30-5 does seem to consistently come back but is causing frequent disruptions to the beam
123187 | 42065.4227 |  | LI15 | 31 | 4 | HARDWARE | HW Problem | 15-3: Intermittent PHJ due to PAD phase mode in "Track"
123200 | 42065.6408 | 1 | LI21 | 51 | 4 | HARDWARE | HW Problem | 21-5 IPA has poor attenuation. Replace IPA and resaturate klystron.
121752 | 41992.3469 | 1 | LI01 | 51 | 4 | HARDWARE | HW Problem | Need to corrected.
123595 | 42081.1361 | 1 | LI30 | 21 | 4 | HARDWARE | HW Problem | 30-2 BVLT has dropped down 50% or more 4 times in last 24 hours.
123595 | 42081.1361 | 2 | LI30 | 21 | 4 | HARDWARE | HW Problem | 30-2 BVLT has dropped down 50% or more 4 times in last 24 hours.
123595 | 42081.1361 |  | LI30 | 21 | 4 | HARDWARE | HW Problem | 30-2 BVLT has dropped down 50% or more 4 times in last 24 hours.
123694 | 42084.5518 | 1 | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 MOD fault no reset.
123761 | 42088.5814 | 1 | LI21 | 31 | 4 | HARDWARE | HW Problem | 21-3: Modulator, modify Enerpro Board
124029 | 42097.2716 |  | LI17 | 41 | 4 | HARDWARE | HW Problem | Klystron LI17-41 PIOP does not IPL. This station had this same fault on 01 April 2015, CATER 123961. The PIOP was swapped, see Solution #1. CATER was Closed. Station is faulted again for same fault, just 3 days later. Station does not respond to multiple IPL reset attempts.
122876 | 42051.0427 | 0.5 | LI21 | 71 | 4 | HARDWARE | HW Problem | 21-7 MOD fault, won't reset
122887 | 42051.6691 | 0.8 | LI21 | 71 | 4 | HARDWARE | HW Problem | Klystron LI21-7 is tripped off and won't reset remotely. This station is running OK. /phill
122889 | 42051.8601 | 0.8 | LI28 | 81 | 4 | HARDWARE | HW Problem | LI28-8 HVR won't come on.
122889 | 42051.8601 | 0.25 | LI28 | 81 | 4 | HARDWARE | HW Problem | LI28-8 HVR won't come on.
122889 | 42051.8601 | 2 | LI28 | 81 | 4 | HARDWARE | HW Problem | LI28-8 HVR won't come on.
123470 | 42076.4316 | 0.5 | LI21 | 61 | 4 | HARDWARE | HW Problem | Look like it been that way for the last 5 days at least. PFN and BV looks great at mod.
123487 | 42076.9089 | 2 | LI13 | 11 | 4 | HARDWARE | HW Problem | 13-1 mod fault.
123494 | 42077.3002 |  | LI00 |  | 4 | HARDWARE | HW Problem | K01 Microperveance is out of tols at 1.72
123502 | 42077.5719 | 1 | LI21 | 21 | 4 | HARDWARE | HW Problem | Klystron LI21-21 (L1-X) has tripped off on MKSU Trigger Enable Fault and will not reset remotely.
123614 | 42081.5874 | 0.25 | LI20 | 71 | 4 | HARDWARE | HW Problem | 20-7 MOD faulted and won't reset remotely
123635 | 42082.4964 | 0.5 | LI20 | 51 | 4 | HARDWARE | HW Problem | 20-5 TCAV0 MOD fault no reset
123660 | 42083.4121 | 1.5 | LI11 | 81 | 4 | HARDWARE | HW Problem | 11-8 KHD.
123786 | 42089.3062 | 3.5 | LI24 | 81 | 4 | HARDWARE | HW Problem | 24-8 install E2V 80 kV thyratron for testing without DC power supplies for reservoir and filament
121926 | 42011.5671 | 1 | LI01 | 21 | 4 | HARDWARE | HW Problem | 1-2 charging diode stack has a broken diode.
122002 | 42013.3944 | 1 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 Mod, adjust station to required beam volts.
122006 | 42013.4212 |  | LI28 | 61 | 4 | HARDWARE | HW Problem | Modulator 28-6 has had at least 3 HVR faults so far on day shift. OPS are able to reset these faults remotely, but the station soon trips again. This is a nuisance, so are taking this station off the beam and swapping in spare station 28-4.
122006 | 42013.4212 | 1 | LI28 | 61 | 4 | HARDWARE | HW Problem | Modulator 28-6 has had at least 3 HVR faults so far on day shift. OPS are able to reset these faults remotely, but the station soon trips again. This is a nuisance, so are taking this station off the beam and swapping in spare station 28-4.
122058 | 42016.4054 | 2 | LI06 | 31 | 4 | HARDWARE | HW Problem | 6-3: AC filter 30VSK6C top near FL2 - loose wires 6-3: K26 contactors are bad. 6-3: T20 primary capacitor C60=25uF - leak. 6-3: Loose strap: Upper (LHS #1).
122210 | 42023.3916 | 2 | LI17 | 61 | 4 | HARDWARE | HW Problem | 17-6: Here is the list below: 1. K26 needs to replace L3. 2. K27 needs to replace L1. 3. T20 AC inputs: H1 and H2 terminals loose bolts. 4. T20 terminal at R3 has a oil leak. 5. Loose strap: lower (Rhs #2). 6. EOLC stack have many caps burnt/cracked.
122212 | 42023.3975 | 1 | LI17 | 81 | 4 | HARDWARE | HW Problem | 17-8: Here is the list below: 1. K26 needs to replace a contactor L2. 2. Loose strap: upper (Lhs #1).
122212 | 42023.3975 | 1 | LI17 | 81 | 4 | HARDWARE | HW Problem | 17-8: Here is the list below: 1. K26 needs to replace a contactor L2. 2. Loose strap: upper (Lhs #1).
122342 | 42028.4834 | 0.5 | LI18 | 71 | 4 | HARDWARE | HW Problem | Kly Mod 18-7: check T20 oil level. Add oil as necessary.
122631 | 42039.7388 | 0.5 | LI20 | 51 | 4 | HARDWARE | HW Problem | 20-5 modulator fault won't reset remotely
123568 | 42080.2655 | 1 | LI28 | 61 | 4 | HARDWARE | HW Problem | 28-6: HVPS Regulator Board needs modification
123756 | 42088.5467 | 1 | LI18 | 71 | 4 | HARDWARE | HW Problem | 18-7 Volt meter for the kly filament is broken and needs to R/R
123936 | 42094.741 | 1 | LI13 | 81 | 4 | HARDWARE | HW Problem | Messaged go between FLK,EMC,TRG
122528 | 42035.5208 |  | LI01 | 21 | 4 | HARDWARE | HW Problem | See CATER 122523 about this same problem. PEM had it cleared after MKSU swap but now the false MOD fault status is back.
122976 | 42055.4319 | 1 | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7 Beam voltage jitter
123066 | 42060.2978 | 0.5 | LI26 | 41 | 4 | HARDWARE | HW Problem | Making tuning diff.
123068 | 42060.306 | 0.5 | LI30 | 51 | 4 | HARDWARE | HW Problem | Making tuning hard. Ranging OK, fault log showing many external faults.
123895 | 42093.7689 | 1 | LI29 | 61 | 4 | HARDWARE | HW Problem | 29-6: DQng Under-Current faults. DQng low at 8 amps.
123899 | 42093.9473 | 2 | LI23 | 21 | 4 | HARDWARE | HW Problem | Found no problems other than one "voltage" at 1346 and 218 "AMPL" . See CATER 123849 for more detail.
124039 | 42097.4865 | 1 | LI30 | 51 | 4 | HARDWARE | HW Problem | Klystron 30-5 frequent mod fault.
121957 | 42012.4028 | 2 | LI29 | 81 | 4 | HARDWARE | HW Problem | 29-8 T20 input snubber wires are burnt
122174 | 42020.3169 | 0.25 | LI22 | 41 | 4 | HARDWARE | HW Problem | 22-4 WGV.
122531 | 42035.7993 | 0.5 | LI21 | 71 | 4 | HARDWARE | HW Problem | 21-7 MOD fault, won't reset.
122654 | 42040.4836 | 1 | LI15 | 41 | 4 | HARDWARE | HW Problem | 15-4: Seeing lots of beam volts jitter.
122786 | 42047.2411 | 1 | LI28 | 61 | 4 | HARDWARE | HW Problem | 28-6: We are able to reset remotely but this klystron has become a nuisance
122808 | 42047.5929 | 0.5 | LI01 | 61 | 4 | HARDWARE | HW Problem | Klystron 1-6 went to HVR but then wouldn't reset. Stuck in MOD fault and we can't keep the beam up without it.
122921 | 42053.5308 | 0.5 | LI06 | 51 | 4 | HARDWARE | HW Problem | Klystron 6-5 mod fault.
122925 | 42053.6555 | 1 | LI23 | 81 | 4 | HARDWARE | HW Problem | Klystron 23-8 is tripped on MOD fault and won't reset remotely. We need this station as a working spare.
123741 | 42087.5934 | 1 | LI03 | 21 | 4 | HARDWARE | HW Problem | 3-2 The fire alarm latching box system ok LED is burnt out. T-20 compartment bulb is burnt out
124053 | 42097.7432 |  | LI25 | 71 | 4 | HARDWARE | HW Problem | Amplt. jitter is 3%. (see pic.s in MCC log.)
124056 | 42097.9324 | 1.5 | LI30 | 51 | 4 | HARDWARE | HW Problem | Klys 30-5 is cycling frequently. over 400 AMPL errors since 08:00 today. (about 14 hrs). Update 4/28/15: This station is still having many dropout issues and is currently off the beam because it causes too many trips.
124056 | 42097.9324 | 1 | LI30 | 51 | 4 | HARDWARE | HW Problem | Klys 30-5 is cycling frequently. over 400 AMPL errors since 08:00 today. (about 14 hrs). Update 4/28/15: This station is still having many dropout issues and is currently off the beam because it causes too many trips.
124056 | 42097.9324 | 1 | LI30 | 51 | 4 | HARDWARE | HW Problem | Klys 30-5 is cycling frequently. over 400 AMPL errors since 08:00 today. (about 14 hrs). Update 4/28/15: This station is still having many dropout issues and is currently off the beam because it causes too many trips.
124103 | 42101.2822 | 0.5 | LI27 | 11 | 4 | HARDWARE | HW Problem | 27-1, 27-4 mod fault won't reset remotely after vvs trip
124184 | 42102.5716 | 0.75 | LI12 | 21 | 4 | HARDWARE | HW Problem | 12-2 K24 relay has a burnt lug that needs to be replaced
121655 | 41986.3936 | 1 | LI22 | 41 | 4 | HARDWARE | HW Problem | 22-4 WGV no reset
121770 | 42006.6582 | 1 | LI23 | 21 | 4 | HARDWARE | HW Problem | Timer is set to 55M, but the station goes directly to "MOD Ready"
121855 | 42010.3201 | 4 | LI26 | 11 | 4 | HARDWARE | HW Problem | 26-1 trips CB4 and blows F7 in the PLC pan
121872 | 42010.5434 | 1 | LI04 | 11 | 4 | HARDWARE | HW Problem | 4-1: T20 has a burnt secondary network resistor at R1
122100 | 42017.2924 | 1.5 | LI02 | 21 | 4 | HARDWARE | HW Problem | 2-2 K17 Relay is broken and K27 contacts need to be replaced
122115 | 42017.6286 | 2 | LI30 | 11 | 4 | HARDWARE | HW Problem | KLYS 30-1 has large Beam Volts jitter (5 kV pulse-to-pulse).
122309 | 42026.922 | 0.5 | LI02 | 11 | 4 | HARDWARE | HW Problem | 2-1 at KHD
124017 | 42096.7853 | 0.5 | LI22 | 81 | 4 | HARDWARE | HW Problem | no remote reset
124032 | 42097.3665 |  | LI24 | 51 | 4 | HARDWARE | HW Problem | 24-5 WGV fault Mod 24-5 has bad WGV unit the new one with digital display. James Craft informed instrumentation dept. For replacing this unit.
124034 | 42097.3987 | 1 | LI24 | 51 | 4 | HARDWARE | HW Problem | 24-5: WG VGC shows "pr0" state and will not juice. Controller or gauge?
124049 | 42097.6393 | 0.5 | LI26 | 71 | 4 | HARDWARE | HW Problem | Klys 26-7 has a Mod VVS Voltage Fault
122040 | 42015.1014 | 2 | LI21 | 61 | 4 | HARDWARE | HW Problem | 21-6 has high phase jitter. Phase FTP makes it look like the phase is two-stating rather than just jumping randomly.
122040 | 42015.1014 | 2 | LI21 | 61 | 4 | HARDWARE | HW Problem | 21-6 has high phase jitter. Phase FTP makes it look like the phase is two-stating rather than just jumping randomly.
122509 | 42034.5727 | 1 | LI27 | 21 | 4 | HARDWARE | HW Problem | 27-2 modulator tripping on ACOV
122680 | 42041.4346 | 1 | LI03 | 51 | 4 | HARDWARE | HW Problem | 3-5: Replaced T20. Need VVS breaker closed for check out.
122783 | 42047.0856 | 0.75 | LI12 | 21 | 4 | HARDWARE | HW Problem | 12-2 MOD fault, won't reset
122862 | 42049.6946 |  | LI28 | 31 | 4 | HARDWARE | HW Problem | Klystron 28-3 experiencing phase jumps. Last time it jumped 105 degrees and stayed there.
122862 | 42049.6946 | 0.8 | LI28 | 31 | 4 | HARDWARE | HW Problem | Klystron 28-3 experiencing phase jumps. Last time it jumped 105 degrees and stayed there.
123593 | 42080.8135 | 1 | LI20 | 41 | 4 | HARDWARE | HW Problem | FACET Xtcav repeated MOD faults
123742 | 42087.595 | 1 | LI04 | 31 | 4 | HARDWARE | HW Problem | 4-3 Mod fire alarm latching box LED is burnt out
124062 | 42098.337 | 4 | LI25 | 71 | 4 | HARDWARE | HW Problem | Kly Mod 25-7: klystron history shows frequent bouts over the last 16 hours(since repair of CB 7) of extreme beam volts jitter. Beam volts rapidly fluctuating between 310KV and 346KV.
124062 | 42098.337 |  | LI25 | 71 | 4 | HARDWARE | HW Problem | Kly Mod 25-7: klystron history shows frequent bouts over the last 16 hours(since repair of CB 7) of extreme beam volts jitter. Beam volts rapidly fluctuating between 310KV and 346KV.
124065 | 42098.4689 |  | LI14 | 81 | 4 | HARDWARE | HW Problem | station 14-8 Mod faulting repeatedly
124129 | 42101.5582 | 0.75 | LI13 | 81 | 4 | HARDWARE | HW Problem | 13-8 Mod Fire alarm latch System ok LED is burnt out.
124140 | 42101.5703 | 1 | LI04 | 81 | 4 | HARDWARE | HW Problem | 4-8 1 pair of K26 contacts need to be replaced 1 PFN strap needs R/R (lower, LHS, #2)
122259 | 42025.226 | 5 | LI29 | 81 | 4 | HARDWARE | HW Problem | 29-8 AC input A phase at K26 is badly overheated/burnt. the main AC contactor is damaged as well as AC wiring from main breaker to contactor and will need to be replaced.
122834 | 42048.5802 | 0.5 | LI18 | 41 | 4 | HARDWARE | HW Problem | 18-4 - Station is not firing properly due to frequent cable faults. Not happening to other stations on the same PDU. See the message 'Cable msg timeout- perhaps bad 119MHz or PDU'.
123577 | 42080.5225 | 8 | LI07 | 51 | 4 | HARDWARE | HW Problem | 7-5 Modulator, Remove T20 and Install in 3-7
123634 | 42082.4949 | 0.5 | LI25 | 11 | 4 | HARDWARE | HW Problem | 25-1 phase not stable 360deg FTP shows PAD near null at current OP point HSTB of phase and KPHR shows the trim cycle is chasing a poor PAD measurement. Request stub installed in phase reference side of PAD in coordination with MCC to disable phase trim and regold at the present KPHR setting.
123986 | 42095.6684 | 2 | LI24 | 41 | 4 | HARDWARE | HW Problem | SEE Downtime job, CATER # 123872 Modulator 24-4 120 VAC controllers and current sensing mods needed. Snubbers need to be added to the output of the 4 120 VAC controllers in the PLC pan. The three current sense transducers (Kly Fil, Thy Fil and Thy Res) are monitoring the return line. These three transducers should be monitoring the hot line. The stability of 24-4 will be enhanced when these changes are made.
124075 | 42099.6436 | 1 | LI13 | 11 | 4 | HARDWARE | HW Problem | Mod 13-1 MOD fault does not reset
124075 | 42099.6436 | 1 | LI13 | 11 | 4 | HARDWARE | HW Problem | Mod 13-1 MOD fault does not reset
122140 | 42018.669 | 0.5 | LI21 | 21 | 4 | HARDWARE | HW Problem | There was a valve replaced during the shutdown, to address the frequent flow adjustments that were required in November-December 2014. The temperature rise is much slower now, but has drifted up by ~3 deg in the past 10 days.
122140 | 42018.669 | 1 | LI21 | 21 | 4 | HARDWARE | HW Problem | There was a valve replaced during the shutdown, to address the frequent flow adjustments that were required in November-December 2014. The temperature rise is much slower now, but has drifted up by ~3 deg in the past 10 days.
122140 | 42018.669 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | There was a valve replaced during the shutdown, to address the frequent flow adjustments that were required in November-December 2014. The temperature rise is much slower now, but has drifted up by ~3 deg in the past 10 days.
122140 | 42018.669 | 0.5 | LI21 | 21 | 4 | HARDWARE | HW Problem | There was a valve replaced during the shutdown, to address the frequent flow adjustments that were required in November-December 2014. The temperature rise is much slower now, but has drifted up by ~3 deg in the past 10 days.
123569 | 42080.2667 | 1 | LI29 | 21 | 4 | HARDWARE | HW Problem | 29-2: HVPS Regulator Board needs modification
123806 | 42089.5703 | 0.5 | LI23 | 21 | 4 | HARDWARE | HW Problem | 23-2 long term phase jitter not related to beam voltage
123806 | 42089.5703 |  | LI23 | 21 | 4 | HARDWARE | HW Problem | 23-2 long term phase jitter not related to beam voltage
123806 | 42089.5703 |  | LI23 | 21 | 4 | HARDWARE | HW Problem | 23-2 long term phase jitter not related to beam voltage
121993 | 42013.1809 | 0.5 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4 Phase Loss flt will not clear.
122345 | 42028.6475 |  | LI19 | 21 | 4 | HARDWARE | HW Problem | Kly Mod 19-2: check T20 oil level. Add oil if necessary.
122744 | 42045.2367 |  | LI01 | 51 | 4 | HARDWARE | HW Problem | Unable to fix with IPL thru SCP.. Solution: MCC Ops Swapped PIOP.. IPL worked. Klystron 1-5 back on the beam. Ops noticed it was difficult to remove and insert PIOP
122885 | 42051.5752 | 0.75 | LI11 | 61 | 4 | HARDWARE | HW Problem | Klystron 11-6 mod fault.
122885 | 42051.5752 | 0.75 | LI11 | 61 | 4 | HARDWARE | HW Problem | Klystron 11-6 mod fault.
123719 | 42086.4599 | 1 | LI03 | 41 | 4 | HARDWARE | HW Problem | Seeing large phase jitter on 3-6. On the beam and we are seeing effects downstream resulting from this phase jitter.
123719 | 42086.4599 | 1 | LI03 | 41 | 4 | HARDWARE | HW Problem | Seeing large phase jitter on 3-6. On the beam and we are seeing effects downstream resulting from this phase jitter.
124108 | 42101.3989 |  | LI29 | 61 | 4 | HARDWARE | HW Problem | On 4/6, while setting up for high energy running, 29-6 tripped off, and LCLS could not reach the users' requested beam energy.
124174 | 42102.5273 |  | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4: Seeing some beam volts jitter.
124191 | 42102.5783 | 1 | LI17 | 71 | 4 | HARDWARE | HW Problem | 17-7 Mod Kly trim cooling fan is old and running slow
124194 | 42102.5814 | 0.75 | LI09 | 81 | 4 | HARDWARE | HW Problem | 9-8 Mod CB7 breaker will not latch when actuated 2 out of 3 times it's pulled ..should be replaced.
124195 | 42102.5829 | 0.75 | LI10 | 41 | 4 | HARDWARE | HW Problem | 10-4 Mod grid trigger chassis is broken and needs to be replaced.
124214 | 42103.4865 | 0.75 | LI14 | 11 | 4 | HARDWARE | HW Problem | SB14: RF APML drooping down and finally OOT.
122445 | 42032.5988 | 1 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3 PFN compartment cooling fan is dead
122503 | 42034.3808 | 3 | LI28 | 21 | 4 | HARDWARE | HW Problem | Kly Mod 28-2: excessive phase jitter over last 36 hours. Adjust E2V resistor compensation box back to original setting of 3.7 + 22 ohms to range out jitter by raising tube gas pressure.
122639 | 42040.1877 | 0.5 | LI28 | 81 | 4 | HARDWARE | HW Problem | klystron 28-8 mod fault
123815 | 42089.845 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 had 48 over voltage faults in the last 24 hours
123815 | 42089.845 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 had 48 over voltage faults in the last 24 hours
123862 | 42092.5614 | 0.5 | LI12 | 41 | 4 | HARDWARE | HW Problem | 12-4: Tried to reset but it failed to reset.
124073 | 42099.5387 | 1.5 | LI02 | 81 | 4 | HARDWARE | HW Problem | 2-8 has had many resettable MOD faults
124073 | 42099.5387 | 0.5 | LI02 | 81 | 4 | HARDWARE | HW Problem | 2-8 has had many resettable MOD faults
124073 | 42099.5387 | 0.5 | LI02 | 81 | 4 | HARDWARE | HW Problem | 2-8 has had many resettable MOD faults
124216 | 42103.5075 | 0.1 | LI13 | 21 | 4 | HARDWARE | HW Problem | 13-2: PHJ due to PAD Phase Mode in "Track".
124219 | 42103.555 |  | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3: RF AMPL low OOT. BV dropped 10+ KV at ~5:00 AM and unstable.
124223 | 42103.5823 | 1 | LI25 | 51 | 4 | HARDWARE | HW Problem | 25-5 AIP modulator EPICS screen shows DeQing Overcurrent (readback = 50 A) Went from 21 to 50 Amps at 13:14 on 3/9/15.
124223 | 42103.5823 | 2 | LI25 | 51 | 4 | HARDWARE | HW Problem | 25-5 AIP modulator EPICS screen shows DeQing Overcurrent (readback = 50 A) Went from 21 to 50 Amps at 13:14 on 3/9/15.
121929 | 42011.5787 | 1 | LI02 | 51 | 4 | HARDWARE | HW Problem | 2-5 KMB missing hardware, Snubber board needs to be replaced also. K26 1pair of contacts need to be replaced.
121929 | 42011.5787 | 1 | LI02 | 51 | 4 | HARDWARE | HW Problem | 2-5 KMB missing hardware, Snubber board needs to be replaced also. K26 1pair of contacts need to be replaced.
122150 | 42019.536 | 1.5 | LI09 | 31 | 4 | HARDWARE | HW Problem | 9-3: K26 need to replace three contactors. 9-3: K27 needs to replace L3 contactor. 9-3: HV rectifier stack is showing a burnt mark. 9-3: T20: terminals phase A, B, and C have Oil leak/residure around it. 9-3: Loose straps: lower (Lhs #1, Rhs #1).
122162 | 42019.595 | 1 | LI16 | 61 | 4 | HARDWARE | HW Problem | 16-6: K26 needs to replace contactor L2. 16-6: K27 need to replace contactors L2 and L3. 16-6: HV rectifier stack, the middle stud has a loose screw. 16-6: Primary capacitor network C60 needs to replace. 16-6: T20 at R3: the white wire touching the metal.
122656 | 42040.5357 | 0.5 | LI11 | 21 | 4 | HARDWARE | HW Problem | 11-2 mod fault and won't reset remotely. VVS breaker tripped. Please reset.
122745 | 42045.2557 | 1.2 | LI28 | 31 | 4 | HARDWARE | HW Problem | MKSU PCD03467 from 24-6 When on Beam, trips on KMC and Ext Intlks
123620 | 42081.789 | 1 | LI29 | 41 | 4 | HARDWARE | HW Problem | 29-4 and 29-5 have phase jitter and have been taken off the beam
123762 | 42088.5849 | 1 | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 Snubber board for Phase C needs to be replaced
123863 | 42092.6795 | 2 | LI13 | 21 | 4 | HARDWARE | HW Problem | 13-2: Constant OV faults, suspect intermittent DQ'g box #1.
123863 | 42092.6795 | 0.5 | LI13 | 21 | 4 | HARDWARE | HW Problem | 13-2: Constant OV faults, suspect intermittent DQ'g box #1.
123863 | 42092.6795 | 4 | LI13 | 21 | 4 | HARDWARE | HW Problem | 13-2: Constant OV faults, suspect intermittent DQ'g box #1.
122187 | 42020.5465 | 2 | LI17 | 31 | 4 | HARDWARE | HW Problem | 17-3: Here is the list below: 1. HV divider stack needs to replace. 2. K27 needs to replace contactor L3. 3. T20 has a oil leak/residue around on the floor surface. 4. T20 AC input cables loose. 5. Charging diode stack have the broken caps.
122424 | 42031.9222 |  | LI22 | 81 | 4 | HARDWARE | HW Problem | One of the temperature gauges on klys 22-8 is malfunctioning. It is reads 900-1000 degF.
122537 | 42036.38 | 0.5 | LI23 | 51 | 4 | HARDWARE | HW Problem | 23-5 MOD fault not reset
122653 | 42040.4818 | 0.75 | LI10 | 41 | 4 | HARDWARE | HW Problem | 10-4: Seeing lots of beam volts jitter.
122768 | 42046.5959 | 0.75 | LI05 | 71 | 4 | HARDWARE | HW Problem | 5-7 K13 relay is missing it,s spring
125718 | 42171.3053 | 0.75 | LI30 | 41 | 4 | HARDWARE | HW Problem | 30-4: DQing current low OOT at less than nominal beam volts.
125737 | 42171.7045 | 0.5 | LI25 | 81 | 4 | HARDWARE | HW Problem | station 25-8 is mod faulting repeatedly
125739 | 42171.8898 | 0.5 | LI26 | 61 | 4 | HARDWARE | HW Problem | station 26-6 Mod faulted and not resetting
125747 | 42172.4547 | 2 | LI30 | 31 | 4 | HARDWARE | HW Problem | 30-3 is very jittery and we can clearly see the effects on a downstream BPM. See beam volt jitter/dropouts on FTPs.
125752 | 42172.5405 | 1 | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7 MOD faults much more frequently than the average klystron. This has been happening for a couple of weeks (pretty clear if you look at beam volts history).
125762 | 42172.7338 |  | LI28 | 11 | 4 | HARDWARE | HW Problem | 28-1 fault on ACOC
122042 | 42015.1911 | 0.25 | LI28 | 61 | 4 | HARDWARE | HW Problem | 28-6 MOD faulting frequently. Errors display shows multiple over-voltage faults. The trips are reduced in frequency.
122042 | 42015.1911 | 1 | LI28 | 61 | 4 | HARDWARE | HW Problem | 28-6 MOD faulting frequently. Errors display shows multiple over-voltage faults. The trips are reduced in frequency.
122042 | 42015.1911 | 1.5 | LI28 | 61 | 4 | HARDWARE | HW Problem | 28-6 MOD faulting frequently. Errors display shows multiple over-voltage faults. The trips are reduced in frequency.
122052 | 42015.7245 | 0.75 | LI22 | 81 | 4 | HARDWARE | HW Problem | 22-8 mod faulted and won't reset remotely
122325 | 42027.4671 | 1 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4: Seeing some beam volts jitter
124003 | 42096.526 |  | LI20 | 51 | 4 | HARDWARE | HW Problem | 20-5 (TCAV0): HV CBT again.
124266 | 42105.6008 | 0.5 | LI17 | 61 | 4 | HARDWARE | HW Problem | 17-6 tripped off and we can't get the station back on remotely.
122156 | 42019.559 | 2 | LI09 | 81 | 4 | HARDWARE | HW Problem | 9-8: K26 need to replace the contactors. 9-8: K27 needs to replace contactor L2. 9-8: T20 at R3 terminal loose screw. 9-8: Step start resistors: R258 and R259, their insulation is showing burnt/melt badly.
122184 | 42020.5347 | 2 | LI17 | 11 | 4 | HARDWARE | HW Problem | 17-1: Here is the list below: 1. K26 needs to replace the contactor L3. 2. K27 Need to replace the contactors L1 and L3. 3. Powerline Portable power conditioner has a lead badly. 4. Loose straps: upper (Lhs #1, #4), (Rhs #4). 5. Thyratron compartment: two caps 30uF near DeQing box #2 has a leak.
122346 | 42028.6487 |  | LI19 | 41 | 4 | HARDWARE | HW Problem | Kly Mod 19-4: check T20 oil level. Add oil as necessary.
122686 | 42041.5963 | 0.25 | LI13 | 21 | 4 | HARDWARE | HW Problem | 13-2 MOD fault, no remote reset
123089 | 42060.6917 | 1 | LI24 | 41 | 4 | HARDWARE | HW Problem | 24-4 high phase jitter and amplitude jitter. AMJ and PHJ began at 12:30. BVJ OK so changed shop to AMRF.
123379 | 42073.4718 | 0.75 | LI13 | 81 | 4 | HARDWARE | HW Problem | 13-8 PFN has a burnt/broken strap (lower, RHS, #2 )
123390 | 42073.4847 | 1 | LI01 | 31 | 4 | HARDWARE | HW Problem | 1-3 CB7 plexiglass cover is broken , Top T20 input snubber resistor shows signs of overheating needs r/r befor it get worst
123683 | 42083.7835 | 0.5 | LI22 | 41 | 4 | HARDWARE | HW Problem | 22-4 has a waveguide vacuum fault
123730 | 42087.3547 | 0.3 | LI20 | 51 | 4 | HARDWARE | HW Problem | 20-5 TCAV0 MOD fault no reset
124063 | 42098.3661 |  | LI26 | 21 | 4 | HARDWARE | HW Problem | Reset it a few times. Sorry to make more work for you guys!
121656 | 41986.7529 | 0.75 | LI02 | 21 | 4 | HARDWARE | HW Problem | dfayhadrh adyratyh
121725 | 41990.5019 | 0.5 | LI22 | 81 | 4 | HARDWARE | HW Problem | 22-8 MOD fault no reset.
122386 | 42030.8256 | 0.7 | LI12 | 51 | 4 | HARDWARE | HW Problem | Remove PAD, replace fan and re-install PAD.
122589 | 42038.4124 | 0.8 | LI24 | 61 | 4 | HARDWARE | HW Problem | Seems happy at standby but when assigned to accelerate it trips frequently. Cannot tell from MCC why it is tripping.
122980 | 42055.5763 | 0.5 | LI14 | 81 | 4 | HARDWARE | HW Problem | Klystron 14-8 mod fault.
123004 | 42058.047 |  | LI30 | 61 | 4 | HARDWARE | HW Problem | Klys 30-6 showing phase jitter. Don't see it on AMPL or BVLT, but it shows up frequently on FTPs.
124222 | 42103.5741 | 0.5 | LI21 | 71 | 4 | HARDWARE | HW Problem | 21-7 MOD fault, could not reset remotely
121948 | 42011.8728 | 0.5 | LI22 | 81 | 4 | HARDWARE | HW Problem | 22-8 MOD fault, no remote reset
122099 | 42017.2845 | 2 | LI05 | 71 | 4 | HARDWARE | HW Problem | 5-7: K26 damaged 5-7: T20 H1 primary nut loose 5-7: Transformer primary resistor 10 Ohms R260 measured open; R262 loose screw/nut. 5-7: Transformer secondary network resistor R1 measured open 5-7: Loose strap: upper L #1
122111 | 42017.5961 | 1 | LI15 | 81 | 4 | HARDWARE | HW Problem | 15-8: K38 wire bundle hanging, K26 contacts (C) burnt.
122111 | 42017.5961 | 0.5 | LI15 | 81 | 4 | HARDWARE | HW Problem | 15-8: K38 wire bundle hanging, K26 contacts (C) burnt.
122132 | 42018.5623 | 1 | LI09 | 11 | 4 | HARDWARE | HW Problem | 9-1: Here is the list below: 1. Shunt Trip Board C1 needs to secure by cable tie. 2. K26 need to replace two contactors at L1 and L3. 3. K27 need to replace three contactors: L1-L3. 4. Primary capacitor C60=25uF has a leak. 5. Secondary resistor network: loose resistors at R2 of T20.
122211 | 42023.3953 | 3.5 | LI17 | 71 | 4 | HARDWARE | HW Problem | 17-7: Bad T20 1. T20 has heavy leak at AC input terminals/ base surface / and around it.
122314 | 42027.2573 | 1 | LI03 | 41 | 4 | HARDWARE | HW Problem | 3-4 Control PWR cb would trip at start up
122344 | 42028.6463 |  | LI19 | 11 | 4 | HARDWARE | HW Problem | Kly Mod 19-1: check T20 oil level. Add xformer oil as necessary.
122455 | 42032.9315 | 0.5 | DR03 | 1 | 4 | HARDWARE | HW Problem | 1-7C SSSB blows fuse when powered on.
123640 | 42082.575 | 0.5 | LI26 | 71 | 4 | HARDWARE | HW Problem | 26-7 MOD VVS Voltage Fault, no remote reset.
123645 | 42082.6944 | 0.5 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 MOD fault, won't reset.
123645 | 42082.6944 | 1 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 MOD fault, won't reset.
123693 | 42084.528 | 1 | LI22 | 81 | 4 | HARDWARE | HW Problem | 22-8 MOD fault no reset.
123795 | 42089.4378 | 0.5 | LI23 | 61 | 4 | HARDWARE | HW Problem | The Ethernet cables inside Klystron Modulator 23-6 for the PLC and TP are reversed.
123795 | 42089.4378 |  | LI23 | 61 | 4 | HARDWARE | HW Problem | The Ethernet cables inside Klystron Modulator 23-6 for the PLC and TP are reversed.
123904 | 42094.2356 | 0.5 | LI29 | 21 | 4 | HARDWARE | HW Problem | 29-2 lots of amplitude jitter
124104 | 42101.2832 | 2 | LI29 | 21 | 4 | HARDWARE | HW Problem | 29-2 mod fault won't reset remotely
124286 | 42106.9931 | 2 | LI28 | 11 | 4 | HARDWARE | HW Problem | 28-1 has tripped off 10s of times over the course of this shift. Took off the beam.
124286 | 42106.9931 | 1 | LI28 | 11 | 4 | HARDWARE | HW Problem | 28-1 has tripped off 10s of times over the course of this shift. Took off the beam.
124291 | 42107.3331 | 1 | LI25 | 11 | 4 | HARDWARE | HW Problem | 25-1: MOD fault, won't reset
121978 | 42012.5691 | 1.5 | LI27 | 11 | 4 | HARDWARE | HW Problem | 27-1 Mod De-spiking coil has a broken resistor
122017 | 42013.4982 | 1 | LI25 | 11 | 4 | HARDWARE | HW Problem | 25-1: Beam Volts Jitter (Two-State)
122017 | 42013.4982 | 1.5 | LI25 | 11 | 4 | HARDWARE | HW Problem | 25-1: Beam Volts Jitter (Two-State)
122176 | 42020.3872 | 0.3 | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 Tripping repeatedly
122176 | 42020.3872 | 0.25 | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 Tripping repeatedly
122609 | 42038.8336 | 0.7 | LI25 | 11 | 4 | HARDWARE | HW Problem | Station does read Phase or Amplitude, when trying to saturate, states PIOP failed and has 0 drive. May have more than 1 problem.
123040 | 42059.4204 |  | LI12 | 21 | 4 | HARDWARE | HW Problem | Water Accelerator #1 fault Water Summary Fault PEM called about this one. Trying to get a hold of a instrument tech to take a look.
124102 | 42100.9918 | 0.5 | LI22 | 41 | 4 | HARDWARE | HW Problem | 22-2 WGV. Switched in 22-4 to replace it, which is a station that needs POMM work.
124136 | 42101.5658 | 0.8 | LI02 | 11 | 4 | HARDWARE | HW Problem | 2-1 Mod has a pair of K26 contacts that need to be replaced
124142 | 42101.5714 | 0.75 | LI04 | 61 | 4 | HARDWARE | HW Problem | 4-6 K26 has 1 pair of contacts that need to be replaced.
124161 | 42102.2171 | 0.5 | LI12 | 81 | 4 | HARDWARE | HW Problem | 12-8 Water leaks out of the Klystron drain hole
124198 | 42102.6511 | 2 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6: PFN Tune to reduce jitter.
124206 | 42103.2851 | 1 | LI22 | 71 | 4 | HARDWARE | HW Problem | Also 5 Drive faults.
124218 | 42103.5189 | 2 | LI25 | 81 | 4 | HARDWARE | HW Problem | The issue/klystron was investigated during the 4/7 POMM CATER 123988 and does not appear to be resolved, though Cater 123988 was closed. The Klystron has partial beam volts dropouts which is super noticeable while we run at such a high energy.
124218 | 42103.5189 | 1 | LI25 | 81 | 4 | HARDWARE | HW Problem | The issue/klystron was investigated during the 4/7 POMM CATER 123988 and does not appear to be resolved, though Cater 123988 was closed. The Klystron has partial beam volts dropouts which is super noticeable while we run at such a high energy.
124218 | 42103.5189 | 1 | LI25 | 81 | 4 | HARDWARE | HW Problem | The issue/klystron was investigated during the 4/7 POMM CATER 123988 and does not appear to be resolved, though Cater 123988 was closed. The Klystron has partial beam volts dropouts which is super noticeable while we run at such a high energy.
122499 | 42033.9691 | 0.3 | LI30 | 51 | 4 | HARDWARE | HW Problem | 30-5 has numerous external faults, saw a klystron magnet fault flash on MKSU. Trips reduced. /phill
122524 | 42035.2342 | 1 | LI02 | 21 | 4 | HARDWARE | HW Problem | 2-2 Has a false Mod flt on SCP
122524 | 42035.2342 | 1 | LI02 | 21 | 4 | HARDWARE | HW Problem | 2-2 Has a false Mod flt on SCP
122551 | 42037.415 | 0.5 | LI19 | 11 | 4 | HARDWARE | HW Problem | 19-1: Will not turn on; has EOL overcurrent fault.
122555 | 42037.4717 | 1 | LI14 | 81 | 4 | HARDWARE | HW Problem | Klystron 14-8 has FLK error.
122555 | 42037.4717 | 0.5 | LI14 | 81 | 4 | HARDWARE | HW Problem | Klystron 14-8 has FLK error.
123698 | 42085.0618 | 0.5 | LI11 | 61 | 4 | HARDWARE | HW Problem | Klystron 11-6 mod fault.
123745 | 42087.7057 | 0.5 | LI25 | 71 | 4 | HARDWARE | HW Problem | 25-7 MOD fault, no remote reset
122113 | 42017.598 | 1 | LI14 | 51 | 4 | HARDWARE | HW Problem | 14-5: Door #7 ground hook wire short, core bias variac overheating.
122161 | 42019.5908 | 1 | LI16 | 51 | 4 | HARDWARE | HW Problem | 16-5: Here is the list below: 1. Primary capacitor network C58 and C60, the wires/jumper wire are burnt/melt. 2. T20 at R3 has a oil leak.
122705 | 42043.1296 | 0.5 | LI20 | 41 | 4 | HARDWARE | HW Problem | FACET XTCAV 20-4B mod fault.
122707 | 42043.5189 | 0.5 | LI18 | 11 | 4 | HARDWARE | HW Problem | 18-1: mod faulted
123641 | 42082.577 | 0.5 | LI15 | 11 | 4 | HARDWARE | HW Problem | klys 15-1 mod faulted and won't reset
123652 | 42083.2639 |  | LI25 | 41 | 4 | HARDWARE | HW Problem | Klystron 25-4 frequent mod faults It resets itself 90% of the time.
123691 | 42084.3695 | 1 | LI23 | 81 | 4 | HARDWARE | HW Problem | 23-8 MOD faulted, no reset.
123759 | 42088.5793 | 1 | LI21 | 11 | 4 | HARDWARE | HW Problem | 21-1: Modulator, modify Enerpro Board
123759 | 42088.5793 | 1 | LI21 | 11 | 4 | HARDWARE | HW Problem | 21-1: Modulator, modify Enerpro Board
123930 | 42094.6343 | 0.75 | LI13 | 11 | 4 | HARDWARE | HW Problem | Klystron 13-1 large phase jitter.
122025 | 42013.6472 |  | LI20 | 61 | 4 | HARDWARE | HW Problem | The APC 8 Port AC Power Switch gives a purple status. This switch powers the Gun LLRF PAC and Klys Diagnostic PAD. Same problem as CATER 122183.
122182 | 42020.4733 |  | LI20 | 61 | 4 | HARDWARE | HW Problem | Klys 20-6 (the gun), is undergoing AMPL excursions. The setpoint for the I and Q values jumped up on the Jan 13th around 10:00. See MCC elog, 1/16 @ 11:21
122269 | 42025.4105 | 1 | LI27 | 51 | 4 | HARDWARE | HW Problem | 27-5: keep alive current is over 500mA.
122711 | 42044.312 | 1.5 | LI20 | 41 | 4 | HARDWARE | HW Problem | 20-4B modulator and RF settings were adjusted for RE trips.
123599 | 42081.209 | 0.5 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3 mod faulted, won't reset remotely.
123738 | 42087.5865 | 1 | LI01 | 31 | 4 | HARDWARE | HW Problem | 1-3 K26 has 2 pair of contacts that need to be replaced
124233 | 42104.0633 | 1 | LI25 | 11 | 4 | HARDWARE | HW Problem | No remote reset
124236 | 42104.3631 | 1 | LI04 | 11 | 4 | HARDWARE | HW Problem | Duplicate of 123921. The expected Amplitude is 50 but the value is 36. This station was put back into MNT
122094 | 42017.2706 | 0.75 | LI18 | 11 | 4 | HARDWARE | HW Problem | 18-1 K13 relay is missing the pull back spring
122220 | 42024.3002 | 1 | LI22 | 81 | 4 | HARDWARE | HW Problem | 22-8 MOD fault, no remote reset
122242 | 42024.5622 | 1 | LI24 | 31 | 4 | HARDWARE | HW Problem | See mcc logbook entry at 13:30 today. dropouts have been occurring since 11:00 today
122975 | 42055.4292 | 0.8 | LI23 | 11 | 4 | HARDWARE | HW Problem | 23-1 Beam voltage jitter
122975 | 42055.4292 | 1 | LI23 | 11 | 4 | HARDWARE | HW Problem | 23-1 Beam voltage jitter
122975 | 42055.4292 | 2 | LI23 | 11 | 4 | HARDWARE | HW Problem | 23-1 Beam voltage jitter
124436 | 42111.9766 | 0.75 | LI29 | 81 | 4 | HARDWARE | HW Problem | 29-8 trips off. It can be reset but trips again right away.
124460 | 42114.4493 | 1 | LI19 | 31 | 4 | HARDWARE | HW Problem | 19-3 re-certify NIRP.
124807 | 42129.5731 | 0.5 | LI30 | 21 | 4 | HARDWARE | HW Problem | 30-2 Mod PFN strap needs to be replaced (lower, rhs, #3)
124890 | 42131.7511 | 1 | LI10 | 61 | 4 | HARDWARE | HW Problem | 10-6: ACC1 Water fault - intermittent
125152 | 42142.6571 | 2 | LI00 | 11 | 4 | HARDWARE | HW Problem | Klys 1-1A has phase and amplitude jitter. The beam volts look stable.
125294 | 42147.9179 |  | LI17 | 31 | 4 | HARDWARE | HW Problem | 17-3 SCRTD EOLC signal present not lit
125553 | 42161.4706 | 1 | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 has tripped several times this shift and is currently FLK.
125553 | 42161.4706 |  | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 has tripped several times this shift and is currently FLK.
126144 | 42195.2728 | 0.25 | LI28 | 61 | 4 | HARDWARE | HW Problem | 28-6 mod faulted and won't reset remotely
126281 | 42201.6974 | 1.5 | LI21 | 21 | 4 | HARDWARE | HW Problem | Klystron L1-X (LI21-2) Rf phase jitter.
124695 | 42124.5399 | 1 | LI11 | 21 | 4 | HARDWARE | HW Problem | 11-2: "FLK" on MKOV's
124802 | 42129.5665 | 0.5 | LI30 | 11 | 4 | HARDWARE | HW Problem | 30-1 Mod grid trigger chassis needs to be replaced.
124834 | 42130.507 |  | LI27 | 21 | 4 | HARDWARE | HW Problem | 27-2 showing constant phase jitter, can see on FTP.
125654 | 42166.387 |  | LI26 | 31 | 4 | HARDWARE | HW Problem | We can reset it, but it trips off again within 1 minute. It has tripped 3 times within 5 minutes.
125684 | 42167.671 | 0.5 | LI30 | 31 | 4 | HARDWARE | HW Problem | 30-3 causing energy jumps, seems due to spontaneous AMM faults
125686 | 42167.7033 |  | LI30 | 71 | 4 | HARDWARE | HW Problem | The WVGC measures LO E-10 on the PIOP it reads 3.6 E-9. While the station was off, there was not a change in vacuum. Set point 1 would not be able to effectively pass NIRP.
125691 | 42167.812 | 0.5 | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 MOD faulting repeatedly.
125693 | 42168.1902 | 1 | LI23 | 41 | 4 | HARDWARE | HW Problem | 23-4 bad CBL status Also 24V battery fault, Water summary fault
125715 | 42171.0321 | 0.75 | LI22 | 51 | 4 | HARDWARE | HW Problem | 22-5 frequent trips - REs are showing in fault counter not sure this is why the station is tripping
124532 | 42117.5965 | 1 | LI11 | 61 | 4 | HARDWARE | HW Problem | 11-6 K24 Relay is broken K27 is missing a screw for the contact assem
124533 | 42117.6002 | 1 | LI11 | 51 | 4 | HARDWARE | HW Problem | 11-5 K24 Relay is broken PFN strap needs to be replaced. (upper, LHS, #3 )
124533 | 42117.6002 | 1 | LI11 | 51 | 4 | HARDWARE | HW Problem | 11-5 K24 Relay is broken PFN strap needs to be replaced. (upper, LHS, #3 )
124891 | 42131.7521 | 1 | LI13 | 61 | 4 | HARDWARE | HW Problem | 13-6: Klystron Water fault - solid.
124968 | 42133.7744 | 1 | LI10 | 41 | 4 | HARDWARE | HW Problem | Kly Mod 10-4: mod fault, no reset.
124969 | 42133.9053 | 1 | LI27 | 21 | 4 | HARDWARE | HW Problem | Called in to PEM.
124973 | 42134.2267 | 0.5 | LI27 | 11 | 4 | HARDWARE | HW Problem | 27-1 MOD fault, won't reset
124996 | 42135.5493 | 3 | LI22 | 71 | 4 | HARDWARE | HW Problem | LCLS - Modulators, Beam Volt adjustment: 22-7.
126299 | 42202.785 | 1 | LI23 | 21 | 4 | HARDWARE | HW Problem | Klystron LI23-21 tripping off, but does not give fault indicator. Resets remotely. Fault count display shows 115 Ampl faults at this time. Klystron show higher than normal phase jitter.
126299 | 42202.785 | 1 | LI23 | 21 | 4 | HARDWARE | HW Problem | Klystron LI23-21 tripping off, but does not give fault indicator. Resets remotely. Fault count display shows 115 Ampl faults at this time. Klystron show higher than normal phase jitter.
126307 | 42203.751 | 1 | LI01 | 51 | 4 | HARDWARE | HW Problem | 1-5: Bad Cable Status. Check with PAD Sub-Box.
126307 | 42203.751 |  | LI01 | 51 | 4 | HARDWARE | HW Problem | 1-5: Bad Cable Status. Check with PAD Sub-Box.
126315 | 42204.7158 | 0.1 | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 MOD fault. CB4?
126316 | 42204.7161 | 0.2 | LI30 | 71 | 4 | HARDWARE | HW Problem | 30-7 MOD fault
124583 | 42120.5373 | 0.5 | LI11 | 21 | 4 | HARDWARE | HW Problem | 11-2 MOD faulting repeatedly.
124583 | 42120.5373 | 1 | LI11 | 21 | 4 | HARDWARE | HW Problem | 11-2 MOD faulting repeatedly.
124819 | 42129.7542 | 2 | LI30 | 11 | 4 | HARDWARE | HW Problem | station 30-1 MOD faulted
124821 | 42130.0789 | 0.5 | LI22 | 41 | 4 | HARDWARE | HW Problem | Shows a HVR fault.
125136 | 42141.9534 | 0.75 | LI09 | 41 | 4 | HARDWARE | HW Problem | 9-4 MOD fault, no remote reset
125237 | 42145.5508 | 1 | LI13 | 81 | 4 | HARDWARE | HW Problem | Downtime; 13-8: Seeing some beam volts jitter. Closing this cater for downtime/ repair has been made. phill
125237 | 42145.5508 | 0.8 | LI13 | 81 | 4 | HARDWARE | HW Problem | Downtime; 13-8: Seeing some beam volts jitter. Closing this cater for downtime/ repair has been made. phill
125256 | 42146.1673 | 1 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3 mod fault does not reset
125288 | 42147.8405 | 0.75 | LI10 | 71 | 4 | HARDWARE | HW Problem | 10-7 MOD fault.
125313 | 42149.3694 | 1 | LI29 | 81 | 4 | HARDWARE | HW Problem | Klys 29-8 was cycling frequently, and is now FLK
125328 | 42150.3663 | 2 | LI24 | 41 | 4 | HARDWARE | HW Problem | 24-4 modulator trips about once per hour
125328 | 42150.3663 | 0.5 | LI24 | 41 | 4 | HARDWARE | HW Problem | 24-4 modulator trips about once per hour
125328 | 42150.3663 | 0.25 | LI24 | 41 | 4 | HARDWARE | HW Problem | 24-4 modulator trips about once per hour
125359 | 42151.5886 | 1 | LI23 | 41 | 4 | HARDWARE | HW Problem | 23-4 Mod has a bad PFN strap ( Lower, LHS, #3 ) Despiking coil has a broken resistor
125404 | 42153.4783 | 0.5 | LI21 | 21 | 4 | HARDWARE | HW Problem | PEM dispatched
125404 | 42153.4783 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | PEM dispatched
125404 | 42153.4783 | 2 | LI21 | 21 | 4 | HARDWARE | HW Problem | PEM dispatched
124340 | 42109.3919 | 1 | LI06 | 31 | 4 | HARDWARE | HW Problem | 6-3 MOD faulted.
124443 | 42112.9951 | 0.75 | LI03 | 81 | 4 | HARDWARE | HW Problem | Klystron 3-8 mod fault.
124506 | 42116.5976 | 1 | LI04 | 71 | 4 | HARDWARE | HW Problem | 4-7 Core bias Variac is degraded and attaching lug needs to be replaced.
125358 | 42151.5863 | 1 | LI23 | 61 | 4 | HARDWARE | HW Problem | 23-6 Mod has 2 ea PFN straps that need to be replaced (upper, RHS, #2 ) (Lower, Center #4 )
126317 | 42204.72 | 0.5 | LI24 | 41 | 4 | HARDWARE | HW Problem | 24-4 MOD faulted
126339 | 42205.5454 | 0.5 | LI25 | 21 | 4 | HARDWARE | HW Problem | Klystron 25-2 mod faulted and won't reset remotely
126343 | 42205.6407 | 0.5 | LI22 | 11 | 4 | HARDWARE | HW Problem | 22-1 MOD VVS voltage fault
126347 | 42205.7248 | 2 | LI29 | 31 | 4 | HARDWARE | HW Problem | Downtime 29-3 Modulator, Primary Filter Caps, Change to newer values
124384 | 42110.6009 | 0.2 | LI22 | 61 | 4 | HARDWARE | HW Problem | 22-6 Modulator, Take T20 oil sample.
124541 | 42117.9088 | 0.5 | LI26 | 61 | 4 | HARDWARE | HW Problem | Klystron 26-6 frequent mod faults Able to reset each time. Took it off the beam due to its frequency of trips.
124645 | 42122.8321 | 0.75 | LI27 | 11 | 4 | HARDWARE | HW Problem | 27-1 has mod faulted and will not reset remotely
124645 | 42122.8321 | 1.5 | LI27 | 11 | 4 | HARDWARE | HW Problem | 27-1 has mod faulted and will not reset remotely
124855 | 42131.032 | 0.5 | LI23 | 81 | 4 | HARDWARE | HW Problem | 23-8 lots of beam volts jitter
124947 | 42132.638 | 0.1 | LI20 | 41 | 4 | HARDWARE | HW Problem | FACET XTCAV Kly 20-4 PAC out but no drive.
125508 | 42158.9412 | 1 | LI30 | 31 | 4 | HARDWARE | HW Problem | Klystron 30-3 repeating mod fault.
125514 | 42159.3529 | 1 | LI27 | 21 | 4 | HARDWARE | HW Problem | Kly. LI 27 - 2 phase jitter has gone from 0.23 to 1.43
125525 | 42159.74 | 0.5 | LI19 | 31 | 4 | HARDWARE | HW Problem | 19-3: Noted that on 13 May 2015 abot the time of a FACET turn-on the Focus Current went from 14.5A (spec) to 15.5A
125769 | 42173.3711 |  | LI23 | 11 | 4 | HARDWARE | HW Problem | 23-1 has high phase jitter
125804 | 42174.5678 | 1 | LI04 | 61 | 4 | HARDWARE | HW Problem | 4-6. K27 contacts need to be replaced EOLC stack needs to be replaced PFN strap bad ( Lower, Center jumper strap )
125804 | 42174.5678 | 1 | LI04 | 61 | 4 | HARDWARE | HW Problem | 4-6. K27 contacts need to be replaced EOLC stack needs to be replaced PFN strap bad ( Lower, Center jumper strap )
125808 | 42174.5846 | 1.5 | LI03 | 71 | 4 | HARDWARE | HW Problem | 3-7 K26 and K27 contacts need to be replaced PFN straps Bad (upper , RHS, # 1 ) ( Lower, RHS, #2,3 )
125810 | 42174.5889 | 1 | LI03 | 41 | 4 | HARDWARE | HW Problem | 3-4 EOLC stack needs to be replaced K27 Contacts need to be replaced
125839 | 42177.2973 | 1.5 | LI02 | 61 | 4 | HARDWARE | HW Problem | 2-6 K26/K27 contacts need to be replaced EOLC stack needs to be replaced PFN strap ( Lower, RHS, # 3)
125841 | 42177.3064 | 1 | LI02 | 31 | 4 | HARDWARE | HW Problem | 2-3 EOLC stack Needs to be replaced Thyratron Cables need to be secured from blower motor
125842 | 42177.3085 | 5 | LI02 | 11 | 4 | HARDWARE | HW Problem | 2-1 EOLC stack needs to be replaced. Secure Thyratron cable away from blower motor
125842 | 42177.3085 | 1 | LI02 | 11 | 4 | HARDWARE | HW Problem | 2-1 EOLC stack needs to be replaced. Secure Thyratron cable away from blower motor
125845 | 42177.3195 | 1 | LI01 | 41 | 4 | HARDWARE | HW Problem | 1-4 EOLC stack needs to be replaced 1 Pair of K27 contacts need to be replaced Bleeder resistors in PFN comp have burnt lugs that need to be replaced.
125845 | 42177.3195 | 0.5 | LI01 | 41 | 4 | HARDWARE | HW Problem | 1-4 EOLC stack needs to be replaced 1 Pair of K27 contacts need to be replaced Bleeder resistors in PFN comp have burnt lugs that need to be replaced.
125867 | 42178.5955 | 1 | LI09 | 31 | 4 | HARDWARE | HW Problem | 9-3 charging Diode and EOLC stacks need to be replaced.
124697 | 42124.575 | 16 | LI28 | 41 | 4 | HARDWARE | HW Problem | Electro-Magnetic current fault on klystron 28-4
124697 | 42124.575 | 1 | LI28 | 41 | 4 | HARDWARE | HW Problem | Electro-Magnetic current fault on klystron 28-4
125255 | 42146.0167 | 1 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3 mod faulted and won't reset remotely
125594 | 42163.7095 | 1 | LI23 | 21 | 4 | HARDWARE | HW Problem | 23-2 Waveguide Water Temp KLYS:LI23:21:DL_WG_TEMP Temp is too low and is alarming on the Low Alarm Limit. The Low alarm limit is 112.0 deg F and the current temp value is 112.0 deg F. The location is listed as S23-F5A (23-5)
125835 | 42177.0645 | 0.5 | LI30 | 31 | 4 | HARDWARE | HW Problem | 30-3 MOD fault, won't reset
125836 | 42177.2066 | 1 | LI23 | 41 | 4 | HARDWARE | HW Problem | 23-4 is experiencing a lot of external interlocks recently. Possible loose connection in AC distribution
125853 | 42177.6646 | 1 | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 MOD faulting frequently
125853 | 42177.6646 | 1 | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 MOD faulting frequently
125865 | 42178.5889 | 1 | LI06 | 61 | 4 | HARDWARE | HW Problem | 6-6 Mod PFN straps burnt/worn ( U, RHS, # 3 ) ( L, LHS, # 3 )
126108 | 42192.4934 | 1 | LI09 | 21 | 4 | HARDWARE | HW Problem | 9-2 waveguide vacuum fault
126428 | 42210.419 | 2 | LI22 | 41 | 4 | HARDWARE | HW Problem | Klystron 22-4 RF jitter seen on FTPs. Infrequent jumps in RF phase.
126428 | 42210.419 | 0.5 | LI22 | 41 | 4 | HARDWARE | HW Problem | Klystron 22-4 RF jitter seen on FTPs. Infrequent jumps in RF phase.
126432 | 42211.2112 | 0.5 | LI22 | 11 | 4 | HARDWARE | HW Problem | Discovered this problem while trying to phase scan the station. See this MCC elog entry for a reference image: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=795712
126432 | 42211.2112 |  | LI22 | 11 | 4 | HARDWARE | HW Problem | Discovered this problem while trying to phase scan the station. See this MCC elog entry for a reference image: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=795712
126432 | 42211.2112 |  | LI22 | 11 | 4 | HARDWARE | HW Problem | Discovered this problem while trying to phase scan the station. See this MCC elog entry for a reference image: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=795712
126453 | 42212.5511 | 75 | LI29 | 41 | 4 | HARDWARE | HW Problem | 29-4: Seeing some beam volts jitter.
126460 | 42213.0601 | 1 | LI29 | 41 | 4 | HARDWARE | HW Problem | 29-4 is WGV faulted
124507 | 42116.5989 | 1 | LI03 | 71 | 4 | HARDWARE | HW Problem | 3-7 Core Bias variac needs t o be replaced
124733 | 42125.6463 | 1 | LI21 | 21 | 4 | HARDWARE | HW Problem | Kly Mod 21-2 (L1X): beam volts unstable, similar jumps as yesterday.
124733 | 42125.6463 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | Kly Mod 21-2 (L1X): beam volts unstable, similar jumps as yesterday.
125937 | 42181.9715 | 2.5 | LI26 | 71 | 4 | HARDWARE | HW Problem | 26-7 mod faults due to ACOC, see 125824
126024 | 42187.5241 | 1 | LI22 | 71 | 4 | HARDWARE | HW Problem | 22-7 tripped of klys filament heater fault
124314 | 42108.3251 | 0.5 | LI13 | 61 | 4 | HARDWARE | HW Problem | 13-6: Tried to reset 3 time but no reset.
124338 | 42109.3116 | 0.75 | LI26 | 11 | 4 | HARDWARE | HW Problem | Klys 26-1 showing beam volts jitter.
124426 | 42111.5862 |  | LI29 | 51 | 4 | HARDWARE | HW Problem | 29-5: does not have AC filters.
124737 | 42126.0245 | 1 | LI20 | 51 | 4 | HARDWARE | HW Problem | TCAV0 (20-5) modulator is tripped off, won't reset.
125665 | 42166.5477 | 1 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3: DQing current low OOT at less than nominal beam volts.
125803 | 42174.5656 | 1 | LI04 | 41 | 4 | HARDWARE | HW Problem | 4-7 EOLC Stack needs to be replaced DE-spiking coil resistor is broken
126044 | 42188.5868 | 1 | LI12 | 11 | 4 | HARDWARE | HW Problem | 12-5 K27 contacts need to be replaced.
126067 | 42191.2121 |  | LI30 | 81 | 4 | HARDWARE | HW Problem | Kly. LI - 30 - 8 AMPLT. jitter is 1.2 Started to increase around 6 - 15
126119 | 42193.402 |  | LI29 | 71 | 4 | HARDWARE | HW Problem | 29-7 HVR will not turn on. Severe, immediate ACOC
126556 | 42216.6255 | 1 | LI26 | 51 | 4 | HARDWARE | HW Problem | 26-5 needs PFN tune due to low BV. BV = 347.5KV, PW = 3.09uS, Slope = 0.7, Ripple = 0.3
126571 | 42218.7205 |  | LI22 | 71 | 4 | HARDWARE | HW Problem | When attempting to calibrate the filament. When the filament cable was removed, oil started leaking from the socket. It stopped leaking when I re-attached the cable.
126578 | 42219.3888 | 0.25 | LI20 | 71 | 4 | HARDWARE | HW Problem | Klys 20-7 is KHD
126608 | 42220.5048 | 1 | LI24 | 41 | 4 | HARDWARE | HW Problem | Klystron 24-4 has large phase jitter as seen on the FTP. There is slight jitter on the beam volts.
126608 | 42220.5048 | 1.5 | LI24 | 41 | 4 | HARDWARE | HW Problem | Klystron 24-4 has large phase jitter as seen on the FTP. There is slight jitter on the beam volts.
126637 | 42221.5233 | 0.75 | LI27 | 81 | 4 | HARDWARE | HW Problem | 27-8: Seeing some beam volts jitter.
126647 | 42222.4388 | 1 | LI22 | 31 | 4 | HARDWARE | HW Problem | 22-3: Seeing some beam volts jitter.
126656 | 42222.5543 | 1 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4 relocate deQ divider monitor location to thyratron anode to check for improved pfn stability.
126658 | 42222.6086 | 0.5 | LI22 | 11 | 4 | HARDWARE | HW Problem | Klystron 22-1 VVS mod fault
126669 | 42223.3044 |  | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 MOD faults frequently. See 126540.
126740 | 42237.4522 |  | LI22 | 71 | 4 | HARDWARE | HW Problem | 22-7 1 pair of K26 contacts need to be replaced
124341 | 42109.3954 |  | LI29 | 51 | 4 | HARDWARE | HW Problem | 29-5: RF AMPL low OOT due to beam volts low. 124271 Closed.
124916 | 42132.357 | 0.5 | LI12 | 21 | 4 | HARDWARE | HW Problem | 12-2 is off on a mod fault, unable to reset remotely.
125031 | 42136.997 |  | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7 We are seeing off energy shots from this station dropping out
125031 | 42136.997 |  | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7 We are seeing off energy shots from this station dropping out
125131 | 42141.1159 | 2 | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7 --DUP--BVLT drop outs. see 125031 ~4kV drop outs causing large beam jitter.
125131 | 42141.1159 | 0.5 | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7 --DUP--BVLT drop outs. see 125031 ~4kV drop outs causing large beam jitter.
125656 | 42166.4118 | 1 | LI21 | 11 | 4 | HARDWARE | HW Problem | 21-1 (AKA L1S) Modulator, Clean-up wire harness in rear cabinet. Many wires left loose after recent repair.
126075 | 42191.5264 | 0.3 | LI20 | 71 | 4 | HARDWARE | HW Problem | Station is generally running fine but the PIOP BV and FTP show no BV signal.
126095 | 42192.2685 | 1 | LI15 | 71 | 4 | HARDWARE | HW Problem | 15-7 K27 Contacts need to be replaced Mod charging Diode stack needs to be replaced.
126130 | 42193.5462 | 1 | LI17 | 11 | 4 | HARDWARE | HW Problem | 17-1 1 K38 bleeder resistor needs to be replaced K26 and K27 Contacts need to be replaced
126130 | 42193.5462 |  | LI17 | 11 | 4 | HARDWARE | HW Problem | 17-1 1 K38 bleeder resistor needs to be replaced K26 and K27 Contacts need to be replaced
126154 | 42195.4679 | 1 | LI24 | 41 | 4 | HARDWARE | HW Problem | 24-4 HV regulator board has wrong resistor value.
126156 | 42195.587 | 1 | LI17 | 71 | 4 | HARDWARE | HW Problem | 17-7 Core bias variac has a burnt/broken connection
126159 | 42195.5919 | 1 | LI18 | 41 | 4 | HARDWARE | HW Problem | 18-4 K27 Contacts need to be replaced
126166 | 42195.6042 | 3 | LI18 | 11 | 4 | HARDWARE | HW Problem | 18-1 The red standoff for assembly sitting above the EOLC Is badly eroded from arcing and needs to be replaced
124745 | 42126.8709 | 0.75 | LI11 | 61 | 4 | HARDWARE | HW Problem | 11-6 in MOD fault
124972 | 42133.974 | 0.6 | LI25 | 11 | 4 | HARDWARE | HW Problem | Kly. LI 25 - 1 tripped on "MOD"fault. Won't reset. Reported to PEM.
126507 | 42214.5171 | 2 | LI25 | 81 | 4 | HARDWARE | HW Problem | Klystron 25-8 repeatedly tripping, going FLK. RE's on fault counter. See 126253. *A. Egger here. I will say that it is still tripping repeatedly on reflected energy at end of swing shift 7/29. Not sure it's ready to close.
126507 | 42214.5171 | 0.5 | LI25 | 81 | 4 | HARDWARE | HW Problem | Klystron 25-8 repeatedly tripping, going FLK. RE's on fault counter. See 126253. *A. Egger here. I will say that it is still tripping repeatedly on reflected energy at end of swing shift 7/29. Not sure it's ready to close.
126564 | 42217.35 | 0.5 | LI27 | 31 | 4 | HARDWARE | HW Problem | 27-3 Keeps faulting on HTR time out. reset twice.
126589 | 42219.4393 | 2 | LI24 | 31 | 4 | HARDWARE | HW Problem | Klys 24-3 is tripping constantly, no more spares in L2. The machine wants downtime to start about a week early, and/or drive ops and PEM into a deep chasm of despair. Gaaah! Anyway, if PEM could investigate, that would be grand. Modulator, Drop-outs" in REMOTE not in LOCAL"
126589 | 42219.4393 |  | LI24 | 31 | 4 | HARDWARE | HW Problem | Klys 24-3 is tripping constantly, no more spares in L2. The machine wants downtime to start about a week early, and/or drive ops and PEM into a deep chasm of despair. Gaaah! Anyway, if PEM could investigate, that would be grand. Modulator, Drop-outs" in REMOTE not in LOCAL"
126589 | 42219.4393 | 1 | LI24 | 31 | 4 | HARDWARE | HW Problem | Klys 24-3 is tripping constantly, no more spares in L2. The machine wants downtime to start about a week early, and/or drive ops and PEM into a deep chasm of despair. Gaaah! Anyway, if PEM could investigate, that would be grand. Modulator, Drop-outs" in REMOTE not in LOCAL"
126589 | 42219.4393 | 1.2 | LI24 | 31 | 4 | HARDWARE | HW Problem | Klys 24-3 is tripping constantly, no more spares in L2. The machine wants downtime to start about a week early, and/or drive ops and PEM into a deep chasm of despair. Gaaah! Anyway, if PEM could investigate, that would be grand. Modulator, Drop-outs" in REMOTE not in LOCAL"
126589 | 42219.4393 | 0.5 | LI24 | 31 | 4 | HARDWARE | HW Problem | Klys 24-3 is tripping constantly, no more spares in L2. The machine wants downtime to start about a week early, and/or drive ops and PEM into a deep chasm of despair. Gaaah! Anyway, if PEM could investigate, that would be grand. Modulator, Drop-outs" in REMOTE not in LOCAL"
126589 | 42219.4393 | 3 | LI24 | 31 | 4 | HARDWARE | HW Problem | Klys 24-3 is tripping constantly, no more spares in L2. The machine wants downtime to start about a week early, and/or drive ops and PEM into a deep chasm of despair. Gaaah! Anyway, if PEM could investigate, that would be grand. Modulator, Drop-outs" in REMOTE not in LOCAL"
126589 | 42219.4393 | 0.3 | LI24 | 31 | 4 | HARDWARE | HW Problem | Klys 24-3 is tripping constantly, no more spares in L2. The machine wants downtime to start about a week early, and/or drive ops and PEM into a deep chasm of despair. Gaaah! Anyway, if PEM could investigate, that would be grand. Modulator, Drop-outs" in REMOTE not in LOCAL"
125382 | 42152.4605 |  | LI27 | 11 | 4 | HARDWARE | HW Problem | station 27-1 Mod Faulted
125534 | 42160.3989 |  | LI29 | 21 | 4 | HARDWARE | HW Problem | Klys 29-2 keeps MOD/HVR faulting We can sometimes reset it, but it quickly turns off again
125534 | 42160.3989 | 1 | LI29 | 21 | 4 | HARDWARE | HW Problem | Klys 29-2 keeps MOD/HVR faulting We can sometimes reset it, but it quickly turns off again
125534 | 42160.3989 | 1 | LI29 | 21 | 4 | HARDWARE | HW Problem | Klys 29-2 keeps MOD/HVR faulting We can sometimes reset it, but it quickly turns off again
125534 | 42160.3989 | 2 | LI29 | 21 | 4 | HARDWARE | HW Problem | Klys 29-2 keeps MOD/HVR faulting We can sometimes reset it, but it quickly turns off again
125812 | 42174.5908 | 1 | LI03 | 31 | 4 | HARDWARE | HW Problem | 3-3 PFN straps need to be replaced ( upper, Center jumper strap ) ( Lower, RHS, # 3 )
126220 | 42199.5968 | 2 | LI09 | 51 | 4 | HARDWARE | HW Problem | 9-5 Lug on Breaker is showing signs of overheating No work to start before Fiscal year 2016, Oct 1st, 2015
126253 | 42200.6068 | 1 | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 No REPK reading on PIOP display.
126301 | 42203.3098 | 1 | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 HVR no remote on
126320 | 42205.2176 | 0.5 | LI23 | 81 | 4 | HARDWARE | HW Problem | 23-8 FIAT rack:36-pins Amp connector that plug to PIOP module, C2 push-in
125181 | 42143.7637 | 1.5 | LI21 | 41 | 4 | HARDWARE | HW Problem | 21-4 T20 compartment fan not working
125479 | 42158.0924 | 1 | LI15 | 41 | 4 | HARDWARE | HW Problem | WGV fault on 15-4.
125799 | 42174.5585 | 2 | LI05 | 81 | 4 | HARDWARE | HW Problem | 5-8 EOLC stack needs to be repaced CB7 Plexiglass shield is broken K24 has a burnt lug Thyratron cables need to be secured.
125802 | 42174.5642 | 1 | LI05 | 31 | 4 | HARDWARE | HW Problem | 5-3 K26 contactor 'A' phase cable lug is burnt and loose
125806 | 42174.581 | 2 | LI04 | 41 | 4 | HARDWARE | HW Problem | 4-4 T20 Input snubber wire is badly burnt Mod 120VAC input filter at top of t20 comp has a burnt lug De- spiking coil had a broken resistor Charging diode stack needs to be replaced.
126052 | 42188.8017 |  | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 MOD faulted and won't reset remotely
126092 | 42192.2641 | 1 | LI15 | 21 | 4 | HARDWARE | HW Problem | 15-2 K26 and K27 Contacts need to be replaced. D Q ing Cap Wire/lug is badly burnt
126103 | 42192.4559 | 1 | LI28 | 31 | 4 | HARDWARE | HW Problem | 10 faults counts
126124 | 42193.5375 | 1 | LI16 | 31 | 4 | HARDWARE | HW Problem | 16-3 Mod EOLC stack needs to be replaced K27 Contacts need to be replaced.
126124 | 42193.5375 | 1 | LI16 | 31 | 4 | HARDWARE | HW Problem | 16-3 Mod EOLC stack needs to be replaced K27 Contacts need to be replaced.
126132 | 42193.6543 | 1.5 | LI27 | 21 | 4 | HARDWARE | HW Problem | 27-2: Thyratron filament curret at 95A. Jumper from 75A. See attached document.
126540 | 42216.2385 | 1 | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 VVS on PIOP"
126620 | 42220.9495 | 2 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3 HV not regulating. Lowered HV setpoint from 22.45 to 21.8 KV, HV stayed around 22KV Ask for station during low energy run
126620 | 42220.9495 |  | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3 HV not regulating. Lowered HV setpoint from 22.45 to 21.8 KV, HV stayed around 22KV Ask for station during low energy run
126744 | 42237.4597 | 1 | LI23 | 11 | 4 | HARDWARE | HW Problem | 23-1 PFN strap needs to be replaced (Lower, Center jumper )
126747 | 42237.4633 | 1.5 | LI24 | 21 | 4 | HARDWARE | HW Problem | Downtime: 24-2 Modulator, Has hardware problems. +1 pair of K26 burnt contacts +DE spiking coil has a broken resistor. +Thy comp D Q ing cap has a burnt lug
126748 | 42237.465 | 0.25 | LI24 | 31 | 4 | HARDWARE | HW Problem | Downtime: 24-3 Modulator has a broken DE spiking coil resistor.
124488 | 42116.1547 |  | LI27 | 11 | 4 | HARDWARE | HW Problem | 27-1 MOD fault and won't reset
124488 | 42116.1547 | 0.3 | LI27 | 11 | 4 | HARDWARE | HW Problem | 27-1 MOD fault and won't reset
124488 | 42116.1547 | 1.5 | LI27 | 11 | 4 | HARDWARE | HW Problem | 27-1 MOD fault and won't reset
124800 | 42129.564 | 1 | LI27 | 11 | 4 | HARDWARE | HW Problem | 27-1 Thy compartment D Q ing cap has a badly burnt wire/lug
124800 | 42129.564 | 1 | LI27 | 11 | 4 | HARDWARE | HW Problem | 27-1 Thy compartment D Q ing cap has a badly burnt wire/lug
124912 | 42132.3023 | 0.01 | LI03 | 71 | 4 | HARDWARE | HW Problem | 3-7 doesn't have any PAD phase readback, and shows several other PAD problems: PAD RF amplitude reads a nonsense number (-790%), PAD phase mode is 'LOST', Mixer gain is 0, and DAC status is 'FALSE'.
125413 | 42154.1904 | 0.5 | LI30 | 51 | 4 | HARDWARE | HW Problem | 30-5 HVR, frequent trips
125598 | 42163.8917 |  | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2 HVR faulting often
126353 | 42205.813 | 0.1 | LI25 | 61 | 4 | HARDWARE | HW Problem | 25-6 MOD fault, won't reset
126393 | 42207.9561 |  | LI25 | 71 | 4 | HARDWARE | HW Problem | Duplicate see 126395 . There are times when the amplitude is all over the place, BV stable and also the forward energy. The amplitude F_T_P is messed up. see elog entry 795177 for data.
126394 | 42208.004 | 3 | LI23 | 61 | 4 | HARDWARE | HW Problem | Klys 23-6 has been tripping relatively frequently over the past few days. Would be nice to have this adjusted (if possible) before the 12 keV run. Thanks!
126394 | 42208.004 | 1 | LI23 | 61 | 4 | HARDWARE | HW Problem | Klys 23-6 has been tripping relatively frequently over the past few days. Would be nice to have this adjusted (if possible) before the 12 keV run. Thanks!
126394 | 42208.004 | 1 | LI23 | 61 | 4 | HARDWARE | HW Problem | Klys 23-6 has been tripping relatively frequently over the past few days. Would be nice to have this adjusted (if possible) before the 12 keV run. Thanks!
126412 | 42209.2447 | 1.5 | LI25 | 71 | 4 | HARDWARE | HW Problem | 25-7 appears to have a bad PAD, Cable fault that clears when PAD substitution box is used.
126417 | 42209.4729 | 0.25 | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2 MOD HVR and won't reset remotely
126454 | 42212.5523 | 1 | LI22 | 11 | 4 | HARDWARE | HW Problem | Klys 22-1 VVS fault
124511 | 42116.6022 | 1 | LI02 | 31 | 4 | HARDWARE | HW Problem | 2-3 The Mod Thy Filament Heater Xformer has a burnt 120VAC lug that needs to be replaced.
124531 | 42117.5942 | 1 | LI11 | 81 | 4 | HARDWARE | HW Problem | 11-8 K24 relay is broken and needs to be replaced.
124747 | 42127.0116 | 0.75 | LI20 | 51 | 4 | HARDWARE | HW Problem | Klystron 20-5 TCAV0 mod fault.
125293 | 42147.917 |  | LI17 | 21 | 4 | HARDWARE | HW Problem | 17-2 SCRTD EOLC signal present not lit
125557 | 42162.4118 | 2 | LI25 | 11 | 4 | HARDWARE | HW Problem | 25-1 amplitude readback is very unstable, beam volts looks OK. Plots https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=786995
125827 | 42175.2496 | 5 | LI28 | 11 | 4 | HARDWARE | HW Problem | 28-1 CBL fault, doesn't clear upon IPLing
126037 | 42187.7375 | 1 | LI30 | 51 | 4 | HARDWARE | HW Problem | 30-5 tripped ~6 times in the first hour of swing shift 7/2/15.
126121 | 42193.4949 | 0.2 | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X (21-2) is MOD faulted, won't reset
125712 | 42170.6508 | 1 | LI25 | 11 | 4 | HARDWARE | HW Problem | LCLS: BV adjustment for 25-1, 29-2
125761 | 42172.7321 | 1 | LI26 | 11 | 4 | HARDWARE | HW Problem | 26-1 mod HV on but not firing
125831 | 42176.4782 | 0.5 | LI22 | 71 | 4 | HARDWARE | HW Problem | 22-7 Klystron temperature fault.
125831 | 42176.4782 | 2 | LI22 | 71 | 4 | HARDWARE | HW Problem | 22-7 Klystron temperature fault.
126448 | 42212.429 | 1 | LI30 | 51 | 4 | HARDWARE | HW Problem | 30-5 Check kly focus connections in kly shunt box and kly magnet breaker box
126515 | 42214.7426 | 2 | LI23 | 51 | 4 | HARDWARE | HW Problem | 23-5 dequing current reading is not changing as control pot is adjusted. BV is changing so dequing circuit seems to work OK.
124412 | 42111.3939 | 1 | LI21 | 21 | 4 | HARDWARE | HW Problem | Kly Mod 21-2 (L1X): PLC date and time are wrong. Fault log and system clock shows today's date as March 26, 2015 instead of correct date as April 17, 2015. Also, time is 3 minutes behind actual time.
124635 | 42122.5535 | 1 | LI29 | 31 | 4 | HARDWARE | HW Problem | 29-3: Seeing some beam volts jitter.
124637 | 42122.5916 |  | LI30 | 51 | 4 | HARDWARE | HW Problem | 30-5 Record KLYS:LI30:51:PIOP_MSGSUB: PIOP Processor Crashed. Reset required".
125327 | 42150.3625 | 0.8 | LI21 | 51 | 4 | HARDWARE | HW Problem | 21-5: BV needs to adjust to spec of 340KV.
125411 | 42153.823 | 1 | LI02 | 21 | 4 | HARDWARE | HW Problem | CBL fault on 2-2. IPL of piop didn't help.
125578 | 42163.5641 | 0.2 | LI29 | 31 | 4 | HARDWARE | HW Problem | 29-3 MKSU trigger Enable Fault The station is no longer faulting.
125595 | 42163.7109 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 FLK no reset
125605 | 42164.3563 | 4.5 | LI21 | 11 | 4 | HARDWARE | HW Problem | We do not see any evidence of this on the beam volts, however there is a strong correlation between 21-1 phase and jitter on the energy BPM in the BC2 chicane.
126603 | 42220.2708 | 1 | LI27 | 81 | 4 | HARDWARE | HW Problem | 27-8 Touch Panel needs to be replaced due to bad network card.
126603 | 42220.2708 | 0.5 | LI27 | 81 | 4 | HARDWARE | HW Problem | 27-8 Touch Panel needs to be replaced due to bad network card.
126711 | 42229.4253 | 1 | LI18 | 71 | 4 | HARDWARE | HW Problem | 18-7: KVGC normal light is off, power is on. Replace KVGC
126751 | 42237.5606 | 1 | LI26 | 81 | 4 | HARDWARE | HW Problem | 26-8 Mod has 2ea burnt carborundums that need to be replaced.
126752 | 42237.5619 | 1 | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2 Mod PFN strap needs to be replaced (Lower, RHS, #2 )
126753 | 42237.5631 | 0.75 | LI28 | 31 | 4 | HARDWARE | HW Problem | Downtime: 28-3 Modulator EOLC stack damaged.
126757 | 42237.5757 | 1 | LI29 | 31 | 4 | HARDWARE | HW Problem | 29-3 Mod need EOLC stack replaced (top rung < 5 Meg )
126759 | 42237.5817 | 1 | LI29 | 51 | 4 | HARDWARE | HW Problem | 29-5 Pulse cable needs to be replaced. HV rectifier cable to DC filter cap insulation is corroded.
126761 | 42237.5884 | 1 | LI30 | 31 | 4 | HARDWARE | HW Problem | 30-3 PFN strap will not tighten R/R ( Lower LHS #3 ) Thy viewing div cable lug burnt. Thy Comp D Qing cap Lug burnt
126768 | 42241.3502 |  | LI29 | 51 | 4 | HARDWARE | HW Problem | 29-5: Oil leak from the pulse cable connection at the klystron tank.
126773 | 42241.5446 | 0.5 | LI21 | 71 | 4 | HARDWARE | HW Problem | Downtime: 21-7 Modulator, Thy Door key broken in lock.
126774 | 42241.546 | 1 | LI21 | 51 | 4 | HARDWARE | HW Problem | 21-5 Mod PFN strap needs to be replaced (Lower, Center jum )
124446 | 42113.5973 | 0.75 | LI30 | 41 | 4 | HARDWARE | HW Problem | Klystron 30-4 RF Phase jitter. The jitter has a periodic motion on the FTP. Unsure of the cause.
124908 | 42132.2579 | 1 | LI27 | 81 | 4 | HARDWARE | HW Problem | 27-8 PFN caps needs replace 1. RHU 4th down. 2. RHL 4th up
124908 | 42132.2579 | 1.5 | LI27 | 81 | 4 | HARDWARE | HW Problem | 27-8 PFN caps needs replace 1. RHU 4th down. 2. RHL 4th up
126584 | 42219.4316 | 1 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 DQing current = 3.5A. Needs to adjust DQing while maintain BV to spec.
126584 | 42219.4316 | 0.5 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 DQing current = 3.5A. Needs to adjust DQing while maintain BV to spec.
126598 | 42219.5614 |  | LI29 | 41 | 4 | HARDWARE | HW Problem | 29-4 Replace IPA; power attenuator inside IPA may have deteriorated.
126648 | 42222.4411 | 0.5 | LI29 | 51 | 4 | HARDWARE | HW Problem | 29-5: Constant "AMJ" on the CUD. No BVJ seen.
126657 | 42222.5544 | 1 | LI24 | 81 | 4 | HARDWARE | HW Problem | 24-8 relocate deQ divider monitor location to thyratron anode to check for improved pfn stability.
126664 | 42222.93 | 2.5 | LI25 | 61 | 4 | HARDWARE | HW Problem | 25-6 mod fault, HV on - not firing
126664 | 42222.93 | 1 | LI25 | 61 | 4 | HARDWARE | HW Problem | 25-6 mod fault, HV on - not firing
126673 | 42224.1309 |  | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3 drops out repeatedly, but does not show any fault. See CATER # 126589
126673 | 42224.1309 | 2 | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3 drops out repeatedly, but does not show any fault. See CATER # 126589
126675 | 42225.1874 | 2 | LI27 | 31 | 4 | HARDWARE | HW Problem | 27-3 MOD fault, won't reset
124445 | 42113.348 | 0.2 | LI28 | 61 | 4 | HARDWARE | HW Problem | 28-6 HVR, no remote reset
124956 | 42132.875 | 0.1 | LI22 | 61 | 4 | HARDWARE | HW Problem | 22-6 MOD fault.
124976 | 42134.5293 | 2 | LI23 | 81 | 4 | HARDWARE | HW Problem | 23-8: Won't reset.
124976 | 42134.5293 | 1 | LI23 | 81 | 4 | HARDWARE | HW Problem | 23-8: Won't reset.
124976 | 42134.5293 | 1 | LI23 | 81 | 4 | HARDWARE | HW Problem | 23-8: Won't reset.
125469 | 42157.6731 | 1 | LI28 | 51 | 4 | HARDWARE | HW Problem | Klystron 28-5 mod fault
125484 | 42158.3421 |  | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3 Low beam voltage due to low DeQuing at 350kV.
126090 | 42192.2323 | 1 | LI28 | 41 | 4 | HARDWARE | HW Problem | noticed that the phase jitter has been creeping back up and now near old level of 1.5 I found an open circuit when ohming the PFN viewing divider and after reseating BNC connections, it cleared up for a while. suggest spraying contact cleaner on BNC connections and checking for intermittent open in the PFN viewing circuit.
126427 | 42210.3955 | 1 | LI24 | 31 | 4 | HARDWARE | HW Problem | Klystron 24-3 large RF phase oscillations On jitter FTPs. Not sure if real. Request someone to take a look.
126511 | 42214.6222 | 0.5 | LI28 | 21 | 4 | HARDWARE | HW Problem | Klys 28-2 HVR faulted and will not reset remotely
124361 | 42110.093 | 1 | LI03 | 21 | 4 | HARDWARE | HW Problem | 3-2 mod faulted and won't reset remotely
124854 | 42131.0166 | 1 | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 tripping repeatedly
124926 | 42132.4414 | 0.1 | LI27 | 41 | 4 | HARDWARE | HW Problem | Klys 27-4 HVR and won't reset remotely
125039 | 42137.913 | 2 | LI25 | 51 | 4 | HARDWARE | HW Problem | Kly. LI 25 - 5 tripping a lot. resets ok but not showing why it tripped.
125259 | 42146.3168 | 0.4 | LI21 | 21 | 4 | HARDWARE | HW Problem | L1-X (21-2) MOD fault, won't reset
125394 | 42152.7372 | 0.5 | LI12 | 41 | 4 | HARDWARE | HW Problem | Klystron 12-4 mod fault.
125562 | 42163.1852 |  | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 has MOD faulted many times this shift
125781 | 42173.6577 | 2 | LI21 | 31 | 4 | HARDWARE | HW Problem | 21-3: Replace Broken Control Arm in SLED.
126032 | 42187.5849 | 0.75 | LI11 | 11 | 4 | HARDWARE | HW Problem | 11-1 K27 Contacts need to be replaced
126138 | 42194.1613 | 1 | LI30 | 51 | 4 | HARDWARE | HW Problem | 30-5 FLK, no remote reset
126138 | 42194.1613 | 2 | LI30 | 51 | 4 | HARDWARE | HW Problem | 30-5 FLK, no remote reset
126221 | 42199.598 | 2 | LI09 | 61 | 4 | HARDWARE | HW Problem | 9-6 KMB breaker Lug showing signs of bad over heating No work to start before Fiscal year 2016, Oct 1st, 2015
126274 | 42201.3178 | 0.5 | LI29 | 81 | 4 | HARDWARE | HW Problem | 29-8 Bad high powered attenuator, replace IPA then resaturate station.
126345 | 42205.7143 | 2 | LI29 | 51 | 4 | HARDWARE | HW Problem | 29-5 has mod faulted and will not reset remotely.
126478 | 42213.5129 |  | LI18 | 71 | 4 | HARDWARE | HW Problem | 18-7: "KLV" fault on the MKSU/PIOP data
126478 | 42213.5129 | 1 | LI18 | 71 | 4 | HARDWARE | HW Problem | 18-7: "KLV" fault on the MKSU/PIOP data
126491 | 42213.8399 | 6 | LI22 | 81 | 4 | HARDWARE | HW Problem | 22-8 trip on ACOC at turn on. Has bad T20
126617 | 42220.7932 | 1 | LI24 | 41 | 4 | HARDWARE | HW Problem | Downtime 24-4 has a EOLC 20 ohm carborundum reststor burnt
124828 | 42130.4167 | 8 | LI28 | 21 | 4 | HARDWARE | HW Problem | Kly Mod 28-2: TS/R HVOCs and backswing spikes. When HV reaches around 18KV, get hard HVOCs. EOLC current has severe spikes throughout entire backswing when this occurs. Thyratron maybe on the verge of going bad as well.
124828 | 42130.4167 | 1 | LI28 | 21 | 4 | HARDWARE | HW Problem | Kly Mod 28-2: TS/R HVOCs and backswing spikes. When HV reaches around 18KV, get hard HVOCs. EOLC current has severe spikes throughout entire backswing when this occurs. Thyratron maybe on the verge of going bad as well.
124913 | 42132.3462 |  | LI20 | 41 | 4 | HARDWARE | HW Problem | Mod fault on facet xtcav (20-4b). Can't reset remotely. EDIT: gets power from VVS10. Probably fine.
124953 | 42132.7494 | 0.75 | LI25 | 31 | 4 | HARDWARE | HW Problem | FTPs look OK right now, but station contributes greatly to beam jitter, so would like PEM to investigate.
125025 | 42136.5954 | 0.75 | LI05 | 51 | 4 | HARDWARE | HW Problem | 5-5 Alarm latch LED is burnt out
125087 | 42138.7929 | 1 | LI10 | 41 | 4 | HARDWARE | HW Problem | Replace Cassel box. HV off leaking electrons which biases K7 so it's vibration sensitive. NOTE!! Reconnect yellow wire to TB2-19.
125124 | 42140.2048 | 1 | LI24 | 41 | 4 | HARDWARE | HW Problem | FTP shows beam volts jitter of about 1kV
125172 | 42143.5582 | 0.75 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4: Seeing some beam volts jitter.
125843 | 42177.3114 |  | DR13 | 1 | 4 | HARDWARE | HW Problem | 1-8 K27 Contacts need to be replaced. Fire alarm latch LED out PFN strap ( Upper, RHS, #2 ) ELP for station Missing
125899 | 42179.9608 | 1 | LI21 | 21 | 4 | HARDWARE | HW Problem | Seems like we have a big L1X phase jump, and L1S follows it and screws up the BC1 energy/bunch length. We have to restore all actuators, and somehow all the feedback time slots get screwed up and we have to zero offsets.
126025 | 42187.5743 | 0.75 | LI13 | 11 | 4 | HARDWARE | HW Problem | 13-1 D Q ing cap wire/lug badly burnt PFN strap bad ( Lower, RHS, # 2)
126026 | 42187.5755 | 1 | LI12 | 71 | 4 | HARDWARE | HW Problem | 12-7 K 27 contacts need to be replaced
126028 | 42187.5781 | 1 | LI12 | 31 | 4 | HARDWARE | HW Problem | 12-3 K27 contacts need to be replaced
126107 | 42192.4926 | 1 | LI06 | 31 | 4 | HARDWARE | HW Problem | 6-3 waveguide vacuum fault
124559 | 42118.6133 | 2 | LI29 | 21 | 4 | HARDWARE | HW Problem | PEM looked at the station locally and the modulator is working fine. They are going to try swapping the MKSU.
124744 | 42126.5467 | 1.5 | LI25 | 71 | 4 | HARDWARE | HW Problem | Error report shows 39 faults since 08:00--I suspect most of them happened in the last hour or so.
124757 | 42127.658 |  | LI26 | 61 | 4 | HARDWARE | HW Problem | 26-6: 1. Need four AC filters. 2. Swap hot vs neutral wires.
125847 | 42177.3288 | 4 | LI00 | 11 | 4 | HARDWARE | HW Problem | 01-1A The D Q ing resistor assem on top of mod needs to be replaced as all the connection links are badly burnt .
125729 | 42171.5361 | 0.75 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3: RF AMPL low OOT due to beam volts low, DQing=19 amps.
125735 | 42171.5904 | 2 | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2 remove Ross divider from thyratron anode
125735 | 42171.5904 | 1.5 | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2 remove Ross divider from thyratron anode
125840 | 42177.3045 | 1 | LI02 | 41 | 4 | HARDWARE | HW Problem | 2-4 Core bias amp meter and lug need to be replaced. Thyratron cables need to be secured from blower motor
126007 | 42186.5647 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | Noticed a phase change in L1-X (21-2) It appears like the temperature stepped down when the beam was off during the POMM. During 9 am to 11:45 am. DLWG body temperature fluctuations.
126046 | 42188.5897 | 1 | LI14 | 21 | 4 | HARDWARE | HW Problem | 14-2 Mod Fire alarm latching unit LED is burnt out.
126104 | 42192.478 | 0.5 | LI02 | 61 | 4 | HARDWARE | HW Problem | 2-6 showing klystron vacuum fault.
126104 | 42192.478 | 0.2 | LI02 | 61 | 4 | HARDWARE | HW Problem | 2-6 showing klystron vacuum fault.
126111 | 42192.4974 | 1 | LI17 | 41 | 4 | HARDWARE | HW Problem | 17-4 waveguide vacuum fault
126155 | 42195.4696 | 2 | LI28 | 71 | 4 | HARDWARE | HW Problem | 28-7: HVR, no remote reset. Needs to check Enerpro board.
124437 | 42112.0237 | 1 | LI22 | 11 | 4 | HARDWARE | HW Problem | 22-1 mod faulted and won't reset remotely
124437 | 42112.0237 | 1.5 | LI22 | 11 | 4 | HARDWARE | HW Problem | 22-1 mod faulted and won't reset remotely
124437 | 42112.0237 | 2 | LI22 | 11 | 4 | HARDWARE | HW Problem | 22-1 mod faulted and won't reset remotely
124437 | 42112.0237 | 6 | LI22 | 11 | 4 | HARDWARE | HW Problem | 22-1 mod faulted and won't reset remotely
124437 | 42112.0237 | 2 | LI22 | 11 | 4 | HARDWARE | HW Problem | 22-1 mod faulted and won't reset remotely
124680 | 42124.3257 | 0.5 | LI13 | 21 | 4 | HARDWARE | HW Problem | Klystron LI13-21 has bouts of high amplitude instability.
124776 | 42128.6727 | 0.5 | LI22 | 81 | 4 | HARDWARE | HW Problem | station 22-8 mod faulted
124797 | 42129.5596 | 1 | LI26 | 11 | 4 | HARDWARE | HW Problem | 26-1 Mod EOLC stack is severely burn and needs to be ASAP.
125113 | 42139.627 | 0.8 | LI26 | 21 | 4 | HARDWARE | HW Problem | Klystron 26-2 large beam volt jitter. As seen on the FTPs.
125113 | 42139.627 | 0.5 | LI26 | 21 | 4 | HARDWARE | HW Problem | Klystron 26-2 large beam volt jitter. As seen on the FTPs.
125113 | 42139.627 | 1 | LI26 | 21 | 4 | HARDWARE | HW Problem | Klystron 26-2 large beam volt jitter. As seen on the FTPs.
125600 | 42163.9407 | 0.5 | LI22 | 81 | 4 | HARDWARE | HW Problem | 22-8 MOD fault, won't reset
124462 | 42114.8265 | 1.5 | LI26 | 11 | 4 | HARDWARE | HW Problem | 26-1 FLK no reset also hint of VESDA no human perceivable smoke.
124573 | 42119.3795 | 3.5 | LI24 | 41 | 4 | HARDWARE | HW Problem | Klys 24-4 trips off on simultaneous HVOC and ACOC faults when HV above 16-18KV.
124573 | 42119.3795 | 3 | LI24 | 41 | 4 | HARDWARE | HW Problem | Klys 24-4 trips off on simultaneous HVOC and ACOC faults when HV above 16-18KV.
124573 | 42119.3795 | 0.5 | LI24 | 41 | 4 | HARDWARE | HW Problem | Klys 24-4 trips off on simultaneous HVOC and ACOC faults when HV above 16-18KV.
124808 | 42129.575 | 0.5 | LI30 | 61 | 4 | HARDWARE | HW Problem | 30-6 Mod PFN strap needs to be replaced (lower, lhs, #3)
124809 | 42129.5761 | 1 | LI30 | 81 | 4 | HARDWARE | HW Problem | 30-8 Mod EOLC stack needs to be replaced.
125009 | 42136.2871 | 2 | LI23 | 81 | 4 | HARDWARE | HW Problem | 23-8 MOD fault
125009 | 42136.2871 | 1 | LI23 | 81 | 4 | HARDWARE | HW Problem | 23-8 MOD fault
125009 | 42136.2871 | 1 | LI23 | 81 | 4 | HARDWARE | HW Problem | 23-8 MOD fault
125009 | 42136.2871 | 6 | LI23 | 81 | 4 | HARDWARE | HW Problem | 23-8 MOD fault
125009 | 42136.2871 |  | LI23 | 81 | 4 | HARDWARE | HW Problem | 23-8 MOD fault
125199 | 42144.6037 | 0 | LI30 | 61 | 4 | HARDWARE | HW Problem | 30-6: Station has excessive BV dropouts. Looks like HV oscillating.
125599 | 42163.9311 | 2 | LI28 | 41 | 4 | HARDWARE | HW Problem | The klystron Magnet breaker tripped. Turned KCS water up a 1/4 turn. Float already at the top. Ran for some time and tripped again.
125599 | 42163.9311 |  | LI28 | 41 | 4 | HARDWARE | HW Problem | The klystron Magnet breaker tripped. Turned KCS water up a 1/4 turn. Float already at the top. Ran for some time and tripped again.
125603 | 42164.0979 | 1 | LI28 | 11 | 4 | HARDWARE | HW Problem | Klys 28-1 Tripping off the beam frequently. Goes MOD and resets but happens every few minutes.
126255 | 42200.6405 | 1 | LI28 | 71 | 4 | HARDWARE | HW Problem | Jumped up from 114 to 115.8 when modulator beam volts was raised on 7/9. Flow to accelerator section should be increased if there is room to adjust.
126255 | 42200.6405 | 0 | LI28 | 71 | 4 | HARDWARE | HW Problem | Jumped up from 114 to 115.8 when modulator beam volts was raised on 7/9. Flow to accelerator section should be increased if there is room to adjust.
126375 | 42206.5925 | 1 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6: causing FEL jitter Duplicate, See 127376
126396 | 42208.1835 | 2 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4 has intermittent phase, amplitude & beam volt jitter that seems to go away when o'scope is connected to PFN monitoring jack
126396 | 42208.1835 | 1 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4 has intermittent phase, amplitude & beam volt jitter that seems to go away when o'scope is connected to PFN monitoring jack
131085 | 42455.103 |  | DR03 | 1 | 4 | HARDWARE | HW Problem | Time accounting - DR03 compressor saw a phase jump, which railed the SREX feedback. It seems like the phase was set manually... No trouble found.
124405 | 42111.0843 | 1 | LI03 | 51 | 4 | HARDWARE | HW Problem | 3-5 Mod Flt will not remote reset
124405 | 42111.0843 | 1 | LI03 | 51 | 4 | HARDWARE | HW Problem | 3-5 Mod Flt will not remote reset
124510 | 42116.6016 | 1 | LI05 | 61 | 4 | HARDWARE | HW Problem | 5-6 1) Thy comp fan rubbing, noisy 2) Wire to thy for DQ meter res needs retermination 3) T20 Network res lugs heated and cracking
124984 | 42135.2659 | 7 | LI00 | 2 | 4 | HARDWARE | HW Problem | 0-2 MOD fault, won't reset.
125023 | 42136.5567 |  | LI23 | 71 | 4 | HARDWARE | HW Problem | Klystron Li23-71 is tripped off and does not reset remotely.
125153 | 42142.8385 | 2 | LI18 | 11 | 4 | HARDWARE | HW Problem | Station had 5 trips on Klystron Vacuum
124499 | 42116.5096 |  | LI30 | 51 | 4 | HARDWARE | HW Problem | 30-5 showing many AMPL dropouts. CEFOI check WG2 switch. See CATER 124056, now closed, for past work on this station. pms
125813 | 42174.5923 | 1 | LI02 | 71 | 4 | HARDWARE | HW Problem | 2-7 Charging diode stack needs to be replaced K27 contacts need to be replaced
125813 | 42174.5923 | 1 | LI02 | 71 | 4 | HARDWARE | HW Problem | 2-7 Charging diode stack needs to be replaced K27 contacts need to be replaced
125832 | 42176.6325 | 0.1 | LI23 | 41 | 4 | HARDWARE | HW Problem | 23-4 phase & amplitude readback broken. Does appear to be contributing energy to the beam. Lots of concurrent PIOP processor crashes with external interlock faults on KMC. Bad PIOP? See 125816.
126051 | 42188.7699 | 3 | LI25 | 21 | 4 | HARDWARE | HW Problem | 25-2 needs PFN tune, slope -.8
126162 | 42195.599 | 1 | LI19 | 31 | 4 | HARDWARE | HW Problem | 19-3 K26 and K27 contacts need to be replaced, ATTN: This job requires a RFWCF before doing work.. Please see Jim Craft for Form.
126228 | 42199.8755 | 1 | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 trips CB4 after about half hour.
126229 | 42199.877 | 2 | LI28 | 81 | 4 | HARDWARE | HW Problem | 28-8 needs PFN tune after cap replacement, -.5 slope.
126437 | 42211.5628 | 0.5 | LI22 | 71 | 4 | HARDWARE | HW Problem | 22-7 mod fault does not reset
126438 | 42211.5865 | 0.5 | LI29 | 11 | 4 | HARDWARE | HW Problem | Klystron 29-1 KMC fault
124403 | 42110.9449 | 4 | LI29 | 51 | 4 | HARDWARE | HW Problem | See FTP in MCC ELOG entry https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=776787
124403 | 42110.9449 | 0.8 | LI29 | 51 | 4 | HARDWARE | HW Problem | See FTP in MCC ELOG entry https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=776787
124407 | 42111.1243 | 1 | LI16 | 51 | 4 | HARDWARE | HW Problem | 16-5 MOD fault, no remote reset
124633 | 42122.5488 | 0.5 | LI22 | 51 | 4 | HARDWARE | HW Problem | 22-5: DQ is measured at 32A. Needs to adjust T20 tap to lower it down
124880 | 42131.643 | 2.5 | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2 Needs new snubber board (phase C); AC filters; hot vs neutral wires need to be swapped.
124892 | 42131.7544 |  | LI18 | 11 | 4 | HARDWARE | HW Problem | 18-1: Thyratron not firing. HV on ok.
124892 | 42131.7544 | 1 | LI18 | 11 | 4 | HARDWARE | HW Problem | 18-1: Thyratron not firing. HV on ok.
126794 | 42243.3602 | 1 | LI08 | 61 | 4 | HARDWARE | HW Problem | 8-6 Disconnect T20 , HV rectifier stack, disconnect Snubber resistors, and fire wire.
126794 | 42243.3602 | 3 | LI08 | 61 | 4 | HARDWARE | HW Problem | 8-6 Disconnect T20 , HV rectifier stack, disconnect Snubber resistors, and fire wire.
126795 | 42243.3636 | 1 | LI08 | 81 | 4 | HARDWARE | HW Problem | 8-8 Disconnect T20 , HV rectifier stack, disconnect Snubber resistors, and fire wire.
126795 | 42243.3636 | 4 | LI08 | 81 | 4 | HARDWARE | HW Problem | 8-8 Disconnect T20 , HV rectifier stack, disconnect Snubber resistors, and fire wire.
126797 | 42243.6145 | 1 | LI22 | 11 | 4 | HARDWARE | HW Problem | 22-1 PIOP indicates KLV. Pressure looks good.
126803 | 42245.8228 | 1.5 | LI21 | 31 | 4 | HARDWARE | HW Problem | 21-3 klystron magnet supply dead at KG21-70-08
126871 | 42255.307 | 1 | LI26 | 41 | 4 | HARDWARE | HW Problem | 26-4 Modulator, LVPS fault won't clear.
124463 | 42114.8608 | 1 | LI21 | 61 | 4 | HARDWARE | HW Problem | Klystron 21-6 phase jitter. Checking the FTP reveals noticeable phase jitter.
124803 | 42129.568 | 0.5 | LI29 | 21 | 4 | HARDWARE | HW Problem | 29-2 Mod de-spiking coil has a broken resistor.
124857 | 42131.2021 | 0.5 | LI27 | 11 | 4 | HARDWARE | HW Problem | 27-1 MOD fault,won't reset
125045 | 42138.2535 | 1 | LI27 | 11 | 4 | HARDWARE | HW Problem | 27-1 Mod trips on HVOC
125067 | 42138.489 | 0.5 | LI13 | 31 | 4 | HARDWARE | HW Problem | Klystron 13-3 mod fault.
125128 | 42140.7403 | 0.75 | LI17 | 71 | 4 | HARDWARE | HW Problem | Klystron 17-7 MOD faulted and won't reset remotely
125134 | 42141.6142 | 1 | LI05 | 21 | 4 | HARDWARE | HW Problem | 5-2 FLK
125134 | 42141.6142 | 1 | LI05 | 21 | 4 | HARDWARE | HW Problem | 5-2 FLK
125160 | 42143.3015 | 1 | LI20 | 81 | 4 | HARDWARE | HW Problem | 20-8 Mod flt
125305 | 42148.5672 | 2 | LI30 | 11 | 4 | HARDWARE | HW Problem | Station 30-1 is cycling very frequently.
125305 | 42148.5672 | 1 | LI30 | 11 | 4 | HARDWARE | HW Problem | Station 30-1 is cycling very frequently.
125800 | 42174.5605 | 0.75 | LI05 | 71 | 4 | HARDWARE | HW Problem | 5-7 HV dc cable link in the PFN compartment is missing a 1/4 20 bolt.
126084 | 42191.6697 | 1 | LI20 | 71 | 4 | HARDWARE | HW Problem | 20-7: WVGC Low Trip Point Setting Erractic (NIRP test).
126206 | 42199.3279 | 0.2 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 SLED needs tuning due to diurnal amplitude changes.
124805 | 42129.5706 | 1 | LI29 | 61 | 4 | HARDWARE | HW Problem | 29-6 Mod charging diode stack needs to be replaced.
125083 | 42138.7646 | 0.5 | LI10 | 71 | 4 | HARDWARE | HW Problem | Informed PEM.
125362 | 42151.5953 | 2 | LI22 | 81 | 4 | HARDWARE | HW Problem | 22-8 Mod T20 Network resistor is badly burnt Grid trigger chassis is broken and needs to be R/R
125425 | 42155.676 | 1 | LI23 | 31 | 4 | HARDWARE | HW Problem | Klystron 23-3 mod fault.
125426 | 42155.7152 | 0.75 | LI12 | 21 | 4 | HARDWARE | HW Problem | Klystron Li12-21 is tripped on WGV fault. Does not reset remotely. Vacuum in area seems fine. Vacuum burst earlier tripped off 11 and 31 also, but they recovered by themselves.
125440 | 42156.9678 | 0.8 | LI21 | 61 | 4 | HARDWARE | HW Problem | Klystron 21-6 beam volts dropouts. We are seeing it on the FTP beam volts/phase jitter.
125440 | 42156.9678 | 1 | LI21 | 61 | 4 | HARDWARE | HW Problem | Klystron 21-6 beam volts dropouts. We are seeing it on the FTP beam volts/phase jitter.
125457 | 42157.5124 | 0.75 | LI27 | 41 | 4 | HARDWARE | HW Problem | 27-4: DQing Undercurrent alarm on EPICS.
125776 | 42173.5891 | 0.1 | LI28 | 51 | 4 | HARDWARE | HW Problem | Klystron 28-5 mod fault. BVLT is up but no AMPL
125807 | 42174.5825 | 1 | LI04 | 41 | 4 | HARDWARE | HW Problem | 4-2 K26 and K27 contacts need to be replaced De-spiking coil resistor is broken
125849 | 42177.3395 | 1 | DR13 | 1 | 4 | HARDWARE | HW Problem | 1-8 K27 contacts need to be replaced, RSWCF needed to complete this job, Please see Jim Craft for form. PFN straps ( up, RHS, #2 ) (U, LHS, # 3 ) Fire alarm LED burnt out ELP Missing
125849 | 42177.3395 | 1 | DR13 | 1 | 4 | HARDWARE | HW Problem | 1-8 K27 contacts need to be replaced, RSWCF needed to complete this job, Please see Jim Craft for form. PFN straps ( up, RHS, #2 ) (U, LHS, # 3 ) Fire alarm LED burnt out ELP Missing
125849 | 42177.3395 | 1 | DR13 | 1 | 4 | HARDWARE | HW Problem | 1-8 K27 contacts need to be replaced, RSWCF needed to complete this job, Please see Jim Craft for form. PFN straps ( up, RHS, #2 ) (U, LHS, # 3 ) Fire alarm LED burnt out ELP Missing
125925 | 42181.2344 | 1 | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2 HVR would not rmt HV ON
125993 | 42186.3644 | 2 | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X jitter is worse since last week
126031 | 42187.5833 | 1 | LI11 | 21 | 4 | HARDWARE | HW Problem | 11-2 K27 Contacts need to be replaced. PFN strap needs to be replaced.( Lower, LHS, # 3 )
126036 | 42187.6444 | 0.1 | LI22 | 71 | 4 | HARDWARE | HW Problem | 22-7 tipped "no reset"
124344 | 42109.4465 |  | LI24 | 41 | 4 | HARDWARE | HW Problem | station 24-4 mod faulted 4 times in ~10 minutes
124798 | 42129.562 | 0.5 | LI26 | 71 | 4 | HARDWARE | HW Problem | 26-7 Mod K28 Bleeder resistors are disconnected and need to be repaired asap.
125112 | 42139.6237 |  | LI25 | 11 | 4 | HARDWARE | HW Problem | Klystron 25-1 bad ampl/phase jitter as seen on the FTPs. Apparent on the beam volt jitter as well.
125234 | 42145.5366 | 1 | LI02 | 61 | 4 | HARDWARE | HW Problem | Downtime; 2-6: Seeing some beam volts jitter.
125307 | 42148.6473 |  | LI30 | 11 | 4 | HARDWARE | HW Problem | 30-1: Vacuum glitches causing station to drop out.
125429 | 42156.3647 | 2 | LI18 | 11 | 4 | HARDWARE | HW Problem | Downtime; 18 -1 Replace Thyratron
125429 | 42156.3647 | 2 | LI18 | 11 | 4 | HARDWARE | HW Problem | Downtime; 18 -1 Replace Thyratron
126027 | 42187.5771 | 0.75 | LI12 | 41 | 4 | HARDWARE | HW Problem | 12-4 has a broken despiking coil resistor
126333 | 42205.4862 | 1 | LI30 | 61 | 4 | HARDWARE | HW Problem | 30-6: Cable connections to J101 & J102 on the HVPS regulator board were swapped in order to have normal configuration. Needs to TS&R to correct the problem.
126364 | 42206.4105 | 0.5 | LI30 | 51 | 4 | HARDWARE | HW Problem | 30-5 tripping repeatedly on HVR and MKT faults
126408 | 42208.5278 | 1 | LI30 | 61 | 4 | HARDWARE | HW Problem | 30-6: Constant RE's, won't stay on.
126441 | 42211.892 | 0.5 | LI24 | 61 | 4 | HARDWARE | HW Problem | Klys 24-6 HVR, will not come on remotely.
124699 | 42124.594 | 2.5 | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X (21-2) Beam Volts dropping out
125519 | 42159.5437 | 1 | LI09 | 71 | 4 | HARDWARE | HW Problem | 9-7: "External Fault". See 125414. Replace the WG VGC.
125519 | 42159.5437 | 1 | LI09 | 71 | 4 | HARDWARE | HW Problem | 9-7: "External Fault". See 125414. Replace the WG VGC.
125679 | 42167.5352 | 0.5 | LI22 | 81 | 4 | HARDWARE | HW Problem | 22-8 MOD fault, no remote reset
125679 | 42167.5352 | 0.5 | LI22 | 81 | 4 | HARDWARE | HW Problem | 22-8 MOD fault, no remote reset
125767 | 42173.3569 | 0 | LI23 | 11 | 4 | HARDWARE | HW Problem | Does not seem to happen at standby.
126043 | 42188.5273 |  | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 showing CAM error, swapped the PIOP, the error switched to CBL until we cycled the crate for a 28 subooster problem. When the crate came back online 28-3 is once again showing a CAM error even with the new PIOP, so either the replacement PIOP is also bad or something else is the problem.
126045 | 42188.588 |  | LI13 | 11 | 4 | HARDWARE | HW Problem | 13-1 Mod D Q ing Cap wire in thyratron comp are badly burnt.
126048 | 42188.5922 | 0.75 | LI14 | 61 | 4 | HARDWARE | HW Problem | 14-6 Mod K27 Contacts Need to be replaced
126073 | 42191.4824 | 0.5 | LI22 | 51 | 4 | HARDWARE | HW Problem | 22-5 MOD fault.
126128 | 42193.5444 | 1 | LI16 | 71 | 4 | HARDWARE | HW Problem | 16-7 PFN strap needs to be replaced (Lower, RHS, 3 )
126140 | 42194.5426 | 1.5 | LI23 | 71 | 4 | HARDWARE | HW Problem | Klys 23-7 has no phase/ampl readback. IPL didn't help. PAD or PIOP?
126174 | 42198.2559 | 2 | LI30 | 81 | 4 | HARDWARE | HW Problem | They will be upgrading various boards and installing new software for a PLC test panel. They are doing this to alleviate some of the frequent MOD faults this station has been having.
126174 | 42198.2559 |  | LI30 | 81 | 4 | HARDWARE | HW Problem | They will be upgrading various boards and installing new software for a PLC test panel. They are doing this to alleviate some of the frequent MOD faults this station has been having.
126176 | 42198.273 | 0.5 | LI29 | 11 | 4 | HARDWARE | HW Problem | klys 29-1 WGV fault
126222 | 42199.5995 | 1.5 | LI04 | 41 | 4 | HARDWARE | HW Problem | 4-4 KMB breaker Lug shows signs of overheating No work to start before Fiscal year 2016, Oct 1st, 2015
126243 | 42200.3482 |  | LI28 | 41 | 4 | HARDWARE | HW Problem | Klys 28-4 showing a sinusoidal oscillation in beam volts. Not sure if this is real or a readback issue.
126252 | 42200.6033 | 1.5 | LI23 | 11 | 4 | HARDWARE | HW Problem | 23-1 raise beam voltage to 350 kV after modulator repairs
126256 | 42200.6528 |  | LI28 | 51 | 4 | HARDWARE | HW Problem | 28-5: Investigate wide PW on PFN
124452 | 42114.2646 | 3 | LI28 | 81 | 4 | HARDWARE | HW Problem | 28-8 needs PFN tune: PW=0.00uS, Slope= -1.1; Ripl = 0.2
124515 | 42116.6053 | 2.5 | LI20 | 71 | 4 | HARDWARE | HW Problem | Won't reset remotely.
124534 | 42117.6016 | 1 | LI09 | 11 | 4 | HARDWARE | HW Problem | 9-1 Mod Alarm latching LED needs to be replaced.
125319 | 42150.0325 | 2 | LI21 | 81 | 4 | HARDWARE | HW Problem | Large phase and amplitude jitter
125352 | 42151.5761 | 1 | LI25 | 11 | 4 | HARDWARE | HW Problem | 25-1 1 pair of K26 and 1 pair of K27 contacts need to be R/R
125474 | 42157.7422 | 0.75 | LI05 | 81 | 4 | HARDWARE | HW Problem | Klystron 5-8 mod fault.
125476 | 42157.8386 | 0.75 | LI10 | 71 | 4 | HARDWARE | HW Problem | Klystron LI10-71 is tripped on Mod Fault and will not reset remotely.
125485 | 42158.3451 | 1 | LI29 | 71 | 4 | HARDWARE | HW Problem | 29-7 BV was lowered due to reported ELOC, KLV and HVOC. Needs investigation.
125844 | 42177.3161 | 2 | LI01 | 61 | 4 | HARDWARE | HW Problem | 1-6 EOLC stack needs to be replaced. T20 Jumper cable worn and needs to be replaced. De-spiking coil resistor needs to be replaced PFN coil needs to be replaced insulator missing (lower, LHS, #3 Thyratron cables need to be secured from blower motor
125844 | 42177.3161 | 1 | LI01 | 61 | 4 | HARDWARE | HW Problem | 1-6 EOLC stack needs to be replaced. T20 Jumper cable worn and needs to be replaced. De-spiking coil resistor needs to be replaced PFN coil needs to be replaced insulator missing (lower, LHS, #3 Thyratron cables need to be secured from blower motor
126049 | 42188.5949 | 1 | LI14 | 81 | 4 | HARDWARE | HW Problem | 14-8 Mod PFN straps Need to be replaced. (upper, RHS, # 3) (Lower, RHS, # 2) ( Lower, Center )
126133 | 42193.657 | 2 | LI26 | 51 | 4 | HARDWARE | HW Problem | 26-5: DQ'g current unstable. See attached document.
126133 | 42193.657 | 0.3 | LI26 | 51 | 4 | HARDWARE | HW Problem | 26-5: DQ'g current unstable. See attached document.
126218 | 42199.5935 | 1 | LI10 | 71 | 4 | HARDWARE | HW Problem | 10-7 KMB heat sink is hanging loose due to broken nylon Screw. No work to start before Fiscal year 2016, Oct 1st, 2015
126410 | 42208.8188 |  | LI23 | 21 | 4 | HARDWARE | HW Problem | 23-2 MOD faulting repeatedly on reflected power during the day.
126410 | 42208.8188 | 0.5 | LI23 | 21 | 4 | HARDWARE | HW Problem | 23-2 MOD faulting repeatedly on reflected power during the day.
125360 | 42151.5906 | 1 | LI23 | 21 | 4 | HARDWARE | HW Problem | 23-2: 1. Mod has a EOLC stack that needs to be replaced 2. BV is low (321KV). Needs to adjust it to 350KV.
125360 | 42151.5906 | 1 | LI23 | 21 | 4 | HARDWARE | HW Problem | 23-2: 1. Mod has a EOLC stack that needs to be replaced 2. BV is low (321KV). Needs to adjust it to 350KV.
125583 | 42163.582 | 1 | LI21 | 11 | 4 | HARDWARE | HW Problem | L1-S HVR, no remote reset
125583 | 42163.582 | 0.5 | LI21 | 11 | 4 | HARDWARE | HW Problem | L1-S HVR, no remote reset
125591 | 42163.6975 | 3.5 | LI21 | 11 | 4 | HARDWARE | HW Problem | PEM dispatched.
125591 | 42163.6975 |  | LI21 | 11 | 4 | HARDWARE | HW Problem | PEM dispatched.
125007 | 42136.2127 | 2 | LI27 | 21 | 4 | HARDWARE | HW Problem | 27-2 HVR fault
125111 | 42139.5025 | 0.5 | LI17 | 31 | 4 | HARDWARE | HW Problem | Klystron 17-3 mod fault
125390 | 42152.6202 | 0.5 | LI20 | 81 | 4 | HARDWARE | HW Problem | L0B phase is jittering by +/- 3 degrees. The modulator beam voltage seems stable while the phase jitter is happening.
125582 | 42163.5793 | 1 | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X phase jumps from -160 degrees to -20 degrees one or two times per second, not in any regular pattern.
125582 | 42163.5793 | 0.5 | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X phase jumps from -160 degrees to -20 degrees one or two times per second, not in any regular pattern.
125582 | 42163.5793 | 1.5 | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X phase jumps from -160 degrees to -20 degrees one or two times per second, not in any regular pattern.
125648 | 42166.177 | 2 | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X tripped, no reset.
125431 | 42156.4386 | 0.25 | DR13 | 1 | 4 | HARDWARE | HW Problem | Won't reset. Informed PEM.
125499 | 42158.4594 | 7 | LI00 | 98 | 4 | HARDWARE | HW Problem | SHB 2 PAD seems to be dead.
125510 | 42159.251 | 0.5 | LI23 | 41 | 4 | HARDWARE | HW Problem | 23-4 keeps tripping off (several times an hour). Took off the beam as we can live without it.
125550 | 42161.0979 | 1 | LI03 | 71 | 4 | HARDWARE | HW Problem | Klys 3-7 MOD faulted and will not reset remotely
125588 | 42163.6579 | 0.5 | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X is tripped on modulator fault lockout.
125657 | 42166.4446 | 0.5 | LI26 | 51 | 4 | HARDWARE | HW Problem | 26-5: DQing current low OOT at less than nominal beam volts.
125657 | 42166.4446 | 1 | LI26 | 51 | 4 | HARDWARE | HW Problem | 26-5: DQing current low OOT at less than nominal beam volts.
125666 | 42166.5554 | 0.75 | LI26 | 61 | 4 | HARDWARE | HW Problem | 26-8: Seeing some beam volts jitter
125690 | 42167.8087 | 2 | LI23 | 21 | 4 | HARDWARE | HW Problem | Large Diurnal swings only seen on the forward Energy, possibly related to the Stripline coupler.
125695 | 42168.8809 | 1 | LI28 | 31 | 4 | HARDWARE | HW Problem | station 28-3 Mod Fault. Huge water leak reported.
125695 | 42168.8809 | 3 | LI28 | 31 | 4 | HARDWARE | HW Problem | station 28-3 Mod Fault. Huge water leak reported.
125695 | 42168.8809 | 2 | LI28 | 31 | 4 | HARDWARE | HW Problem | station 28-3 Mod Fault. Huge water leak reported.
124516 | 42116.6437 | 2.5 | LI28 | 21 | 4 | HARDWARE | HW Problem | Kly Mod 28-2: remove CX1937AX thyratron and control equipment. Replace EEV thyratron with L4888 thyratron previously removed. Adapt DC filament and reservoir supplies to run L4888 thyratron. Make exterior adjustable controls for DC supplies. Test L4888 long term for stability with DC heater supplies. Estimated time to complete re-retrofit: 3-4 days.
124518 | 42116.7221 | 0.75 | LI26 | 31 | 4 | HARDWARE | HW Problem | Klystron 26-3 mod fault. Repeating mod faults.
124518 | 42116.7221 | 2 | LI26 | 31 | 4 | HARDWARE | HW Problem | Klystron 26-3 mod fault. Repeating mod faults.
124847 | 42130.6075 | 2 | LI27 | 51 | 4 | HARDWARE | HW Problem | 27-5 (SEE CATER #123872 Downtime job. for wire swap.) Needs AC filters; hot vs neutral wires need to be reversed
125194 | 42144.5304 | 1.5 | LI26 | 51 | 4 | HARDWARE | HW Problem | 26-5: Tripping. Fault display shows HVOV/External Faults.
125194 | 42144.5304 | 2 | LI26 | 51 | 4 | HARDWARE | HW Problem | 26-5: Tripping. Fault display shows HVOV/External Faults.
125194 | 42144.5304 | 2 | LI26 | 51 | 4 | HARDWARE | HW Problem | 26-5: Tripping. Fault display shows HVOV/External Faults.
125204 | 42144.6424 | 2 | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X amplitude jitter slowly increasing since April. Slow increase likely indicates thyratron delay jitter which is more apparent as amplitude jitter due to narrow beam voltage pulse.
125322 | 42150.2165 |  | LI23 | 71 | 4 | HARDWARE | HW Problem | Kly LI 23 - 7 tripping off. about 10 times in the last 2 hours.(see MCC log for plot.)
125653 | 42166.3532 |  | LI21 | 11 | 4 | HARDWARE | HW Problem | 21-1 It has FLK faults and HVR faults that we can reset, but it trips maybe 1 minute later.
125788 | 42174.3598 | 1 | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 L1X - Dequing under current fault flashing red on epics modulator display
125788 | 42174.3598 | 1 | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 L1X - Dequing under current fault flashing red on epics modulator display
125805 | 42174.5782 | 1 | LI04 | 51 | 4 | HARDWARE | HW Problem | 4-5 Mod charging diode stack needs to replaced De- spiking coil insulation hanging loose
125816 | 42174.6142 | 1.5 | LI23 | 41 | 4 | HARDWARE | HW Problem | 23-4: This station is cycling almost constantly. Took it off the beam as it was unbearable. I see large excursions in the phase..
125824 | 42174.7497 | 0.5 | LI26 | 71 | 4 | HARDWARE | HW Problem | 26-7 showing numerous AMPL errors. Drops of about 2-3% regularly. See 125937 Closed.
125854 | 42177.6904 | 1 | LI23 | 11 | 4 | HARDWARE | HW Problem | 23-1 trips on ACOC after a single external fault then requires reset.
125854 | 42177.6904 | 1 | LI23 | 11 | 4 | HARDWARE | HW Problem | 23-1 trips on ACOC after a single external fault then requires reset.
124311 | 42108.1605 | 1 | LI28 | 61 | 4 | HARDWARE | HW Problem | Took 28-6 off the beam; AMPL is sagging badly.
125626 | 42165.112 |  | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7 frequently MOD faulting
125640 | 42165.5866 | 3 | LI28 | 61 | 4 | HARDWARE | HW Problem | 28-6 Modulator, Tune PFN network: Currently PW=2.98uS, Slope=0.8, Ripl=0.3. Completed claimed. Evaluate tune with engneers.
125609 | 42164.4758 | 0.5 | LI20 | 51 | 4 | HARDWARE | HW Problem | TCAV0 (20-5) mod fault won't reset remotely
125765 | 42172.7383 | 0.8 | LI28 | 11 | 4 | HARDWARE | HW Problem | 28-1 Modulator in FLK, No reset.
125765 | 42172.7383 | 4 | LI28 | 11 | 4 | HARDWARE | HW Problem | 28-1 Modulator in FLK, No reset.
125786 | 42173.9306 | 0.5 | LI22 | 71 | 4 | HARDWARE | HW Problem | 22-7 temperature fault; WG temp. shows steep 3 degree drop.
125786 | 42173.9306 | 1 | LI22 | 71 | 4 | HARDWARE | HW Problem | 22-7 temperature fault; WG temp. shows steep 3 degree drop.
125863 | 42178.5836 | 1 | LI06 | 41 | 4 | HARDWARE | HW Problem | 6-4 PFN strap needs to be replaced ( upper, RHS, # 1 ) K24 has a burnt lug that needs to be replaced.
125866 | 42178.59 | 1 | LI06 | 81 | 4 | HARDWARE | HW Problem | 6-8 Mod EOLC stack needs to be replaced
125868 | 42178.5966 | 0.75 | LI09 | 41 | 4 | HARDWARE | HW Problem | 9-4 Mod K27 contacts need to be replaced.
125896 | 42179.9305 | 1 | LI21 | 21 | 4 | HARDWARE | HW Problem | See multiple messages about HV fcous supply faults and Beam voltage exceeds Limits. as well as NO RF DRIVE. Ask PEM to look at kly.
125896 | 42179.9305 | 3.5 | LI21 | 21 | 4 | HARDWARE | HW Problem | See multiple messages about HV fcous supply faults and Beam voltage exceeds Limits. as well as NO RF DRIVE. Ask PEM to look at kly.
124347 | 42109.5334 | 0.75 | LI29 | 81 | 4 | HARDWARE | HW Problem | 29-8: Seeing some beam volts jitter.
124390 | 42110.6428 | 0.5 | LI29 | 51 | 4 | HARDWARE | HW Problem | 29-5 Modulator, Tune PFN
124424 | 42111.5511 |  | LI28 | 71 | 4 | HARDWARE | HW Problem | Klystron 28-7 DL_WG_Temp alarm, at 114.1 degree. warning is at 113.9 and high alarm is at 114.0
124473 | 42115.5838 | 2 | LI21 | 71 | 4 | HARDWARE | HW Problem | 21-7 Mod T20 Phase B & C Jumper cables and brass fittings Need to be replaced due to looseness and overheating Mod Thy Heater Xformer (under thy chassis) has a burnt Lug and rusted screws that need attention.
125043 | 42138.0864 | 1 | LI30 | 11 | 4 | HARDWARE | HW Problem | 30-1 MOD faults excessively. High phase jitter.
125862 | 42178.5816 | 1 | LI06 | 11 | 4 | HARDWARE | HW Problem | 6-1 Mod de- spiking coil has a broken resistor
125864 | 42178.5855 | 1 | LI06 | 51 | 4 | HARDWARE | HW Problem | 6-5 D q ing Cap wire/lug is badly burnt .
125876 | 42179.1438 | 1 | LI22 | 81 | 4 | HARDWARE | HW Problem | 22-8 MOD fault.
125945 | 42183.6399 | 2 | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3 trips too frequently to use. See 125923. See HSTB in MCC ELOG: See 125923(closed). CATER # 125923 Comments 'Found PIOP" messages showing the unit has 139 PIOP crashes.'
125945 | 42183.6399 | 1 | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3 trips too frequently to use. See 125923. See HSTB in MCC ELOG: See 125923(closed). CATER # 125923 Comments 'Found PIOP" messages showing the unit has 139 PIOP crashes.'
125947 | 42184.3256 | 1 | LI09 | 81 | 4 | HARDWARE | HW Problem | 9-8 Mod charging stack needs to be replaced.
125949 | 42184.3286 | 2 | LI10 | 61 | 4 | HARDWARE | HW Problem | 10-6 K26 and K27 Contacts needs to be replaced Mod EOLC stack needs to be replaced .
124330 | 42108.6977 | 0.5 | LI26 | 71 | 4 | HARDWARE | HW Problem | Klys 26-7 VVS fault - won't reset. 2 CATERs open.
124696 | 42124.5418 | 1 | LI18 | 41 | 4 | HARDWARE | HW Problem | 18-4: PAD "CBL".
124746 | 42126.9656 | 0.5 | LI18 | 31 | 4 | HARDWARE | HW Problem | 18-3 WGV fault.
124962 | 42133.4684 | 0.5 | LI04 | 51 | 4 | HARDWARE | HW Problem | 4-5 PAD not working. Lots of red on the piop page.
125957 | 42184.5799 | 0.5 | LI22 | 71 | 4 | HARDWARE | HW Problem | 22-7 KLT. See 125831, 125938 POMM job.
125968 | 42184.638 | 0.5 | LI21 | 41 | 4 | HARDWARE | HW Problem | 21-4 being mighty jittery.
125974 | 42185.2669 | 1.5 | LI28 | 61 | 4 | HARDWARE | HW Problem | Klys 28-6 comes on fine, but usually trips within a minute. HVPS Reg. Bd. suspected.
125984 | 42185.5385 | 2 | LI22 | 11 | 4 | HARDWARE | HW Problem | 22-1 is reading Mod VVS Voltage Fault and will not reset remotely
125984 | 42185.5385 | 1 | LI22 | 11 | 4 | HARDWARE | HW Problem | 22-1 is reading Mod VVS Voltage Fault and will not reset remotely
124596 | 42121.4369 | 0.25 | LI21 | 51 | 4 | HARDWARE | HW Problem | 21-5 MOD fault no remote reset
124634 | 42122.5525 | 1 | LI27 | 31 | 4 | HARDWARE | HW Problem | 27-3: Seeing some beam volts jitter.
124674 | 42123.9114 | 2 | LI30 | 51 | 4 | HARDWARE | HW Problem | Everything looks ok, could probably be improved by sled tuning.
124693 | 42124.5365 | 1 | LI03 | 61 | 4 | HARDWARE | HW Problem | 3-6: Seeing lots of beam volts jitter.
124693 | 42124.5365 | 1 | LI03 | 61 | 4 | HARDWARE | HW Problem | 3-6: Seeing lots of beam volts jitter.
125991 | 42185.7351 | 1.5 | LI30 | 61 | 4 | HARDWARE | HW Problem | 30-6 mod faulted and won't reset remotely
125999 | 42186.4631 |  | LI23 | 21 | 4 | HARDWARE | HW Problem | 23-2 HVR fault, won't reset
126005 | 42186.5245 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | Suspect standby PAC or modulator timing is not correct. Noticed this by observing the 6x6 commands for L1S phase and amplitude when changing rates. Temperature change is small, <0.5 deg F, and 6x6 is able to compensate. Need invasive time (POMM) to investigate (need beam rate to 10 hz, and disable LLRF feedback).
126009 | 42186.6602 | 3 | LI25 | 21 | 4 | HARDWARE | HW Problem | Klystron 25-2 has tripped off and will not reset remotely
126012 | 42186.8963 | 0.1 | LI27 | 31 | 4 | HARDWARE | HW Problem | 27-3 HVR no reset
126014 | 42187.0188 |  | LI25 | 21 | 4 | HARDWARE | HW Problem | 25-2 MOD: Tried to reset 3 times but no change was detected.
124353 | 42109.6398 | 2.5 | LI13 | 61 | 4 | HARDWARE | HW Problem | 13-6 MOD fault, won't remotely reset.
124542 | 42117.9435 | 1 | LI20 | 81 | 4 | HARDWARE | HW Problem | L0B modulator tripped, won't reset.
124678 | 42124.2619 | 0.1 | LI27 | 11 | 4 | HARDWARE | HW Problem | Klystron 27-1 mod fault.
124692 | 42124.5089 | 4 | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 (aka L1X): Although the klystron tank reservoir site glass shows plenty of oil, thetank seems to be leaking a lot.
124692 | 42124.5089 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 (aka L1X): Although the klystron tank reservoir site glass shows plenty of oil, thetank seems to be leaking a lot.
124692 | 42124.5089 | 2 | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 (aka L1X): Although the klystron tank reservoir site glass shows plenty of oil, thetank seems to be leaking a lot.
124804 | 42129.5692 | 0.5 | LI29 | 31 | 4 | HARDWARE | HW Problem | 29-3 T20 input snubber wires are badly burnt.
124967 | 42133.6599 | 2.5 | LI01 | 61 | 4 | HARDWARE | HW Problem | KLYS 1-6 MOD fault no reset
125090 | 42139.2547 |  | LI03 | 61 | 4 | HARDWARE | HW Problem | 3-6 in MOD fault.
125214 | 42144.8723 | 1 | LI20 | 81 | 4 | HARDWARE | HW Problem | L0B showing amplitude jitter. (2 state)
125266 | 42146.4391 | 0.5 | LI12 | 21 | 4 | HARDWARE | HW Problem | The first few times it cleared on its own but the KLV error keeps occurring.
125266 | 42146.4391 |  | LI12 | 21 | 4 | HARDWARE | HW Problem | The first few times it cleared on its own but the KLV error keeps occurring.
125266 | 42146.4391 |  | LI12 | 21 | 4 | HARDWARE | HW Problem | The first few times it cleared on its own but the KLV error keeps occurring.
125990 | 42185.6805 | 1.5 | LI27 | 51 | 4 | HARDWARE | HW Problem | klystron 27-5 mod faulted and won't reset remotely. Has core bias current fault, current sensor bad.
126000 | 42186.4839 | 0 | LI22 | 11 | 4 | HARDWARE | HW Problem | 22-1 VVS fault
126013 | 42186.9243 | 1 | LI22 | 71 | 4 | HARDWARE | HW Problem | 22-7 Klystron temperature fault
126016 | 42187.0383 |  | LI22 | 71 | 4 | HARDWARE | HW Problem | KLYS:LI22:71:DELTTEMP = 5.3 then numerous other channels started to fault. Swinging to large - / + values.
126018 | 42187.1904 | 0.5 | LI30 | 61 | 4 | HARDWARE | HW Problem | 30-6 Plot shows that the unit has dropped out before.
126030 | 42187.5813 | 1 | LI11 | 71 | 4 | HARDWARE | HW Problem | 11-7 Mod has a bad PFN strap ( Lower, RHS, # 2 ) EOLC stack needs to be replaced.
126041 | 42188.2108 | 0.5 | LI24 | 81 | 4 | HARDWARE | HW Problem | LI 24-8, TS/R water leak inside FIAT rack.
126063 | 42190.3801 | 0.5 | LI26 | 51 | 4 | HARDWARE | HW Problem | Klystron 26-5 mod fault.
126065 | 42191.0089 | 1 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4: PHJ & BVJ. Looks like it started around 1900 Sun. evening
125354 | 42151.5789 | 1 | LI24 | 51 | 4 | HARDWARE | HW Problem | 24-5 Mod despiking coil has a broken resistor Thy door clip for grn hook needs to be replaced.
125368 | 42151.8092 | 1 | LI30 | 41 | 4 | HARDWARE | HW Problem | Believed to be caused by the filament supply was very unstable before the PAMM and never settled down after turn-on See elog 785041
125480 | 42158.2692 |  | LI14 | 31 | 4 | HARDWARE | HW Problem | LI14-3 water faults, please check station
125628 | 42165.1662 | 1 | LI21 | 61 | 4 | HARDWARE | HW Problem | 21-6 high phase jitter
125628 | 42165.1662 | 0.5 | LI21 | 61 | 4 | HARDWARE | HW Problem | 21-6 high phase jitter
125628 | 42165.1662 | 0.8 | LI21 | 61 | 4 | HARDWARE | HW Problem | 21-6 high phase jitter
126029 | 42187.5793 | 1 | LI12 | 11 | 4 | HARDWARE | HW Problem | 12-1 Mod needs EOLC stack replaced
126047 | 42188.5908 | 0.75 | LI14 | 51 | 4 | HARDWARE | HW Problem | 15-5 Mod K27 contacts need to be replaced.
126059 | 42189.3495 |  | LI30 | 51 | 4 | HARDWARE | HW Problem | Klys 30-5 has lots of AMPL dropouts BV look okay, but every minuet or so the amplitude drops outs.
126066 | 42191.1093 | 1 | LI24 | 41 | 4 | HARDWARE | HW Problem | 24-4 is tripping frequently Will HVR or MOD fault every few minutes or so. Issue cleared up, Thyratron replaced SEE: CATER # 126608
126066 | 42191.1093 | 1 | LI24 | 41 | 4 | HARDWARE | HW Problem | 24-4 is tripping frequently Will HVR or MOD fault every few minutes or so. Issue cleared up, Thyratron replaced SEE: CATER # 126608
126066 | 42191.1093 | 1 | LI24 | 41 | 4 | HARDWARE | HW Problem | 24-4 is tripping frequently Will HVR or MOD fault every few minutes or so. Issue cleared up, Thyratron replaced SEE: CATER # 126608
126069 | 42191.3467 | 1 | LI22 | 61 | 4 | HARDWARE | HW Problem | 22-6 Klystron vacuum low at 1.0E-7
126236 | 42200.2138 | 1 | LI22 | 71 | 4 | HARDWARE | HW Problem | Klys 22-7 MOD faulted and won't reset remotely
126300 | 42203.0082 | 1 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 tripped off, can't reset remotely.
126387 | 42207.5112 | 2.5 | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2 tripping on ACOC/HVOC
126541 | 42216.2489 | 1 | LI29 | 81 | 4 | HARDWARE | HW Problem | 29-8 Modulator, Primary Filter Caps, Change to newer values
126541 | 42216.2489 | 2 | LI29 | 81 | 4 | HARDWARE | HW Problem | 29-8 Modulator, Primary Filter Caps, Change to newer values
126563 | 42217.3119 | 1 | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 frequent MOD faults. Both reflected energy faults and core bias U/C faults. See noise or instablity on core bias current plot.
124316 | 42108.5387 | 5 | LI25 | 71 | 4 | HARDWARE | HW Problem | Kly Mod 25-7: repair severe BV swings/oscillations. 3 CATERs open.
125285 | 42147.4885 |  | LI27 | 21 | 4 | HARDWARE | HW Problem | 27-2 needs PFN tune: PW = 0.0uS, Slope=1.1, Ripl=0.2
125337 | 42150.6583 | 2 | LI22 | 21 | 4 | HARDWARE | HW Problem | 22-2 tripping frequently on mod trigger overcurrent faults.
125357 | 42151.5838 |  | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7 Dup, Job moved to 125031 Mod Pulse cable near end of life and should be replaced
125439 | 42156.9091 | 2 | LI28 | 61 | 4 | HARDWARE | HW Problem | Klystron LI28-61 has elevated Beam Volt Jitter. History plot shows a jump around 2:30 A.M. from .1 to .2
125439 | 42156.9091 |  | LI28 | 61 | 4 | HARDWARE | HW Problem | Klystron LI28-61 has elevated Beam Volt Jitter. History plot shows a jump around 2:30 A.M. from .1 to .2
125439 | 42156.9091 | 3 | LI28 | 61 | 4 | HARDWARE | HW Problem | Klystron LI28-61 has elevated Beam Volt Jitter. History plot shows a jump around 2:30 A.M. from .1 to .2
125439 | 42156.9091 | 5 | LI28 | 61 | 4 | HARDWARE | HW Problem | Klystron LI28-61 has elevated Beam Volt Jitter. History plot shows a jump around 2:30 A.M. from .1 to .2
125439 | 42156.9091 | 0.8 | LI28 | 61 | 4 | HARDWARE | HW Problem | Klystron LI28-61 has elevated Beam Volt Jitter. History plot shows a jump around 2:30 A.M. from .1 to .2
125439 | 42156.9091 | 3 | LI28 | 61 | 4 | HARDWARE | HW Problem | Klystron LI28-61 has elevated Beam Volt Jitter. History plot shows a jump around 2:30 A.M. from .1 to .2
126070 | 42191.3647 | 3 | LI20 | 71 | 4 | HARDWARE | HW Problem | 20-7 (L0-A) MKSU trigger enable fault.
126477 | 42213.5113 | 1 | LI22 | 11 | 4 | HARDWARE | HW Problem | 22-1: Tripped and reset "Phase Loss" fault twice in the period of 1/2 hour. It is off again as I write less than 15 minutes later.
126487 | 42213.6989 |  | LI29 | 41 | 4 | HARDWARE | HW Problem | Klys 29-4 cycling repeatedly, turned it off. Saw EVO during cycles.
126708 | 42229.2984 | 2 | LI30 | 71 | 4 | HARDWARE | HW Problem | 30-7 Modulator, load side B phase conductor is overheated.
126754 | 42237.5657 | 1 | LI28 | 71 | 4 | HARDWARE | HW Problem | 28-7 Thy Grid trigger chassis needs to be replaced
124379 | 42110.4772 | 0.5 | LI09 | 11 | 4 | HARDWARE | HW Problem | 9-1: MOD fault, no reset.
124584 | 42120.5659 | 0.2 | LI09 | 81 | 4 | HARDWARE | HW Problem | 9-8 FLK. Won't reset.
124851 | 42130.7171 | 0.5 | LI17 | 41 | 4 | HARDWARE | HW Problem | station 17-4 has Waveguide Vacuum Fault
125289 | 42147.9134 |  | LI10 | 71 | 4 | HARDWARE | HW Problem | 10-7 SCRTD EOLC signal present not lit.
125292 | 42147.9158 |  | LI15 | 11 | 4 | HARDWARE | HW Problem | 15-1 SCRTD EOLC signal present not lit
125662 | 42166.5152 | 0.5 | LI23 | 81 | 4 | HARDWARE | HW Problem | Klys 23-8 KMC -- electro-magnet current fault.
126152 | 42195.4228 | 1 | LI30 | 51 | 4 | HARDWARE | HW Problem | 30-5 many bogus OC, focus and re faults and PIOP crashes. This station has been up with no trips for the last three days PMS
126152 | 42195.4228 | 1 | LI30 | 51 | 4 | HARDWARE | HW Problem | 30-5 many bogus OC, focus and re faults and PIOP crashes. This station has been up with no trips for the last three days PMS
126211 | 42199.3686 | 0.5 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 Modulator, Noise when on beam
126217 | 42199.5913 | 1 | LI17 | 21 | 4 | HARDWARE | HW Problem | 17-2 KMB has a badly burnt connection wire No work to start before Fiscal year 2016, Oct 1st, 2015.
126321 | 42205.2381 | 25 | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7 RF amplitude became erratic and dropped to near zero
126370 | 42206.561 | 2 | LI30 | 51 | 4 | HARDWARE | HW Problem | 30-5: Facilities can't see which switch is causing the WSF fault.
124749 | 42127.1705 |  | LI01 | 61 | 4 | HARDWARE | HW Problem | 1-6 has a false EVC status on SCP
124837 | 42130.5267 | 1 | LI20 | 51 | 4 | HARDWARE | HW Problem | TCAV0 (20-5) in FLK. Looks like "WGV".
125117 | 42139.8921 | 0.75 | LI25 | 81 | 4 | HARDWARE | HW Problem | Kly. LI 25 - 8 is tripping on "MKSU Or "External Fault". Took it off the beam.
126102 | 42192.4284 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | KLYS 21-1 L1S 21-2 L1X and 21-3 all have the same PDU in the same crate We see problems with L1X AND also 21-3 starting at that time. As measured with jitter FPT at raise time of beam volts waveform. lclshome -> RF -> Ops Klys Display -> Jitter Reports -> TS Jitter -> Trigger Jitter (one per station) KLYS:LI21:11:MKBVTRIGJITT KLYS:LI21:21:MKBVTRIGJITT KLYS:LI21:31:MKBVTRIGJITT
126208 | 42199.3303 | 0.2 | LI29 | 41 | 4 | HARDWARE | HW Problem | 29-4 SLED needs tuning due to diurnal amplitude changes.
126429 | 42210.6258 | 0.75 | LI21 | 41 | 4 | HARDWARE | HW Problem | station 21-4 mod fault
126741 | 42237.4537 | 1 | LI22 | 31 | 4 | HARDWARE | HW Problem | 22-3 Mod Thy grid trigger chassis Needs to be replaced
126742 | 42237.4548 | 0.25 | LI22 | 81 | 4 | HARDWARE | HW Problem | Downtime: 22-8 Modulator, Broken despiking resistor
126743 | 42237.4581 | 1 | LI23 | 71 | 4 | HARDWARE | HW Problem | Downtime: 23-7 Modulator, Has hardware problems. Replace PFN strap ( Upper, LHS, #2 ) Replace 1 Despiking resistor. Burnt D Q ing cap lug. Broken Thyratron compartment ground hook door clip.
126756 | 42237.5721 | 1 | LI29 | 11 | 4 | HARDWARE | HW Problem | 29-1 Mod has a burnt/broken Carborundum resistor
126762 | 42237.5896 | 1 | LI30 | 61 | 4 | HARDWARE | HW Problem | 30-6 Mod has a broken Carborundum resistor
124408 | 42111.1561 | 1 | LI19 | 41 | 4 | HARDWARE | HW Problem | 19-4 MOD fault, no remote reset
125284 | 42147.378 | 1 | LI09 | 61 | 4 | HARDWARE | HW Problem | Kly Mod 9-6: mod faulted, no reset. HVOC's cause.
125295 | 42147.9188 |  | LI17 | 41 | 4 | HARDWARE | HW Problem | 17-4 SCRTD EOLC signal present not lit
125787 | 42174.0236 |  | LI21 | 81 | 4 | HARDWARE | HW Problem | Feels like something is screwed up with the timing. Stanek note: The SBI is not seeing Standby triggers (channel 7 of the PDU). This could be a PDU problem, or a SBI problem. STBY" LED on SBI not lit. No RF standby pulse on scope (J. Craft). - edit tallmike - SBI swapped, still no STANDBY light on SBI. "
125801 | 42174.5624 | 1.5 | LI05 | 41 | 4 | HARDWARE | HW Problem | 5-4 Grid trigger chassis needs to be replaced EOLC stack needs to be replaced K24 has a burnt lug
125809 | 42174.5874 | 2.5 | LI03 | 61 | 4 | HARDWARE | HW Problem | 3-6 PFN straps need to be replaced ( lower, RHS # 1, 3 ) ( Lower center jumper strap ) De-spiking coil lead needs heat shrink or tape.
125828 | 42175.3509 | 0.5 | LI28 | 11 | 4 | HARDWARE | HW Problem | 28-1 cable fault returned
125828 | 42175.3509 | 2 | LI28 | 11 | 4 | HARDWARE | HW Problem | 28-1 cable fault returned
125828 | 42175.3509 | 1 | LI28 | 11 | 4 | HARDWARE | HW Problem | 28-1 cable fault returned
125828 | 42175.3509 | 1 | LI28 | 11 | 4 | HARDWARE | HW Problem | 28-1 cable fault returned
126636 | 42221.5149 | 0.5 | LI22 | 21 | 4 | HARDWARE | HW Problem | 22-2: DQing current = 33A. Needs to lower T20 tap.
124840 | 42130.5517 |  | LI23 | 81 | 4 | HARDWARE | HW Problem | 23-8: DQing current low. Red on EPICS
125424 | 42155.5917 | 0.3 | LI20 | 51 | 4 | HARDWARE | HW Problem | T-CAV0 mod fault. Kly 20-5
125477 | 42157.9206 |  | LI21 | 81 | 4 | HARDWARE | HW Problem | Duplicate 125319 Klystron 21-8 phase jitter as seen on the FTP. Beam volt jitter looks fine.. So the phase jitter might not be real. Requires further investigation.
125477 | 42157.9206 | 0.8 | LI21 | 81 | 4 | HARDWARE | HW Problem | Duplicate 125319 Klystron 21-8 phase jitter as seen on the FTP. Beam volt jitter looks fine.. So the phase jitter might not be real. Requires further investigation.
125551 | 42161.2008 | 1 | LI03 | 31 | 4 | HARDWARE | HW Problem | 3-3 MOD flt
125615 | 42164.5651 | 1 | LI24 | 41 | 4 | HARDWARE | HW Problem | 24-4 modulator, problem founds: De-spiking coil, EOLC and charging diode stacks need to replace. 1. EOLC and charging diode stacks measured low resistance. 2. De-spiking coil shows burnt marks and cracks 3. Cable that connected to De-spiking coil needs to be replaced. 4. Thyratron voltage divider lead needs to be re-terminated.
125615 | 42164.5651 | 1 | LI24 | 41 | 4 | HARDWARE | HW Problem | 24-4 modulator, problem founds: De-spiking coil, EOLC and charging diode stacks need to replace. 1. EOLC and charging diode stacks measured low resistance. 2. De-spiking coil shows burnt marks and cracks 3. Cable that connected to De-spiking coil needs to be replaced. 4. Thyratron voltage divider lead needs to be re-terminated.
125615 | 42164.5651 | 1 | LI24 | 41 | 4 | HARDWARE | HW Problem | 24-4 modulator, problem founds: De-spiking coil, EOLC and charging diode stacks need to replace. 1. EOLC and charging diode stacks measured low resistance. 2. De-spiking coil shows burnt marks and cracks 3. Cable that connected to De-spiking coil needs to be replaced. 4. Thyratron voltage divider lead needs to be re-terminated.
125615 | 42164.5651 | 1 | LI24 | 41 | 4 | HARDWARE | HW Problem | 24-4 modulator, problem founds: De-spiking coil, EOLC and charging diode stacks need to replace. 1. EOLC and charging diode stacks measured low resistance. 2. De-spiking coil shows burnt marks and cracks 3. Cable that connected to De-spiking coil needs to be replaced. 4. Thyratron voltage divider lead needs to be re-terminated.
125615 | 42164.5651 | 1 | LI24 | 41 | 4 | HARDWARE | HW Problem | 24-4 modulator, problem founds: De-spiking coil, EOLC and charging diode stacks need to replace. 1. EOLC and charging diode stacks measured low resistance. 2. De-spiking coil shows burnt marks and cracks 3. Cable that connected to De-spiking coil needs to be replaced. 4. Thyratron voltage divider lead needs to be re-terminated.
125650 | 42166.1928 | 1 | LI21 | 41 | 4 | HARDWARE | HW Problem | Klystron LI21-41 is tripped off and will not reset remotely.
125711 | 42170.6203 | 0.1 | LI20 | 93 | 4 | HARDWARE | HW Problem | 20-3C is CBL on the klystron CUD.
125738 | 42171.8191 | 0.1 | LI23 | 41 | 4 | HARDWARE | HW Problem | Klys 23-4 showing CBL fault, IPL no help.
125738 | 42171.8191 | 2 | LI23 | 41 | 4 | HARDWARE | HW Problem | Klys 23-4 showing CBL fault, IPL no help.
125763 | 42172.7352 | 0.25 | LI30 | 31 | 4 | HARDWARE | HW Problem | Klys 30-3 showing beam volts jitter.
125784 | 42173.8865 | 0.5 | LI21 | 81 | 4 | HARDWARE | HW Problem | station 21-8 has had Amplitude of zero since 21:00 tonight
125830 | 42176.3661 | 1 | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8: Thyratron compartment light out. Bulb swapped, still out.
125923 | 42180.8222 |  | LI24 | 31 | 4 | HARDWARE | HW Problem | Found PIOP" messages showing the unit has 139 PIOP crashes.
125948 | 42184.3267 | 1 | LI10 | 21 | 4 | HARDWARE | HW Problem | 10-2 Mod EOLC stack needs to be replaced
124649 | 42123.2034 |  | LI23 | 11 | 4 | HARDWARE | HW Problem | DUPLICATE of 125854 - Station was running fine then started reporting no accelerate triggers. Attempted reset and IPL PIOP as well as PDU reset problem persists. Need to put station to standby until opportunity to investigate further. Duplicate of 125767?
125220 | 42145.2921 |  | LI26 | 11 | 4 | HARDWARE | HW Problem | 26-1 PAD sample gate shows jitter
125558 | 42162.4411 | 0.5 | LI25 | 81 | 4 | HARDWARE | HW Problem | Klys 25-8 having frequent amplitude drop outs and faults. It was doing this yesterday morning also but was ok throughout the night. SEE CATER #125553 for previous efforts.
125558 | 42162.4411 |  | LI25 | 81 | 4 | HARDWARE | HW Problem | Klys 25-8 having frequent amplitude drop outs and faults. It was doing this yesterday morning also but was ok throughout the night. SEE CATER #125553 for previous efforts.
125846 | 42177.3245 | 1 | LI01 | 21 | 4 | HARDWARE | HW Problem | 1-2 All K27 contacts need to be replaced
125848 | 42177.3325 | 2 | LI00 | 1 | 4 | HARDWARE | HW Problem | 0-01 D Q ing resistor assem at top of mod needs to be replace as all the connection links are badly burnt.
125869 | 42178.6172 | 1 | LI26 | 11 | 4 | HARDWARE | HW Problem | 26-1 tripping frequently. Errors display shows multiple over-current faults since 08:00 this morning.
125871 | 42178.7049 | 1 | LI28 | 41 | 4 | HARDWARE | HW Problem | 06-18-15 it was 0.2 %
126197 | 42198.5926 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | Klys 21-2, L1X experiencing bad AMPL/phase jitter. Returned reservior voltage to 3.4 volts (value before Decker lowered it).
126246 | 42200.4644 | 2.5 | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X AMPL jitter very bad.
126309 | 42203.8131 | 0.5 | LI14 | 31 | 4 | HARDWARE | HW Problem | 14-3: Bad Cable Status. Check with PAD Sub-Box.
126367 | 42206.4601 | 2 | LI25 | 61 | 4 | HARDWARE | HW Problem | Klystron 25-6 mod faulted and won't reset remotely
126411 | 42209.0287 | 2 | LI24 | 61 | 4 | HARDWARE | HW Problem | Klystron 24-6 repeating mod fault. see 127376
126411 | 42209.0287 |  | LI24 | 61 | 4 | HARDWARE | HW Problem | Klystron 24-6 repeating mod fault. see 127376
124354 | 42109.678 | 1 | LI20 | 71 | 4 | HARDWARE | HW Problem | L0A's beam volts has many frequent 1kV dips. Shows up on FEL as big downward spikes. Probably means thyratron needs to be ranged.
124486 | 42116.0927 | 0.75 | LI30 | 61 | 4 | HARDWARE | HW Problem | 30-6 FTPs show high phase jitter. The phase jitter FTP is showing pronounced two-state behavior; not clear whether or not this is physical.
124576 | 42119.67 | 2.5 | LI15 | 31 | 4 | HARDWARE | HW Problem | 15-3 has mod faulted and will not reset remotely
124735 | 42125.6908 |  | LI17 | 81 | 4 | HARDWARE | HW Problem | 17-8 HVR fault, won't reset. Fails remote turn on.
124742 | 42126.4541 | 1 | LI20 | 51 | 4 | HARDWARE | HW Problem | Kly Mod 20-5 (TCav): mod faulted, no reset.
124850 | 42130.7001 |  | LI25 | 11 | 4 | HARDWARE | HW Problem | other window temps in that sector (other klystrons) show 100DeGF. Called AMRF?PEM to have a look.
126352 | 42205.7974 | 0.5 | LI21 | 81 | 4 | HARDWARE | HW Problem | 21-8 is having bad, sporadic jitter. For the high energy config we are currently running, we are unable to run without 21-8 so it is still on the beam in use.
126352 | 42205.7974 | 2.5 | LI21 | 81 | 4 | HARDWARE | HW Problem | 21-8 is having bad, sporadic jitter. For the high energy config we are currently running, we are unable to run without 21-8 so it is still on the beam in use.
126352 | 42205.7974 | 3 | LI21 | 81 | 4 | HARDWARE | HW Problem | 21-8 is having bad, sporadic jitter. For the high energy config we are currently running, we are unable to run without 21-8 so it is still on the beam in use.
126492 | 42214.0129 | 1.5 | LI25 | 71 | 4 | HARDWARE | HW Problem | 25-7 is showing MOD, CBL and TRG errors and will not reset remotely.
126492 | 42214.0129 | 1 | LI25 | 71 | 4 | HARDWARE | HW Problem | 25-7 is showing MOD, CBL and TRG errors and will not reset remotely.
126763 | 42237.5908 | 0.5 | LI22 | 71 | 4 | HARDWARE | HW Problem | Downtime: 22-3 Modulator, Thyratron compartment fan INOP.
124554 | 42118.4504 |  | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 install toroid on back of MKSU focus monitor input cable to reduce noise.
125179 | 42143.6921 | 1 | LI24 | 11 | 4 | HARDWARE | HW Problem | 24-1 It may be some time slot jitter or something because we do not see any beam volts jitter.
125321 | 42150.1837 | 1 | LI25 | 81 | 4 | HARDWARE | HW Problem | (see plots in MCC log.) edit: when active, 25-8 modulator trips about every five minutes.
125414 | 42154.4526 | 2 | LI09 | 71 | 4 | HARDWARE | HW Problem | Klystron 9-7 MKSU trigger enable fault. No reset.
125475 | 42157.8059 | 0.75 | LI20 | 51 | 4 | HARDWARE | HW Problem | Klystron 20-5 T-CAV0 mod fault
125641 | 42165.5914 | 1.5 | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 Modulator, Tune PFN network: Currently PW=2.8uS, slope=0.4, Ripl=0.3
126224 | 42199.6431 | 2 | LI28 | 81 | 4 | HARDWARE | HW Problem | 28-8: Constant ACOCHVOC trips
126226 | 42199.6722 | 2 | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3: Arcing sound when turning on.
126389 | 42207.548 | 1 | LI30 | 41 | 4 | HARDWARE | HW Problem | 30-4: Seeing some beam volts jitter
126496 | 42214.3892 | 2 | LI23 | 51 | 4 | HARDWARE | HW Problem | 23-5 Beam voltage dropped 6KV 0n 7-28-15 at 17:40 and has not recovered.
126749 | 42237.4669 | 1 | LI26 | 11 | 4 | HARDWARE | HW Problem | 26-1 Mod PFN strap needs to be replaced ( Lower, RHS, #3 )
126758 | 42237.5766 | 1 | LI29 | 41 | 4 | HARDWARE | HW Problem | 29-4 Mod needs to have EOLC stack replaced
126760 | 42237.5835 | 1 | LI29 | 71 | 4 | HARDWARE | HW Problem | 29-7 EOLC stack needs to be replaced D Qing cap Lug is badly burnt Carborundum resistor is broken
126775 | 42241.5475 | 1 | LI21 | 41 | 4 | HARDWARE | HW Problem | 21-4 PFN strap (U, RHS, #3 ) (Lower, RHS, #1 )
124308 | 42108.0486 | 2 | LI24 | 41 | 4 | HARDWARE | HW Problem | 24-4 had BVLT drop to zero 10 times in 8 hours. Only 3 trips in the previous 7 days.
124694 | 42124.5378 | 1 | LI04 | 11 | 4 | HARDWARE | HW Problem | 4-1: Seeing lots of beam volts jitter.
125238 | 42145.5525 | 1 | LI18 | 41 | 4 | HARDWARE | HW Problem | 18-4 PAD "Lost Phase"
125238 | 42145.5525 | 0.5 | LI18 | 41 | 4 | HARDWARE | HW Problem | 18-4 PAD "Lost Phase"
125238 | 42145.5525 | 1 | LI18 | 41 | 4 | HARDWARE | HW Problem | 18-4 PAD "Lost Phase"
126518 | 42215.1874 | 1 | LI22 | 61 | 4 | HARDWARE | HW Problem | 22-6 has mod faulted and will not reset remotely
126546 | 42216.4221 | 0.5 | LI27 | 11 | 4 | HARDWARE | HW Problem | 27-1 Adjust beam voltage to 327KV and update IKE. RP to re-adjust radiation boundaries.
126676 | 42225.5073 | 1.8 | LI22 | 11 | 4 | HARDWARE | HW Problem | Klystron 22-1 MOD VVS voltage fault. 2 additional CATERS are active 126639 and 126658.
124741 | 42126.3503 | 7 | LI28 | 21 | 4 | HARDWARE | HW Problem | Kly Mod 28-2: next PAMM, install decoupling capacitors on thyratron base and DC power supplies. Examine EOLC circuit and PFN cap tuning coils for signs of micro-arcing leading to noise on BV pulse.
124901 | 42131.8791 | 1 | LI10 | 61 | 4 | HARDWARE | HW Problem | (see pic.s in MCC log.)
124952 | 42132.7471 | 0.75 | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3 large RF amplitude jitter.
125363 | 42151.5976 | 2 | LI22 | 61 | 4 | HARDWARE | HW Problem | 22-6 Mod T20 network resistors are badly burnt Mod HV Meter cap is blown..R/R
126091 | 42192.2623 | 2.5 | LI15 | 11 | 4 | HARDWARE | HW Problem | 15-1 K26 and K27 Contacts need to be replaced. Pulse cable needs to be replaced EOLC stack needs to be replaced
126219 | 42199.595 | 4 | LI03 | 71 | 4 | HARDWARE | HW Problem | 3-7 KMB has a badly overheated wire / Lug that needs to be Repaired. No work to start before Fiscal year 2016, Oct 1st, 2015
126234 | 42199.9322 | 0.5 | LI23 | 71 | 4 | HARDWARE | HW Problem | Klystron 23-7 repeating mod faults. Station tripped x4 times in 30 minutes.
126234 | 42199.9322 |  | LI23 | 71 | 4 | HARDWARE | HW Problem | Klystron 23-7 repeating mod faults. Station tripped x4 times in 30 minutes.
126234 | 42199.9322 | 1 | LI23 | 71 | 4 | HARDWARE | HW Problem | Klystron 23-7 repeating mod faults. Station tripped x4 times in 30 minutes.
126380 | 42207.038 | 1 | LI23 | 11 | 4 | HARDWARE | HW Problem | Klystron LI23-11 is tripped off and will not reset remotely.
124933 | 42132.5301 | 0.5 | LI13 | 21 | 4 | HARDWARE | HW Problem | 13-2 tripped on a mod fault; unable to reset remotely.
125453 | 42157.4715 | 0.75 | LI26 | 51 | 4 | HARDWARE | HW Problem | 26-5: DQing Undercurrent alarm on EPICS.
125453 | 42157.4715 | 0.5 | LI26 | 51 | 4 | HARDWARE | HW Problem | 26-5: DQing Undercurrent alarm on EPICS.
126093 | 42192.2652 | 1 | LI15 | 31 | 4 | HARDWARE | HW Problem | 15-3 K25 and K27 Contacts need to be replaced.
126110 | 42192.4967 | 1 | LI16 | 51 | 4 | HARDWARE | HW Problem | 16-5 waveguide vacuum fault
126126 | 42193.5416 | 1 | LI16 | 51 | 4 | HARDWARE | HW Problem | 16-5 Pulse cable badly corroded and needs to be replaced. PFN straps ( upper, RHS, # 3 ) K27 contacts need to be replaced
126126 | 42193.5416 |  | LI16 | 51 | 4 | HARDWARE | HW Problem | 16-5 Pulse cable badly corroded and needs to be replaced. PFN straps ( upper, RHS, # 3 ) K27 contacts need to be replaced
126165 | 42195.6021 | 1 | LI19 | 41 | 4 | HARDWARE | HW Problem | 19-4 K26 and K27 contacts need to be replaced ATTN: RFWCF required before working on mod Please see Jim Craft for Documentation
126223 | 42199.6014 | 2 | LI11 | 81 | 4 | HARDWARE | HW Problem | 11-8 KMB breaker Lug shows signs of excessive over heating No work to start before Fiscal year 2016, Oct 1st, 2015
126290 | 42201.9899 |  | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 MOD faulting frequently. Happened three times in the space of an hour at the end of 7/16 swing shift.
126292 | 42202.3627 | 0.5 | LI20 | 51 | 4 | HARDWARE | HW Problem | Klystron mod fault 20-5 (TCAV0)
126341 | 42205.6163 | 2 | LI22 | 71 | 4 | HARDWARE | HW Problem | Klystron 22-7 mod faulted and won't reset remotely
124736 | 42125.8872 | 0.75 | LI10 | 41 | 4 | HARDWARE | HW Problem | 10-4 MOD fault, won't reset
124806 | 42129.5717 | 1 | LI29 | 81 | 4 | HARDWARE | HW Problem | 29-8 Mod charging diode stack needs to be replaced.
125006 | 42136.1488 | 0.5 | LI21 | 81 | 4 | HARDWARE | HW Problem | 21-8 spontaneous amplitude jitter.
125159 | 42143.1959 | 1 | LI21 | 41 | 4 | HARDWARE | HW Problem | 21-4 main Control Power red LED indicating lamp is not lit. Cabinet lights also do not work. Need to check T20 compartment lamp fuse.
125176 | 42143.6352 | 4 | LI28 | 21 | 4 | HARDWARE | HW Problem | Kly Mod 28-2: station does not turn on in remote mode. Works well in local mode. Check connections and interface board.
125290 | 42147.9142 |  | LI12 | 31 | 4 | HARDWARE | HW Problem | 12-3 SCRTD EOLC signal present not lit
125630 | 42165.1859 | 1 | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X beam volts are momentarily dropping to 0 about once per hours.
125630 | 42165.1859 | 2 | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X beam volts are momentarily dropping to 0 about once per hours.
126109 | 42192.4951 | 1 | LI12 | 21 | 4 | HARDWARE | HW Problem | 12-2 waveguide vacuum fault
126204 | 42199.3258 | 0.2 | LI22 | 21 | 4 | HARDWARE | HW Problem | 22-2 SLED needs tuning due to diurnal amplitude changes.
126207 | 42199.3291 | 0.2 | LI26 | 11 | 4 | HARDWARE | HW Problem | 26-1 SLED needs tuning due to diurnal amplitude changes.
126237 | 42200.2553 | 2 | LI30 | 61 | 4 | HARDWARE | HW Problem | Downtime 30-6 primary caps are wrong value =25uF, should be 12 uF 660VAC
126520 | 42215.2078 |  | LI29 | 51 | 4 | HARDWARE | HW Problem | 29-5 has high phase and amplitude jitter but beam volt jitter is ok. There is noise on the RF drive signal, amplitude and phase signals from the IPA & PAD
126520 | 42215.2078 | 2 | LI29 | 51 | 4 | HARDWARE | HW Problem | 29-5 has high phase and amplitude jitter but beam volt jitter is ok. There is noise on the RF drive signal, amplitude and phase signals from the IPA & PAD
126520 | 42215.2078 | 1 | LI29 | 51 | 4 | HARDWARE | HW Problem | 29-5 has high phase and amplitude jitter but beam volt jitter is ok. There is noise on the RF drive signal, amplitude and phase signals from the IPA & PAD
126522 | 42215.3527 | 1 | LI26 | 51 | 4 | HARDWARE | HW Problem | Downtime: 26-5 - Replace SCR modules B & C in SCR panel
126576 | 42219.3795 | 0.6 | LI24 | 31 | 4 | HARDWARE | HW Problem | Klystron 24-3 mod fault.
126665 | 42223.0427 | 1 | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2 HVR fault, won't reset. Possible fallout from the power dip earlier.
124346 | 42109.5289 | 0.75 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4: Seeing some beam volts jitter.
124586 | 42120.8072 | 1 | LI30 | 51 | 4 | HARDWARE | HW Problem | 30-5 has frequent beam volts drop outs and 721 Amp. errors since morning.
124677 | 42124.0334 | 1 | LI20 | 51 | 4 | HARDWARE | HW Problem | TCAV0 klystron (20-5) mod fault. Unable to reset remotely.
124778 | 42128.7028 | 0.8 | LI13 | 11 | 4 | HARDWARE | HW Problem | 13-1 MOD fault.
124825 | 42130.3568 | 0 | LI28 | 21 | 4 | HARDWARE | HW Problem | Klys 28-2 showing severe phase, amplitude, and BVLT issues.
125094 | 42139.3201 |  | LI20 | 81 | 4 | HARDWARE | HW Problem | L0-B has mod faulted and will not reset remotely
126338 | 42205.5436 | 1 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 has 24V battery fault
126338 | 42205.5436 | 1 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 has 24V battery fault
126338 | 42205.5436 | 1 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 has 24V battery fault
126667 | 42223.2478 | 1 | LI30 | 51 | 4 | HARDWARE | HW Problem | 30-5 frequent AMM faults.
126776 | 42241.5489 | 1 | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 Thy fil xformer AC input lug is burnt
126776 | 42241.5489 | 1 | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 Thy fil xformer AC input lug is burnt
126778 | 42241.5514 | 1 | LI25 | 71 | 4 | HARDWARE | HW Problem | 25-7 Mod has a burnt PFN strap ( Lower, Center jumper )
125257 | 42146.1912 | 1 | LI22 | 51 | 4 | HARDWARE | HW Problem | 22-5 WGV does not reset dispatched PEM
125291 | 42147.9151 |  | LI13 | 61 | 4 | HARDWARE | HW Problem | 13-6 SCRTD EOLC signal present not lit
125445 | 42157.3543 | 3 | LI21 | 61 | 4 | HARDWARE | HW Problem | Informed PEM.
125451 | 42157.4561 | 0.1 | LI09 | 81 | 4 | HARDWARE | HW Problem | Klys 9-8 MOD fault, no reset.
126169 | 42195.9 | 1 | LI28 | 21 | 4 | HARDWARE | HW Problem | Klys 28-2 HVR and trips on reset
126169 | 42195.9 | 1 | LI28 | 21 | 4 | HARDWARE | HW Problem | Klys 28-2 HVR and trips on reset
126169 | 42195.9 | 1 | LI28 | 21 | 4 | HARDWARE | HW Problem | Klys 28-2 HVR and trips on reset
126169 | 42195.9 | 2 | LI28 | 21 | 4 | HARDWARE | HW Problem | Klys 28-2 HVR and trips on reset
126169 | 42195.9 | 1 | LI28 | 21 | 4 | HARDWARE | HW Problem | Klys 28-2 HVR and trips on reset
126455 | 42212.5574 | 0.5 | LI21 | 51 | 4 | HARDWARE | HW Problem | Klys 21-5 MOD and will not reset
126480 | 42213.5171 | 1 | LI09 | 51 | 4 | HARDWARE | HW Problem | 9-5: WG VGC GF and O/P won't clear No work to start before Fiscal year 2016, Oct 1st, 2015
126544 | 42216.4181 |  | LI29 | 11 | 4 | HARDWARE | HW Problem | 29-1 Adjust beam voltage to 335KV and update IKE. Radiation barriers to be returned to previous positions by RP.
126545 | 42216.4199 | 0.5 | LI28 | 51 | 4 | HARDWARE | HW Problem | 28-5 Adjust beam voltage to 335KV and update IKE. Radiation barriers to be returned to previous positions by RP.
126605 | 42220.3096 | 2 | LI29 | 41 | 4 | HARDWARE | HW Problem | 29-4 Modulator, Replace Thyratron. The Thyratron res is at 8+ volts and unstable. See CATER # 126602
124566 | 42118.8017 | 0.75 | LI03 | 41 | 4 | HARDWARE | HW Problem | Klystron 3-4 MOD fault, won't reset.
124580 | 42119.9315 | 0.75 | LI22 | 81 | 4 | HARDWARE | HW Problem | 22-8 MOD fault, won't reset
125118 | 42139.8987 | 2 | LI24 | 81 | 4 | HARDWARE | HW Problem | Kly. LI 24 - 8 is KHD - 20 min.s Passed to PEM.
125938 | 42182.1242 | 1 | LI22 | 71 | 4 | HARDWARE | HW Problem | Klys 22-7 sees sharp 5 degree jumps in klystron temp, faulting on KLT. Looks like a discrete step, seems unphysical. Klixon? See 125831.
125943 | 42183.4499 | 1 | LI13 | 21 | 4 | HARDWARE | HW Problem | 13-2: Replaced failed Ion Pump PS
126125 | 42193.5393 | 1 | LI16 | 41 | 4 | HARDWARE | HW Problem | 16-4 Following PFN straps need to be replaced ( L, LHS, # 3 ) ( Lower, Center strap )
126127 | 42193.5428 | 1 | LI16 | 61 | 4 | HARDWARE | HW Problem | 16-6 K27 Contacts need to be replaced.
126746 | 42237.4614 | 0.25 | LI24 | 11 | 4 | HARDWARE | HW Problem | Downtime: 24-1 Modulator, burnt lug on D Q ing cap.
124603 | 42121.5577 | 0.5 | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 (L1X) getting beam volts dropouts
125155 | 42143.0161 |  | LI24 | 41 | 4 | HARDWARE | HW Problem | 24-4 is tripping repeatedly. Resets remotely each time.
125410 | 42153.7553 | 0.2 | LI18 | 41 | 4 | HARDWARE | HW Problem | Klystron LI18-41 has lost phase reading. PAD phase will show up for one PIOP update, then disappear again. Not sure if this indicates bad PAD, PIOP, or just a loose connection.
126779 | 42241.5528 | 1 | LI25 | 51 | 4 | HARDWARE | HW Problem | 25-5 PFN strap needs to be replaced ( Lower, LHS , #3 )
124432 | 42111.6928 | 0.5 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4 focus current fault
124474 | 42115.5863 | 1 | LI21 | 61 | 4 | HARDWARE | HW Problem | 21-6 Mod T20 Has 2 Network resistors that need to be replaced due to loose connection overheating
124475 | 42115.5887 | 1 | LI21 | 51 | 4 | HARDWARE | HW Problem | 21-5 Mod PFN Coil and strap need to be replaced due to arcing, Strap may need to be cut off (lower, RHS, #3)
124477 | 42115.5946 | 1 | LI20 | 61 | 4 | HARDWARE | HW Problem | 20-6 Mod Kly fil RTM is broken and needs to be replaced CB7 Plexiglass shield ceramic standoff needs to be replaced.
124612 | 42122.005 | 0.75 | LI02 | 31 | 4 | HARDWARE | HW Problem | 2-3 MOD fault, won't reset
124709 | 42124.9112 | 0.75 | LI11 | 61 | 4 | HARDWARE | HW Problem | 11-6 MOD fault, won't reset.
124899 | 42131.8544 | 2 | LI26 | 71 | 4 | HARDWARE | HW Problem | station 26-7 keeps cycling on and off. Probably a vacuum issue.
124974 | 42134.2538 | 1 | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7 HVR faulting every few minutes UPDATE: 23-7 had a huge jitter problem on 5/12 owl shift
124339 | 42109.3764 | 0 | LI26 | 51 | 4 | HARDWARE | HW Problem | 26-5: BVJ and AMJ
124476 | 42115.5909 | 1 | LI20 | 81 | 4 | HARDWARE | HW Problem | 20-8 All T20 Jumper cables are loose and need to be removed cleaned and tightened along with brass fittings.
124520 | 42116.8953 | 2 | LI26 | 21 | 4 | HARDWARE | HW Problem | 26-2 ACOC with no indication as to why. AC current 91A. Other stations running 103A. Time for the power analyzer.
124592 | 42121.3053 | 7 | LI22 | 71 | 4 | HARDWARE | HW Problem | 22-7 MOD fault no reset
124844 | 42130.5655 | 0.25 | LI22 | 31 | 4 | HARDWARE | HW Problem | Klys 22-3 WGV faulted
124954 | 42132.7508 | 1 | LI25 | 11 | 4 | HARDWARE | HW Problem | 25-1 high amplitude jitter.
124971 | 42133.9239 | 1 | LI26 | 21 | 4 | HARDWARE | HW Problem | Called into PEM. Kly. LI 26 - 2 Phase and Amplitude 1.65 / 0.39
125021 | 42136.493 | 0.5 | LI20 | 51 | 4 | HARDWARE | HW Problem | 20-5 (TCAV0): HVR fault
125297 | 42147.9206 |  | LI18 | 41 | 4 | HARDWARE | HW Problem | 18-4 SCRTD EOLC signal present not lit
126231 | 42199.9137 | 1 | LI29 | 31 | 4 | HARDWARE | HW Problem | Klystron 29-3 & 29-4 large phase jitter. As seen on the FTPs. Beam Volts jitter is noticeable but not as large.
126231 | 42199.9137 | 1 | LI29 | 31 | 4 | HARDWARE | HW Problem | Klystron 29-3 & 29-4 large phase jitter. As seen on the FTPs. Beam Volts jitter is noticeable but not as large.
126348 | 42205.7373 | 0.1 | LI28 | 71 | 4 | HARDWARE | HW Problem | 28-7 is only getting 60 Hz accelerate triggers. Revealed (and exacerbated) problems with the fast longitudinal feedback, caused small amounts of downtime.
124565 | 42118.6946 | 0.1 | LI26 | 81 | 4 | HARDWARE | HW Problem | Klystron 26-8 low forward energy peak Forward RF looks low on FTP
124756 | 42127.6296 | 1 | LI26 | 61 | 4 | HARDWARE | HW Problem | 26-6: need to update PLC firmware and install software. Note: The Enerpro and HV regulator boards are modified.
124766 | 42128.0126 | 0.75 | LI27 | 11 | 4 | HARDWARE | HW Problem | Klystron LI27-11 is tripped off and will not reset remotely.
125040 | 42137.9644 | 1.5 | LI13 | 21 | 4 | HARDWARE | HW Problem | Kly. LI 13 - 2 tripped on MOD" fault and won't reset.
125154 | 42142.8816 | 2 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 is tripping repeatedly. Resets remotely each time. Also see CATER#125562 for current issues
125154 | 42142.8816 | 0.1 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 is tripping repeatedly. Resets remotely each time. Also see CATER#125562 for current issues
125280 | 42146.7868 | 0.2 | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X tripped. Comes on, but trips again immediately.
125306 | 42148.6178 | 1 | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2 Check the PIOP ratios and adjust the klystron beam voltage to 350KV
125772 | 42173.4583 | 1 | LI21 | 31 | 4 | HARDWARE | HW Problem | 21-3 occasional phase jumps
125939 | 42182.2073 | 1 | LI22 | 51 | 4 | HARDWARE | HW Problem | 22-5 MOD fault.
126230 | 42199.8912 | 0.1 | LI21 | 51 | 4 | HARDWARE | HW Problem | Klystron 21-5 mod fault.
126232 | 42199.9156 | 2.5 | LI21 | 81 | 4 | HARDWARE | HW Problem | 21-8 RF amplitude jitter FTPs show occasional dropouts Don't see corresponding beam volts dropouts. Maybe this is a readback problem?
126233 | 42199.9166 | 0.5 | LI29 | 81 | 4 | HARDWARE | HW Problem | 29-8 has occasional beam volt jumps
126313 | 42204.692 | 0.2 | LI29 | 31 | 4 | HARDWARE | HW Problem | 29-3 MKSU faulting often
126600 | 42219.6411 | 0.5 | LI22 | 11 | 4 | HARDWARE | HW Problem | Klys 22-1 tripped off and won't reset remotely. See126477.
124453 | 42114.2926 | 2 | LI29 | 11 | 4 | HARDWARE | HW Problem | 29-1 needs PFN tune: PW=0.00uS, Slope= 1.1; Ripl = 0.2
124780 | 42128.8269 | 0.5 | LI22 | 71 | 4 | HARDWARE | HW Problem | station 22-7 reports MOD EVOC Fault but otherwise operates normally
124780 | 42128.8269 | 3 | LI22 | 71 | 4 | HARDWARE | HW Problem | station 22-7 reports MOD EVOC Fault but otherwise operates normally
124780 | 42128.8269 | 1 | LI22 | 71 | 4 | HARDWARE | HW Problem | station 22-7 reports MOD EVOC Fault but otherwise operates normally
124827 | 42130.4027 | 4 | LI23 | 31 | 4 | HARDWARE | HW Problem | Klys 23-3 MOD and will not reset
124833 | 42130.4501 |  | LI25 | 11 | 4 | HARDWARE | HW Problem | 25-1 window 2 temp is very high (133 degrees F)
124957 | 42132.8872 | 0.2 | LI25 | 11 | 4 | HARDWARE | HW Problem | 25-1 MOD fault.
124966 | 42133.6129 |  | LI26 | 21 | 4 | HARDWARE | HW Problem | Has increased since 5 - 8 - 15 at 2000. 0145 to 1.21
124989 | 42135.3589 | 0 | LI23 | 81 | 4 | HARDWARE | HW Problem | Klystron 23-8 mod fault.
124835 | 42130.5075 |  | LI23 | 81 | 4 | HARDWARE | HW Problem | 23-8 is showing fairly constant amplitude jitter. Can see on FTP.
124839 | 42130.5484 | 0.5 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3: Beam volts is low.
124942 | 42132.6224 |  | LI13 | 21 | 4 | HARDWARE | HW Problem | Klystron 13-2 mod fault
124975 | 42134.3498 | 0 | LI21 | 51 | 4 | HARDWARE | HW Problem | 21-5 lots of amplitude and phase jitter
125412 | 42153.8491 | 0.5 | LI17 | 41 | 4 | HARDWARE | HW Problem | 17-4 WGV readback error: The PIOP reads 2.1 e-5 but WVGC reads 4e-7 and doesn't trip the station. Trip point is set to 1.5e-6.
125420 | 42154.6305 | 3 | LI02 | 41 | 4 | HARDWARE | HW Problem | Klystron LI02-41 is tripped off and will not reset remotely.
126078 | 42191.5811 | 3 | LI23 | 21 | 4 | HARDWARE | HW Problem | LCLS MODs: 23-2, 24-3, 24-4, 24-8,28-1, 28-2, 29-2 - need BV adjust to get 350KV while maintaining adequate DQing current.
126078 | 42191.5811 | 1 | LI23 | 21 | 4 | HARDWARE | HW Problem | LCLS MODs: 23-2, 24-3, 24-4, 24-8,28-1, 28-2, 29-2 - need BV adjust to get 350KV while maintaining adequate DQing current.
126078 | 42191.5811 | 1 | LI23 | 21 | 4 | HARDWARE | HW Problem | LCLS MODs: 23-2, 24-3, 24-4, 24-8,28-1, 28-2, 29-2 - need BV adjust to get 350KV while maintaining adequate DQing current.
126078 | 42191.5811 | 1 | LI23 | 21 | 4 | HARDWARE | HW Problem | LCLS MODs: 23-2, 24-3, 24-4, 24-8,28-1, 28-2, 29-2 - need BV adjust to get 350KV while maintaining adequate DQing current.
126078 | 42191.5811 | 1 | LI23 | 21 | 4 | HARDWARE | HW Problem | LCLS MODs: 23-2, 24-3, 24-4, 24-8,28-1, 28-2, 29-2 - need BV adjust to get 350KV while maintaining adequate DQing current.
126078 | 42191.5811 | 1 | LI23 | 21 | 4 | HARDWARE | HW Problem | LCLS MODs: 23-2, 24-3, 24-4, 24-8,28-1, 28-2, 29-2 - need BV adjust to get 350KV while maintaining adequate DQing current.
126397 | 42208.2246 | 1 | LI24 | 81 | 4 | HARDWARE | HW Problem | 24-8 WGV flt
126479 | 42213.5144 | 1 | LI06 | 31 | 4 | HARDWARE | HW Problem | 6-3: WG VGC HV tripped, won't reset. TS&R. No work to start before Fiscal year 2016, Oct 1st, 2015
126490 | 42213.7853 | 2 | LI30 | 41 | 4 | HARDWARE | HW Problem | 30-4 mod fault, firewire alarm
126607 | 42220.4128 | 0.75 | LI29 | 81 | 4 | HARDWARE | HW Problem | 29-8: Seeing some beam volts jitter.
126720 | 42233.5401 |  | LI20 | 51 | 4 | HARDWARE | HW Problem | 20-5 oil leak at the klystron pulse cable
126777 | 42241.5502 | 1 | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 PFN strap needs to be replaced (Lower, LHS, #3 )
124748 | 42127.1164 | 0.5 | LI26 | 61 | 4 | HARDWARE | HW Problem | Klystron 26-6 repeating HVR fault. Also 123625. Able to reset each time. Off the beam for now.
124748 | 42127.1164 | 0.25 | LI26 | 61 | 4 | HARDWARE | HW Problem | Klystron 26-6 repeating HVR fault. Also 123625. Able to reset each time. Off the beam for now.
124750 | 42127.238 | 1 | LI01 | 61 | 4 | HARDWARE | HW Problem | Klystron 1-6 EVC fault MOD EOLC fault. Klystron is on the beam and working. Has RF drive. This may be a false message. IPL the PIOP did not resolve this issue
124903 | 42132.0914 | 1 | LI02 | 31 | 4 | HARDWARE | HW Problem | 2-3 has a WGV fault, can't reset from the SCP.
125089 | 42138.9122 | 0.75 | LI13 | 11 | 4 | HARDWARE | HW Problem | Won't reset.
125215 | 42144.908 | 0.5 | LI20 | 81 | 4 | HARDWARE | HW Problem | 20-8 has core bias at 16.5 amps
126171 | 42195.9758 | 0.5 | LI28 | 71 | 4 | HARDWARE | HW Problem | Klys 28-7 is MOD and won't reset remotely. Tripped a couple of times before this. RE's reported.
126266 | 42201.3109 | 1.5 | LI28 | 11 | 4 | HARDWARE | HW Problem | 28-1 Bad high powered attenuator, replace IPA then resaturate station.
126319 | 42205.081 | 3 | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 is HVR and won't reset remotely
126377 | 42206.6135 | 1 | LI21 | 51 | 4 | HARDWARE | HW Problem | 21-5 mod faulted and won't reset remotely
126488 | 42213.6993 | 2 | LI29 | 41 | 4 | HARDWARE | HW Problem | 29-4 cycles frequently
126565 | 42217.5087 | 1 | LI27 | 31 | 4 | HARDWARE | HW Problem | Tried to bring back on 3 times but it didn't come back.
126674 | 42224.8931 | 0.5 | LI23 | 51 | 4 | HARDWARE | HW Problem | Klys 23-5 tripped on KMC.
124959 | 42133.36 | 1 | LI13 | 21 | 4 | HARDWARE | HW Problem | 13-2 MOD fault, won't reset
124991 | 42135.4039 | 1 | LI06 | 31 | 4 | HARDWARE | HW Problem | FACET - Beam Volt adjustment for MODs: 6-3, 13-2, and 20-4B
125005 | 42136.1386 |  | LI05 | 11 | 4 | HARDWARE | HW Problem | 5-1 has a CBL fault. IPLing the PIOP did not fix this.
125282 | 42147.0992 |  | LI22 | 81 | 4 | HARDWARE | HW Problem | Will not reset. Informed PEM.
125611 | 42164.4801 |  | LI28 | 11 | 4 | HARDWARE | HW Problem | 28-1 MOD HVR and won't reset remotely
126094 | 42192.2671 | 1 | LI15 | 61 | 4 | HARDWARE | HW Problem | 15-6 K26 and K27 Contacts Need to be replaced. T20 output network resistor is broken. PFN strap ( upper, RHS, # 3 )
126106 | 42192.4898 | 0.5 | LI04 | 31 | 4 | HARDWARE | HW Problem | 4-3 SLED motor fault.
126134 | 42193.8625 | 2 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 core bias voltage reading 88 volts, should be around 4.
126157 | 42195.5894 | 2 | LI18 | 21 | 4 | HARDWARE | HW Problem | 18-2 K26 Contacts need to be replaced CB7 Plexiglas shield is broken
126158 | 42195.5907 | 1 | LI18 | 31 | 4 | HARDWARE | HW Problem | 18-3 K26 contacts need to be replaced
126160 | 42195.5937 | 1 | LI18 | 61 | 4 | HARDWARE | HW Problem | 18-6 K27 contacts need to be replaced
126161 | 42195.5954 | 1 | LI18 | 71 | 4 | HARDWARE | HW Problem | 18-7 K 26 and K27 contacts need to be replaced
126250 | 42200.5782 | 4 | LI21 | 21 | 4 | HARDWARE | HW Problem | PO and beam voltage dropout when the chassis is tapped.
126250 | 42200.5782 | 4 | LI21 | 21 | 4 | HARDWARE | HW Problem | PO and beam voltage dropout when the chassis is tapped.
126257 | 42200.6537 |  | LI30 | 11 | 4 | HARDWARE | HW Problem | 30-1: Investigate wide PW on PFN
125188 | 42144.3898 | 2.5 | LI23 | 41 | 4 | HARDWARE | HW Problem | Kly Mod 23-4: add oil to T20 transformer for safe HV ops. LOTO station, remove pressure relief valve on top of T20 case. Add Sontex oil until oil level above top 480 volt feedthru electrodes.
125296 | 42147.9198 |  | LI17 | 81 | 4 | HARDWARE | HW Problem | 17-8 SCRTD EOLC signal present not li
125298 | 42147.9218 |  | LI20 | 41 | 4 | HARDWARE | HW Problem | 20-4B SCRTD EOLC signal present not lit
127649 | 42285.5985 | 1 | LI21 | 21 | 4 | HARDWARE | HW Problem | Duplicate see: 127622. Channel 0 on the L1X feedback PAD is broken. The measured amplitude does not correlate to the structure temperature. Channel 1 looks fine, and correlates with temperature.
127693 | 42286.617 |  | LI02 | 31 | 4 | HARDWARE | HW Problem | Duplicate see 127690. Klystron LI02-31 has lost PAD phase. Amplitude reads very high. Tried IPLing PIOP with no change in status.
127694 | 42286.6188 |  | LI02 | 41 | 4 | HARDWARE | HW Problem | Duplicate see 127691. Klystron LI02-41 is showing zero PAD amplitude and lost PAD phase.
127698 | 42286.6272 | 1 | DR13 | 1 | 4 | HARDWARE | HW Problem | 1-8 Numerous MKSU faults, No 24 Volts present?
127702 | 42286.6363 | 1 | LI28 | 31 | 4 | HARDWARE | HW Problem | 29-3 tripping frequently "MKT"
127710 | 42286.7834 |  | LI15 | 71 | 4 | HARDWARE | HW Problem | Duplicate see 127667. 15-7 possible bad PIOP.
127715 | 42286.8235 | 1 | LI10 | 61 | 4 | HARDWARE | HW Problem | 10-6 AMM low beam voltage.
127531 | 42280.4777 | 3 | LI30 | 11 | 4 | HARDWARE | HW Problem | Klystron 30-1 is tripping off the beam constantly. Resets, but rapidly faults.
127759 | 42287.9439 | 0.5 | LI17 | 11 | 4 | HARDWARE | HW Problem | 17-1 has low beam voltage
128146 | 42300.9431 | 1 | LI20 | 81 | 4 | HARDWARE | HW Problem | L0B is HVR and won't reset remotely. This happened after the Linac East PPS went to state mismatch, but the PPS permits should be back on this and other injector stations
128235 | 42306.0006 |  | LI25 | 51 | 4 | HARDWARE | HW Problem | Klystron LI25-51 is tripping off frequently. Errors panel shows several vacuum faults.
128289 | 42309.3119 | 1 | LI24 | 51 | 4 | HARDWARE | HW Problem | 24-5 has high phase jitter. Thyratron EOL reported.
127735 | 42287.4408 | 0 | LI22 | 31 | 4 | HARDWARE | HW Problem | Klystron 22-3 is exhibiting a lot of phase jitter, which is negatively impacting the FEL. We have taken this station off of the beam for now.
127736 | 42287.4438 | 2 | LI29 | 81 | 4 | HARDWARE | HW Problem | 29-8 Modulator, Awaiting Thyratron This station is exhibiting a lot of amplitude and phase jitter, which is negatively affecting the FEL. We have taken it off the beam for now.
127736 | 42287.4438 | 1 | LI29 | 81 | 4 | HARDWARE | HW Problem | 29-8 Modulator, Awaiting Thyratron This station is exhibiting a lot of amplitude and phase jitter, which is negatively affecting the FEL. We have taken it off the beam for now.
127736 | 42287.4438 | 1 | LI29 | 81 | 4 | HARDWARE | HW Problem | 29-8 Modulator, Awaiting Thyratron This station is exhibiting a lot of amplitude and phase jitter, which is negatively affecting the FEL. We have taken it off the beam for now.
127740 | 42287.559 | 1 | LI14 |  | 4 | HARDWARE | HW Problem | 14-2, 14-4, 14-5, 14-6 and 14-8 have low beam voltage.
127748 | 42287.6815 | 2 | LI27 | 41 | 4 | HARDWARE | HW Problem | 27-4 mod fault
127779 | 42288.4941 | 1 | LI05 | 71 | 4 | HARDWARE | HW Problem | 5-7 MKSU bad voltages on PIOP.
127780 | 42288.4966 | 0.5 | LI12 | 41 | 4 | HARDWARE | HW Problem | 12-4 waveguide vacuum fault
127796 | 42289.1481 | 0.5 | LI14 | 41 | 4 | HARDWARE | HW Problem | 14-4 is FLK
127825 | 42289.637 | 0.5 | LI27 | 31 | 4 | HARDWARE | HW Problem | 27-3 Modulator, Thyratron is dead. Replace tube when spares inventory increases beyond 2. PMS
127863 | 42290.762 | 0.75 | LI13 | 11 | 4 | HARDWARE | HW Problem | Kly. LI13 - 1 faulted on MOD fault.
127878 | 42290.9835 | 1 | LI05 | 71 | 4 | HARDWARE | HW Problem | Kly. LI05 - 7 PAD "phase' and Amplit. read back is now bad Noted problem after doing a Facet phase scan. Beam still looks ok but the PAD. is bad.
127878 | 42290.9835 | 0.5 | LI05 | 71 | 4 | HARDWARE | HW Problem | Kly. LI05 - 7 PAD "phase' and Amplit. read back is now bad Noted problem after doing a Facet phase scan. Beam still looks ok but the PAD. is bad.
128049 | 42297.8778 | 0.75 | LI13 | 11 | 4 | HARDWARE | HW Problem | MOD fault on klystron 13-1
128223 | 42305.5974 | 1 | LI02 | 81 | 4 | HARDWARE | HW Problem | 2-8 PFN compartment cooling fan is dead
127807 | 42289.2931 | 0.75 | LI18 | 81 | 4 | HARDWARE | HW Problem | 18-8 MOD Fault
127808 | 42289.2954 | 0.5 | LI17 | 61 | 4 | HARDWARE | HW Problem | 17-6 has a KMC fault
127820 | 42289.5914 |  | LI29 | 31 | 4 | HARDWARE | HW Problem | Klys 29-3 tripps on MKT every few minutes. Can not get it to stay on continuously.
127824 | 42289.6356 |  | LI28 | 21 | 4 | HARDWARE | HW Problem | Klys 28-2 keeps going WA2. Putting it into MNT for tonight to remind ops not to use it.
127841 | 42290.3097 | 0.75 | LI03 | 81 | 4 | HARDWARE | HW Problem | station 3-8 mod faulted
127853 | 42290.5899 | 0.8 | LI02 | 11 | 4 | HARDWARE | HW Problem | Klystron 2-1 has mod faulted and will not reset remotely
127857 | 42290.7031 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X work- replacing multipler and local oscillator chassis.
127858 | 42290.7113 | 1 | LI01 | 51 | 4 | HARDWARE | HW Problem | Kly. LI01 - 5 is tripping on Kly. OV.
127860 | 42290.7428 | 0.75 | LI10 | 71 | 4 | HARDWARE | HW Problem | ELECTRO-MAGNET Current. MKSU fault.
126846 | 42250.4887 | 1 | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 Klystron reflected power connecter is loose from the shielding at the klystron end.
126983 | 42264.2623 | 1 | LI22 | 21 | 4 | HARDWARE | HW Problem | 22-2 PFN cap needs replace. RHS 4th down from top. Coil and straps side of caps.
126983 | 42264.2623 | 1 | LI22 | 21 | 4 | HARDWARE | HW Problem | 22-2 PFN cap needs replace. RHS 4th down from top. Coil and straps side of caps.
127176 | 42270.5375 | 2 | LI21 | 11 | 4 | HARDWARE | HW Problem | Kly Mod 21-1 (L1S): when HVPS voltage raised above 20.3 KV, modulator T20 begins to "motorboat". This indicates Enerpro board needs repair/adjustment. Regulator board checks good.
127288 | 42273.3208 | 0.5 | LI02 | 81 | 4 | HARDWARE | HW Problem | 2-8 MKSU has a battery fault
127387 | 42276.4434 | 2.5 | LI20 | 61 | 4 | HARDWARE | HW Problem | 20-6 high voltage has timing jitter. Timing jumps measured at SCRTD on order of 10 ns. Thyratron reliably follows timing jumps. Suspect SCRTD
127495 | 42279.4046 | 1 | LI05 | 21 | 4 | HARDWARE | HW Problem | Klystron water fault
127544 | 42281.3977 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X TWT spiky jitter, phase jumps up to 4 deg
127550 | 42281.8478 | 0.5 | LI21 | 41 | 4 | HARDWARE | HW Problem | Klys 21-4 MKSU Trigger enable fault.
127861 | 42290.7498 | 0.75 | LI12 | 31 | 4 | HARDWARE | HW Problem | Kly. LI12 - 3 faulted on ELECTRO-MAGNET fault.
127862 | 42290.7545 | 1.5 | LI10 | 51 | 4 | HARDWARE | HW Problem | Kly. LI12 - 5 faulted on ELECTRO-MAGNET fault.
127864 | 42290.7641 | 0.5 | LI13 | 61 | 4 | HARDWARE | HW Problem | Kly. LI13 - 6 faulted on MOD fault.
127866 | 42290.7692 |  | LI14 | 11 | 4 | HARDWARE | HW Problem | duplicate see 127840. Kly. LI14 - 1 faulted on MOD fault.
128572 | 42320.1133 | 0.5 | LI23 | 51 | 4 | HARDWARE | HW Problem | 23-5 tripped HVR no remote reset station in local
128579 | 42320.3858 | 0.25 | LI26 | 71 | 4 | HARDWARE | HW Problem | 26-7 getting beam volts calibration. Cater for Phillip.
128579 | 42320.3858 | 0.5 | LI26 | 71 | 4 | HARDWARE | HW Problem | 26-7 getting beam volts calibration. Cater for Phillip.
128579 | 42320.3858 | 1 | LI26 | 71 | 4 | HARDWARE | HW Problem | 26-7 getting beam volts calibration. Cater for Phillip.
128585 | 42320.4173 | 1 | LI09 | 21 | 4 | HARDWARE | HW Problem | Klystron LI09-21 is tripped off and will not reset remotely.
127296 | 42273.765 | 1 | LI30 | 71 | 4 | HARDWARE | HW Problem | 30-7 WKL. Klystron shows 2481 water faults for last 24 hrs.
127418 | 42277.5459 | 0.5 | LI20 | 94 | 4 | HARDWARE | HW Problem | PPS inhibit lights on the Modulator 20-4A does not appear to be working. When we cycle the PPS state where the lights should be inhibited and light up, they do not. I am not sure if that requires modulator power to function, or if it is powered by 24V power.
127418 | 42277.5459 | 1 | LI20 | 94 | 4 | HARDWARE | HW Problem | PPS inhibit lights on the Modulator 20-4A does not appear to be working. When we cycle the PPS state where the lights should be inhibited and light up, they do not. I am not sure if that requires modulator power to function, or if it is powered by 24V power.
127418 | 42277.5459 | 2 | LI20 | 94 | 4 | HARDWARE | HW Problem | PPS inhibit lights on the Modulator 20-4A does not appear to be working. When we cycle the PPS state where the lights should be inhibited and light up, they do not. I am not sure if that requires modulator power to function, or if it is powered by 24V power.
127768 | 42288.2509 | 1 | LI19 | 11 | 4 | HARDWARE | HW Problem | 19-1 has a KMC fault and won't reset
127771 | 42288.3363 | 2 | LI25 | 11 | 4 | HARDWARE | HW Problem | 25-1 is MOD faulted and won't reset - thyratron is dead
127867 | 42290.7726 | 0.5 | LI17 | 61 | 4 | HARDWARE | HW Problem | Kly. LI17 - 6 faulted on ELECTRO_MAGNET MOD fault.
127959 | 42293.9282 | 1 | LI25 | 11 | 4 | HARDWARE | HW Problem | 25-1 mod faults, EOLC faults
127959 | 42293.9282 | 2 | LI25 | 11 | 4 | HARDWARE | HW Problem | 25-1 mod faults, EOLC faults
127871 | 42290.8061 |  | LI02 | 31 | 4 | HARDWARE | HW Problem | duplicate see 127690. Kly. LI02 - 3 "Lost" Phase and can't get it back.
127872 | 42290.8084 |  | LI02 | 41 | 4 | HARDWARE | HW Problem | Duplicate see: 127691. Kly. LI02 - 4 "Lost" Phase and can't get it back.
127874 | 42290.842 | 0.75 | LI02 | 11 | 4 | HARDWARE | HW Problem | 2-1 mod fault, can't reset remotely.
127875 | 42290.9108 |  | LI12 | 61 | 4 | HARDWARE | HW Problem | duplicate see 127709. KLY. LI 12 - 6 PAD DAC status seems to have failed. PAD DAC status = "FALSE".
127876 | 42290.9575 | 0.5 | LI17 | 41 | 4 | HARDWARE | HW Problem | 17-4 sounds like is operating at an unsteady 15 hertz
127881 | 42291.0025 |  | LI02 | 41 | 4 | HARDWARE | HW Problem | duplicate see 127691. Kly. LI02 - 4 PAD "phase' and Amplit. read back is bad.
127886 | 42291.3687 | 1 | LI19 | 11 | 4 | HARDWARE | HW Problem | 19-1 focus magnet power supply has no output
127899 | 42291.6165 | 0.5 | LI23 | 11 | 4 | HARDWARE | HW Problem | 23-1 HVPS tripping frequently couple times per hour.
127902 | 42291.6658 | 1 | LI15 | 71 | 4 | HARDWARE | HW Problem | station 15-7 has KMC fault
127909 | 42291.7698 |  | LI05 | 71 | 4 | HARDWARE | HW Problem | DUPLICATE of 127878: 5-7 accelerate and modulator trigger rates at 28, expected is 30.
127912 | 42291.8246 | 0.75 | LI10 | 21 | 4 | HARDWARE | HW Problem | klystron 10-2 mod faults and resets ok, but then mod faults immediately afterwards.
126876 | 42255.3146 | 1 | LI27 | 71 | 4 | HARDWARE | HW Problem | 27-7 has a Klystron water fault
127135 | 42269.7157 |  | LI30 | 41 | 4 | HARDWARE | HW Problem | 30-4 has bouts of unstability on the klystron heater current and voltage. I checked it current should be and was 3.01A, the voltage was 17.4 should be 11.7V. The voltage has not been correct since the control has been turned back on. duplicate: 127078
127460 | 42278.2331 | 2 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 tripping several times per minute
127764 | 42288.1079 | 0.5 | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7 MOD faults frequently. Has faulted ~once every fifteen minutes this shift.
127781 | 42288.498 | 1 | LI12 | 31 | 4 | HARDWARE | HW Problem | 12-3 klystron water fault
127800 | 42289.1568 | 1 | LI11 | 71 | 4 | HARDWARE | HW Problem | 11-7 is FLK
127939 | 42292.7392 | 1 | LI15 | 31 | 4 | HARDWARE | HW Problem | 15-3 Ampl FTP appears saturated - top of pulse is flat-topped at 100%. Needs calibration?
127939 | 42292.7392 | 0.7 | LI15 | 31 | 4 | HARDWARE | HW Problem | 15-3 Ampl FTP appears saturated - top of pulse is flat-topped at 100%. Needs calibration?
127956 | 42293.551 | 1 | LI05 | 31 | 4 | HARDWARE | HW Problem | 5-3 Mcc reported Vac fault will not clear
127963 | 42294.6036 | 0.5 | LI23 | 31 | 4 | HARDWARE | HW Problem | Station 23-3 tripping frequently on HVR. Resets remotely but is a frequent offender.
126875 | 42255.3138 | 1 | LI27 | 81 | 4 | HARDWARE | HW Problem | 27-8 has a Klystron water fault
126901 | 42257.2167 | 1 | LI27 | 21 | 4 | HARDWARE | HW Problem | 26-2 Klystron Water fault light intermittently flashes
126907 | 42257.4992 | 1 | LI21 | 11 | 4 | HARDWARE | HW Problem | 21-2 reduce PFN from 5 stages to 3 stages. (2 parallel of 5 stages to 2 parallel of 3 stages).
126907 | 42257.4992 | 2 | LI21 | 11 | 4 | HARDWARE | HW Problem | 21-2 reduce PFN from 5 stages to 3 stages. (2 parallel of 5 stages to 2 parallel of 3 stages).
127324 | 42275.1608 | 1 | LI12 | 21 | 4 | HARDWARE | HW Problem | 12-2 has an ACC2 water fault
127369 | 42276.0799 |  | LI30 | 41 | 4 | HARDWARE | HW Problem | Klystron LI30-41 trips off frequently (on the order of minutes); shows over voltage and drive faults.
127873 | 42290.8184 | 0.75 | LI10 | 21 | 4 | HARDWARE | HW Problem | Kly. LI10 - 2 keeps tripping on "MOD" fault.
127916 | 42292.2242 |  | LI18 | 31 | 4 | HARDWARE | HW Problem | 18-3 MOD fault, no remote reset.
127996 | 42296.6028 | 0.5 | LI10 | 41 | 4 | HARDWARE | HW Problem | station 10-4 MOD fault
127964 | 42294.658 |  | LI20 | 71 | 4 | HARDWARE | HW Problem | Time Accounting CATER L0A (klystron 20-7) has tripped on HVR or MOD at least four times on day shift (10/17/15). OPS were able to reset the station remotely each time, but the reset did coincide with a brief vacuum burst in the injector on one occasion. This has caused about 0.2 hours of downtime today (split between XPP and AMO users).
127966 | 42294.6903 | 0.5 | LI23 | 11 | 4 | HARDWARE | HW Problem | Klys 23-1 HVR, will not come on remotely, in local?
127969 | 42295.2528 | 0.5 | LI25 | 51 | 4 | HARDWARE | HW Problem | 25-5 HVR fault, no remote reset.
127988 | 42296.4746 | 0.5 | LI01 | 51 | 4 | HARDWARE | HW Problem | stations 1-5 and 1-6 are reporting very low amplitude
127992 | 42296.5286 | 0.5 | LI11 | 81 | 4 | HARDWARE | HW Problem | 11-8 is MOD faulted and does not reset remotely.
127993 | 42296.5292 | 0.5 | LI13 | 41 | 4 | HARDWARE | HW Problem | 13-4 is MOD faulted and does not reset remotely.
127636 | 42284.9245 | 1 | LI20 | 93 | 4 | HARDWARE | HW Problem | 20-3C T20 cooling fan inoperative
128150 | 42301.3737 | 1 | LI22 | 71 | 4 | HARDWARE | HW Problem | Klys 22-7 showing some beam volts jitter.
128152 | 42301.4745 | 1 | LI23 | 61 | 4 | HARDWARE | HW Problem | Klys 23-6 HVR. Will not come on remotely.
126793 | 42243.3528 | 1 | LI07 | 61 | 4 | HARDWARE | HW Problem | 7-6 Disconnect T20 , HV rectifier stack, disconnect Snubber resistors, and fire wire.
126793 | 42243.3528 | 3 | LI07 | 61 | 4 | HARDWARE | HW Problem | 7-6 Disconnect T20 , HV rectifier stack, disconnect Snubber resistors, and fire wire.
126874 | 42255.3128 | 1 | LI28 | 51 | 4 | HARDWARE | HW Problem | 28-5 has a Klystron water fault
126899 | 42257.2124 | 1 | LI27 | 11 | 4 | HARDWARE | HW Problem | 27-1 Klystron Water fault light intermittently flashes
126903 | 42257.344 | 1 | LI06 | 61 | 4 | HARDWARE | HW Problem | 6-6 has a klystron vacuum fault.
126845 | 42250.4674 | 0.5 | LI27 | 41 | 4 | HARDWARE | HW Problem | 27-4: Klystron Magnet Trim PS out of regulation. Meter stuck.
127323 | 42275.1595 | 1 | LI18 | 61 | 4 | HARDWARE | HW Problem | 18-6 has a Wave Guide Vacuum fault
127538 | 42280.9724 | 0.5 | LI28 | 51 | 4 | HARDWARE | HW Problem | Klys 28-5 MOD fault.
128027 | 42297.5163 |  | LI01 | 51 | 4 | HARDWARE | HW Problem | Dupe. See 127850. Klystron 1-5 was found 30 degrees and trimming does not trim to the expected value. The fox phase shifter also appears to be stuck since about 9am this morning.
128029 | 42297.5211 | 2 | LI28 | 11 | 4 | HARDWARE | HW Problem | 28-1: Off on core bias undercurrent. Drooped to zero yesterday afternoon.
128089 | 42299.0174 |  | LI02 | 41 | 4 | HARDWARE | HW Problem | Duplicate see 127691. Kly LI 02 - 4 has a problem with the PAD Phase, Mixer Gain, AMPLIT. and Phase Mode all are red and not reading back.
128090 | 42299.0264 | 1 | LI25 | 11 | 4 | HARDWARE | HW Problem | 25-1 instantly trips off. Water klystron fault flashes briefly.
127209 | 42271.3449 | 0.5 | LI22 | 71 | 4 | HARDWARE | HW Problem | 22-7 klystron MKT fault, no remote reset.
127210 | 42271.362 |  | LI29 | 11 | 4 | HARDWARE | HW Problem | 29-1 trips off frequently, can reset
127251 | 42272.1012 | 1 | LI24 | 51 | 4 | HARDWARE | HW Problem | Klystron LI24-51 has accumulated 497 water faults on klystron errors panel.
127266 | 42272.4519 | 1 | LI26 | 51 | 4 | HARDWARE | HW Problem | KLYS:LI26:51:WNDW2_TEMP = 141.8 deg.s F.
127266 | 42272.4519 |  | LI26 | 51 | 4 | HARDWARE | HW Problem | KLYS:LI26:51:WNDW2_TEMP = 141.8 deg.s F.
127547 | 42281.669 | 0.2 | LI26 | 11 | 4 | HARDWARE | HW Problem | 26-1: KMC fault.
127709 | 42286.779 | 0.5 | LI12 | 61 | 4 | HARDWARE | HW Problem | 12-6 PAD readbacks are hosed.
127904 | 42291.6919 |  | LI13 | 81 | 4 | HARDWARE | HW Problem | Klys 13-8 beam volts is low, hoping to raise it up to ~328 kV.
128058 | 42298.2747 |  | LI20 | 41 | 4 | HARDWARE | HW Problem | FACET XTCAV not able to reach nominal amplitude. Was making up 35 MV yesterday but fell off to 25 MV between 19:00 and midnight. Appears to be a timing issue, AMRF recommends re-timing. 1/12/15 Steele: Able to reach 34.5MV. Beam voltage is about 1KV low. We will increase it.
128058 | 42298.2747 |  | LI20 | 41 | 4 | HARDWARE | HW Problem | FACET XTCAV not able to reach nominal amplitude. Was making up 35 MV yesterday but fell off to 25 MV between 19:00 and midnight. Appears to be a timing issue, AMRF recommends re-timing. 1/12/15 Steele: Able to reach 34.5MV. Beam voltage is about 1KV low. We will increase it.
128080 | 42298.7333 | 0.5 | LI24 | 31 | 4 | HARDWARE | HW Problem | Klys 24-3 tripping off frequently after POMM on 10/21.
128080 | 42298.7333 |  | LI24 | 31 | 4 | HARDWARE | HW Problem | Klys 24-3 tripping off frequently after POMM on 10/21.
128080 | 42298.7333 | 1 | LI24 | 31 | 4 | HARDWARE | HW Problem | Klys 24-3 tripping off frequently after POMM on 10/21.
128086 | 42298.9265 | 1 | LI13 | 11 | 4 | HARDWARE | HW Problem | 13-1 HV meter not indicating
126933 | 42261.2703 | 1 | LI29 | 61 | 4 | HARDWARE | HW Problem | 29-6 has a PFN cap that appears to leaking oil at the bottom seam
126963 | 42263.2848 | 4 | LI27 | 41 | 4 | HARDWARE | HW Problem | 27-4 on the PLC interlface board, F12 (internal 24v) fuse is blown and won't reset/clear.
126997 | 42265.2815 | 1 | LI25 | 11 | 4 | HARDWARE | HW Problem | 25-1 PFN cap needs replace. LHS 1st on top.
127003 | 42265.5215 | 0.5 | LI22 | 71 | 4 | HARDWARE | HW Problem | 22-7: Fault panel display shows "Intlck Complt OK" and "VVS OK" with the sector VVS off. That is not right and it is unlike every other modulator in the LCLS machine when the VVS's are off.
128050 | 42298.0572 | 1 | LI25 | 51 | 4 | HARDWARE | HW Problem | KLY. LI 25 - 5 tripped on "MOD" fault Says it's ready but fails to reset.
128055 | 42298.2167 |  | LI18 | 41 | 4 | HARDWARE | HW Problem | Duplicate see 127789. 18-4 PAD started reading NAN phase in the middle of a phase scan. Also the +15V is OOT.
128071 | 42298.4994 | 1 | LI12 | 21 | 4 | HARDWARE | HW Problem | Klys 12-2 has some bad AMPL jitter.
128093 | 42299.0452 | 1 | LI12 | 41 | 4 | HARDWARE | HW Problem | KLY. LI 12 - 4 tripped on "WGV" and won't reset.
128105 | 42299.4047 |  | LI20 | 71 | 4 | HARDWARE | HW Problem | Time Accounting: L0A faulting on waveguide vacuum (WGV). Bypassed the gauge's low trip threshold. Suspect faulty VGC and swapped in a new unit. (DT OS 10/22 -2.2hr)
126821 | 42248.6552 | 1 | LI22 | 51 | 4 | HARDWARE | HW Problem | 22-5 Modulator, Had oil on the floor of T20 cabinet and PFN cabinet it has been clean up but need to check on next pamm if it leaks from T20 cabinet and PFN cabinet
126908 | 42257.5109 |  | LI24 | 11 | 4 | HARDWARE | HW Problem | 24-1: Waveguide #1 water fault on the MKSU.
127002 | 42265.4729 |  | LI04 | 11 | 4 | HARDWARE | HW Problem | Water leak reported in Control rack near klystron LI04-11.
127716 | 42286.8274 | 0 | LI12 | 31 | 4 | HARDWARE | HW Problem | 12-3 klystron water fault.
127895 | 42291.5545 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X RF level checks after LO replacement
127911 | 42291.8173 | 0.5 | LI16 | 81 | 4 | HARDWARE | HW Problem | 16-8 MKSU has a water summary fault.
128091 | 42299.0323 | 1 | LI25 | 51 | 4 | HARDWARE | HW Problem | 25-5 has tripped off and will not come back on remotely.
128091 | 42299.0323 | 1 | LI25 | 51 | 4 | HARDWARE | HW Problem | 25-5 has tripped off and will not come back on remotely.
127019 | 42266.2819 | 1 | LI24 | 51 | 4 | HARDWARE | HW Problem | 24-5 Loose wires needs securing
127044 | 42267.2268 |  | LI29 | 41 | 4 | HARDWARE | HW Problem | 29-4 is faulted on Klystron Water
128106 | 42299.4199 |  | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6: HVR faults
128110 | 42299.4945 | 0 | LI29 | 61 | 4 | HARDWARE | HW Problem | station 29-6 Mod faulted
128128 | 42299.9419 | 0.75 | LI11 | 61 | 4 | HARDWARE | HW Problem | Klystron LI11-61 is tripped off and will not reset remotely.
128145 | 42300.7204 |  | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 PTF: Slope, ripple and PW = 0
128162 | 42303.0797 | 0.5 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4 HVR fault, won't reset.
128165 | 42303.3166 | 0.5 | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7 MOD fault doesn't reset remotely.
128165 | 42303.3166 | 0.5 | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7 MOD fault doesn't reset remotely.
128165 | 42303.3166 | 1 | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7 MOD fault doesn't reset remotely.
128165 | 42303.3166 | 0.25 | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7 MOD fault doesn't reset remotely.
127143 | 42269.9376 | 0.1 | LI24 |  | 4 | HARDWARE | HW Problem | 24-1, 2 &3 don't have any Amplitude and don't appear to have any drives. Cycled power on PACs. No help.
127316 | 42275.1518 | 0.5 | LI13 | 21 | 4 | HARDWARE | HW Problem | 13-2 Fire Alarm Latching Box "System Ready" green LED is out.
127372 | 42276.169 | 0.75 | LI09 | 31 | 4 | HARDWARE | HW Problem | 9-3 Fire Alarm Latching Box "System Ready" green LED is out.
127752 | 42287.765 |  | LI05 | 71 | 4 | HARDWARE | HW Problem | Duplicate see 127779. Klystron 5-7 has bad PAD voltage readbacks
128109 | 42299.4834 |  | LI25 | 21 | 4 | HARDWARE | HW Problem | 25-2 Klystron switch water faults
128149 | 42301.2415 | 1 | LI21 | 81 | 4 | HARDWARE | HW Problem | 21-8 has lots of amplitude jitter.
127097 | 42268.666 | 1 | LI30 | 81 | 4 | HARDWARE | HW Problem | 30-8: Station "motorboating". Needs Enerpro/HV Regulator checked out.
127138 | 42269.7435 | 0 | LI30 | 61 | 4 | HARDWARE | HW Problem | 30-6: Thyratron Filament Under/current.
127160 | 42270.4209 | 1 | LI25 | 41 | 4 | HARDWARE | HW Problem | 25-4: Intermittent Klystron Water fault.
127201 | 42271.1564 |  | LI29 | 41 | 4 | HARDWARE | HW Problem | 29-4 has no drive.
127201 | 42271.1564 | 1 | LI29 | 41 | 4 | HARDWARE | HW Problem | 29-4 has no drive.
127205 | 42271.3066 | 1 | LI24 | 51 | 4 | HARDWARE | HW Problem | Klystron LI24-51 has water problems. 575 Water faults indicated on error display for last 24 hrs. Klystron is currently faulted and not assigned to beam.
127219 | 42271.4186 | 0.5 | LI21 | 41 | 4 | HARDWARE | HW Problem | 21-4 MOD fault. No remote reset.
128194 | 42304.4158 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | Time Accounting: LCLS XCAV mod KLYS 21-2 trip (DT DS 10/26 -0.2hr)
128219 | 42305.5112 | 0.5 | LI19 | 11 | 4 | HARDWARE | HW Problem | 19-1 Fails NIRP - the overpressure does not trip the station.
128225 | 42305.6004 | 1 | LI09 | 71 | 4 | HARDWARE | HW Problem | 9-7 PFN Compartment fan is dead
128237 | 42306.0689 | 0 | LI27 | 41 | 4 | HARDWARE | HW Problem | Klystron 27-4 continually tripping off on HVR fault
128247 | 42306.4711 | 1 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3 needs TS from local into "REMOTE"
128247 | 42306.4711 | 1 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3 needs TS from local into "REMOTE"
128250 | 42306.478 | 0.5 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 needs TS from local into "REMOTE"
128250 | 42306.478 | 0.5 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 needs TS from local into "REMOTE"
128250 | 42306.478 | 2 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 needs TS from local into "REMOTE"
128250 | 42306.478 | 1 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 needs TS from local into "REMOTE"
127805 | 42289.2767 | 8 | LI25 | 41 | 4 | HARDWARE | HW Problem | 25-4 repeatedly tripping on klystron heater fault
128157 | 42301.6205 | 3 | LI23 | 11 | 4 | HARDWARE | HW Problem | Kly Mod 23-1: DQ volts waveform not present at front panel BNC monitor jack. Investigate and correct problem during next PAMM.
128161 | 42302.8603 | 1 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4 is tripped; unable to reset remotely.
128161 | 42302.8603 | 1 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4 is tripped; unable to reset remotely.
128220 | 42305.526 | 0 | LI22 | 71 | 4 | HARDWARE | HW Problem | 22-7 BVLT jitter OOT
128236 | 42306.0025 | 4 | LI27 | 41 | 4 | HARDWARE | HW Problem | Klystron LI27-41 trips off immediately upon remote resetting. PIOP shows "Mod EVOC" faulting.
128246 | 42306.4615 | 0.5 | LI23 | 61 | 4 | HARDWARE | HW Problem | 23-6 needs the HV, BV and DQ calibration
128246 | 42306.4615 | 1.5 | LI23 | 61 | 4 | HARDWARE | HW Problem | 23-6 needs the HV, BV and DQ calibration
128246 | 42306.4615 | 1.5 | LI23 | 61 | 4 | HARDWARE | HW Problem | 23-6 needs the HV, BV and DQ calibration
128246 | 42306.4615 | 0.2 | LI23 | 61 | 4 | HARDWARE | HW Problem | 23-6 needs the HV, BV and DQ calibration
128260 | 42307.3832 | 0.5 | LI28 | 61 | 4 | HARDWARE | HW Problem | 28-6 HVR, no remote reset
128264 | 42307.5182 | 1 | LI11 | 61 | 4 | HARDWARE | HW Problem | 11-6 has mod faulted and will not reset remotely.
128268 | 42307.6556 | 0.75 | LI21 | 51 | 4 | HARDWARE | HW Problem | 21-5 Mod fault won't reset remotely
126872 | 42255.3083 | 1.5 | LI26 | 51 | 4 | HARDWARE | HW Problem | 26-5 Modulator, No Klystron filament, thyratron filament or heater power.
126872 | 42255.3083 | 1 | LI26 | 51 | 4 | HARDWARE | HW Problem | 26-5 Modulator, No Klystron filament, thyratron filament or heater power.
127243 | 42271.8018 | 1 | LI26 | 61 | 4 | HARDWARE | HW Problem | 26-6 KMC fault
127243 | 42271.8018 | 1 | LI26 | 61 | 4 | HARDWARE | HW Problem | 26-6 KMC fault
127254 | 42272.1165 |  | LI26 | 51 | 4 | HARDWARE | HW Problem | Klystron LI26-51 trips frequently. Nothing shows on errors display. Klystron resets remotely. Possible intermittent fault in MKSU-to-Modulator status cable.
127265 | 42272.4414 |  | LI25 | 11 | 4 | HARDWARE | HW Problem | KLYS:LI25:11:WNDW1_TEMP = 137.8 deg. F
127053 | 42267.6176 | 0.5 | LI28 | 51 | 4 | HARDWARE | HW Problem | 28-5: Touch Panel display flickering (still readable).
127077 | 42268.1435 | 2 | LI21 | 11 | 4 | HARDWARE | HW Problem | 21-1 has an intermittent SCR klixon fault
127077 | 42268.1435 | 5 | LI21 | 11 | 4 | HARDWARE | HW Problem | 21-1 has an intermittent SCR klixon fault
127078 | 42268.2943 | 6 | LI30 | 41 | 4 | HARDWARE | HW Problem | TBR - 30-4 is tripping on Klystron Filament O/C, times out and does it again. When it resets and filament current / voltage is ramping back up, it is erratic & bouncing around + filament voltage is low for the current.
127078 | 42268.2943 | 1 | LI30 | 41 | 4 | HARDWARE | HW Problem | TBR - 30-4 is tripping on Klystron Filament O/C, times out and does it again. When it resets and filament current / voltage is ramping back up, it is erratic & bouncing around + filament voltage is low for the current.
127192 | 42270.7429 | 2 | LI20 | 71 | 4 | HARDWARE | HW Problem | L0-A has tripped off several times today. It resets fine, but is an important station that we can't afford to have tripping consistently.
127255 | 42272.166 |  | LI22 | 51 | 4 | HARDWARE | HW Problem | Klystron LI22-51 is tripped on MKSU Trigger fault and will not reset remotely.
127255 | 42272.166 |  | LI22 | 51 | 4 | HARDWARE | HW Problem | Klystron LI22-51 is tripped on MKSU Trigger fault and will not reset remotely.
127527 | 42279.8369 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X feedback setpoint is maxing out without making up desired amplitude of 20 MV at 120 Hz after 2-bunch ms separation MD.
127671 | 42286.3209 | 1 | LI20 | 81 | 4 | HARDWARE | HW Problem | 20-8 Need to inspect , T20 input phase 2 input bushing 95% had a very slow oil leak, i tighten down nut while Critt and i were tunning modulator.
126890 | 42256.1459 | 1.5 | LI29 | 81 | 4 | HARDWARE | HW Problem | 29-8 has a Core Bias U/C fault, current is 9.3 amps
127320 | 42275.1566 | 1 | LI18 | 11 | 4 | HARDWARE | HW Problem | 18-1 has a thyratron reservoir U/C fault
127321 | 42275.1574 | 0.5 | LI15 | 51 | 4 | HARDWARE | HW Problem | 15-5 Modulator, Needs thyratron and T20 replaced. T20 Leaking oil from case crack. Oil covers the front cabinet floor, could be a possible fire hazard.
127325 | 42275.1618 | 1 | LI15 | 81 | 4 | HARDWARE | HW Problem | 15-8 has an ACC2 water fault
127326 | 42275.163 | 1 | LI12 | 81 | 4 | HARDWARE | HW Problem | 12-8 has a Klystron Water fault
127327 | 42275.1638 | 1 | LI14 | 21 | 4 | HARDWARE | HW Problem | 14-2 has a Klystron Water fault
127328 | 42275.1646 | 1 | LI16 | 41 | 4 | HARDWARE | HW Problem | 16-4 has a Klystron Water fault
127329 | 42275.1659 | 2 | LI17 |  | 4 | HARDWARE | HW Problem | 17-4, 17-7 & 17-8 have Klystron Water faults
127330 | 42275.1671 | 2 | LI18 |  | 4 | HARDWARE | HW Problem | 18-1, 18-2, 18-3, 18-5, 18-6, 18-7 & 18-8 have Klystron Water faults
127355 | 42275.5411 | 1 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4 electromagnet current fault.
126785 | 42242.5521 | 1 | LI02 | 81 | 4 | HARDWARE | HW Problem | 2-8, PIOP/MKSU, 24V Battery Fault on PIOP data. No work to start before Fiscal year 2016, Oct 1st, 2015
126932 | 42261.0737 |  | LI28 |  | 4 | HARDWARE | HW Problem | Crate LI28 CR03 has ac power but appears dead
127064 | 42268.0349 | 1 | LI21 | 11 | 4 | HARDWARE | HW Problem | The PAD -15 Volts value for L1S or 21-1 is higher than expected and outside the required tolerance. The expected is -15V while the actual is -13.9V.
127079 | 42268.2972 | 0.5 | LI30 | 31 | 4 | HARDWARE | HW Problem | 30-3 is tripping on external interlocks - saw klystron Delta Temp fault lamp flicker on MKSU several times.
127103 | 42268.9812 | 3 | LI20 | 71 | 4 | HARDWARE | HW Problem | 20-7 L0A false MOD fault status. Station is running.
127125 | 42269.5579 |  | LI30 | 41 | 4 | HARDWARE | HW Problem | 30--4: EOLC faults see CATER 127135.
127175 | 42270.5369 | 2 | LI21 | 11 | 4 | HARDWARE | HW Problem | 21-1 SCR oscillation at 340 - 350 kV klystron beam voltage (corresponding to about 21 kV charge).
127366 | 42276.0335 | 1 | LI09 | 81 | 4 | HARDWARE | HW Problem | 9-8 Faulty VGC needs replace. Loud relay chatter when Modulator is keyed to green light.
127370 | 42276.0973 |  | LI21 | 11 | 4 | HARDWARE | HW Problem | L1S phase and amplitude change in response to LCLS rate changes. Presumably indicates that standby timing is not set up properly. See this elog entry: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=803194
127376 | 42276.2816 | 6.5 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 trips when set to accelerate runs fine otherwise. See MCCELOG entry: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=803413 DO NOT CLOSE THIS CATER UNTIL WE HAVE A SOLUTION! SEE SOLUTIONS 2, 3 and 4 FOR HISTORY OF THIS PROBLEM.. From; CATER Id: 128850 Time accounting - excessive klystron trips causing downtime for XPP.. As a start, the most obvious culprits are 21-4, 21-5, 24-5, and 24-6. There are many others that are not too far from being just as bad. Note that the PIOP error count is under reporting, possibly because they are fast. Wouldn't be surprised if these are mostly stations that need a new Thyratron Also see cater:129178
127376 | 42276.2816 | 1 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 trips when set to accelerate runs fine otherwise. See MCCELOG entry: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=803413 DO NOT CLOSE THIS CATER UNTIL WE HAVE A SOLUTION! SEE SOLUTIONS 2, 3 and 4 FOR HISTORY OF THIS PROBLEM.. From; CATER Id: 128850 Time accounting - excessive klystron trips causing downtime for XPP.. As a start, the most obvious culprits are 21-4, 21-5, 24-5, and 24-6. There are many others that are not too far from being just as bad. Note that the PIOP error count is under reporting, possibly because they are fast. Wouldn't be surprised if these are mostly stations that need a new Thyratron Also see cater:129178
127376 | 42276.2816 | 1 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 trips when set to accelerate runs fine otherwise. See MCCELOG entry: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=803413 DO NOT CLOSE THIS CATER UNTIL WE HAVE A SOLUTION! SEE SOLUTIONS 2, 3 and 4 FOR HISTORY OF THIS PROBLEM.. From; CATER Id: 128850 Time accounting - excessive klystron trips causing downtime for XPP.. As a start, the most obvious culprits are 21-4, 21-5, 24-5, and 24-6. There are many others that are not too far from being just as bad. Note that the PIOP error count is under reporting, possibly because they are fast. Wouldn't be surprised if these are mostly stations that need a new Thyratron Also see cater:129178
127376 | 42276.2816 | 2 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 trips when set to accelerate runs fine otherwise. See MCCELOG entry: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=803413 DO NOT CLOSE THIS CATER UNTIL WE HAVE A SOLUTION! SEE SOLUTIONS 2, 3 and 4 FOR HISTORY OF THIS PROBLEM.. From; CATER Id: 128850 Time accounting - excessive klystron trips causing downtime for XPP.. As a start, the most obvious culprits are 21-4, 21-5, 24-5, and 24-6. There are many others that are not too far from being just as bad. Note that the PIOP error count is under reporting, possibly because they are fast. Wouldn't be surprised if these are mostly stations that need a new Thyratron Also see cater:129178
127376 | 42276.2816 |  | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 trips when set to accelerate runs fine otherwise. See MCCELOG entry: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=803413 DO NOT CLOSE THIS CATER UNTIL WE HAVE A SOLUTION! SEE SOLUTIONS 2, 3 and 4 FOR HISTORY OF THIS PROBLEM.. From; CATER Id: 128850 Time accounting - excessive klystron trips causing downtime for XPP.. As a start, the most obvious culprits are 21-4, 21-5, 24-5, and 24-6. There are many others that are not too far from being just as bad. Note that the PIOP error count is under reporting, possibly because they are fast. Wouldn't be surprised if these are mostly stations that need a new Thyratron Also see cater:129178
127376 | 42276.2816 |  | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 trips when set to accelerate runs fine otherwise. See MCCELOG entry: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=803413 DO NOT CLOSE THIS CATER UNTIL WE HAVE A SOLUTION! SEE SOLUTIONS 2, 3 and 4 FOR HISTORY OF THIS PROBLEM.. From; CATER Id: 128850 Time accounting - excessive klystron trips causing downtime for XPP.. As a start, the most obvious culprits are 21-4, 21-5, 24-5, and 24-6. There are many others that are not too far from being just as bad. Note that the PIOP error count is under reporting, possibly because they are fast. Wouldn't be surprised if these are mostly stations that need a new Thyratron Also see cater:129178
127376 | 42276.2816 |  | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 trips when set to accelerate runs fine otherwise. See MCCELOG entry: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=803413 DO NOT CLOSE THIS CATER UNTIL WE HAVE A SOLUTION! SEE SOLUTIONS 2, 3 and 4 FOR HISTORY OF THIS PROBLEM.. From; CATER Id: 128850 Time accounting - excessive klystron trips causing downtime for XPP.. As a start, the most obvious culprits are 21-4, 21-5, 24-5, and 24-6. There are many others that are not too far from being just as bad. Note that the PIOP error count is under reporting, possibly because they are fast. Wouldn't be surprised if these are mostly stations that need a new Thyratron Also see cater:129178
127376 | 42276.2816 |  | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 trips when set to accelerate runs fine otherwise. See MCCELOG entry: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=803413 DO NOT CLOSE THIS CATER UNTIL WE HAVE A SOLUTION! SEE SOLUTIONS 2, 3 and 4 FOR HISTORY OF THIS PROBLEM.. From; CATER Id: 128850 Time accounting - excessive klystron trips causing downtime for XPP.. As a start, the most obvious culprits are 21-4, 21-5, 24-5, and 24-6. There are many others that are not too far from being just as bad. Note that the PIOP error count is under reporting, possibly because they are fast. Wouldn't be surprised if these are mostly stations that need a new Thyratron Also see cater:129178
127376 | 42276.2816 |  | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 trips when set to accelerate runs fine otherwise. See MCCELOG entry: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=803413 DO NOT CLOSE THIS CATER UNTIL WE HAVE A SOLUTION! SEE SOLUTIONS 2, 3 and 4 FOR HISTORY OF THIS PROBLEM.. From; CATER Id: 128850 Time accounting - excessive klystron trips causing downtime for XPP.. As a start, the most obvious culprits are 21-4, 21-5, 24-5, and 24-6. There are many others that are not too far from being just as bad. Note that the PIOP error count is under reporting, possibly because they are fast. Wouldn't be surprised if these are mostly stations that need a new Thyratron Also see cater:129178
127376 | 42276.2816 | 1.5 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 trips when set to accelerate runs fine otherwise. See MCCELOG entry: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=803413 DO NOT CLOSE THIS CATER UNTIL WE HAVE A SOLUTION! SEE SOLUTIONS 2, 3 and 4 FOR HISTORY OF THIS PROBLEM.. From; CATER Id: 128850 Time accounting - excessive klystron trips causing downtime for XPP.. As a start, the most obvious culprits are 21-4, 21-5, 24-5, and 24-6. There are many others that are not too far from being just as bad. Note that the PIOP error count is under reporting, possibly because they are fast. Wouldn't be surprised if these are mostly stations that need a new Thyratron Also see cater:129178
127381 | 42276.3514 | 1 | LI23 | 61 | 4 | HARDWARE | HW Problem | Kly.s LI 23 - 6 Phase jitter = 1.7
127410 | 42277.2986 | 1 | LI06 | 11 | 4 | HARDWARE | HW Problem | 6-1 No Green light on. When HV off is pushed in the green light comes on.
127436 | 42277.8062 | 1 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3 Modulator DeQuing current is zero.
127449 | 42278.0373 | 1 | LI04 | 81 | 4 | HARDWARE | HW Problem | 4-8 Klys Vac P/S meter reads 10 to minus 8 Torr.
127450 | 42278.0388 | 1 | LI05 | 71 | 4 | HARDWARE | HW Problem | 5-7 Klys Vac P/S meter reads 5 DC micro amp.
127456 | 42278.1986 | 1 | LI26 | 61 | 4 | HARDWARE | HW Problem | 26-6 HVR and will not reset remotely.
127134 | 42269.6564 | 2 | LI24 | 11 | 4 | HARDWARE | HW Problem | 24-1: Waveguide 1 Water fault (intermittent).
127233 | 42271.5331 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | L1-X HVR no remote reset
127235 | 42271.6218 |  | LI29 | 31 | 4 | HARDWARE | HW Problem | 29-3 MKT fault. No remote reset.
127253 | 42272.1139 | 1 | LI22 | 71 | 4 | HARDWARE | HW Problem | Klystron LI22-71 is tripped off; shows MKSU trigger fault. PIOP display shows Electro-magnet current problem.
127294 | 42273.4702 | 1 | LI25 | 51 | 4 | HARDWARE | HW Problem | Klystron LI25-51 says it's tripped off on KMC, but PIOP data is a mess. Beam current shows 10x expected value, beam volts 0ver 7K, focus supply near zero, and many red status items that should read okay. Tried IPLing PIOP but no improvement.
127301 | 42274.0378 |  | LI24 | 51 | 4 | HARDWARE | HW Problem | Duplicate see 127337. Klystron LI24-51 shows high phase and amplitude jitter.
127339 | 42275.3215 | 0.5 | LI22 | 61 | 4 | HARDWARE | HW Problem | 22-6 MKT fault no reset.
127554 | 42282.0013 | 3 | LI20 | 81 | 4 | HARDWARE | HW Problem | 20-8 is leaking oil from the front, T20 area out of the cabinet and onto the floor.
127457 | 42278.2074 | 1 | LI03 | 61 | 4 | HARDWARE | HW Problem | 3-6 Will not Key to green light. Stuck in HVOC light.
127457 | 42278.2074 | 1 | LI03 | 61 | 4 | HARDWARE | HW Problem | 3-6 Will not Key to green light. Stuck in HVOC light.
127464 | 42278.3714 | 0.5 | LI23 | 41 | 4 | HARDWARE | HW Problem | Klys 23-4 is tripping repeatedly
127464 | 42278.3714 | 1 | LI23 | 41 | 4 | HARDWARE | HW Problem | Klys 23-4 is tripping repeatedly
127473 | 42278.6443 |  | LI20 | 93 | 4 | HARDWARE | HW Problem | 20-3C modulator failed PPS certification for chain-A permit. Seems to be indicating permit always on, by viewing modulator PPS lamps. Note: must use not permitted" lamps on modulator since PPS interface box LEDs don't work. Chain-A lamp never lights.
127489 | 42279.3329 |  | LI28 | 41 | 4 | HARDWARE | HW Problem | Klystron water switch fault @ 8.5 on rotor meter
127494 | 42279.4017 | 1 | LI11 | 81 | 4 | HARDWARE | HW Problem | 11-8 ACC2 Water fault
127529 | 42279.9788 | 1 | LI20 | 81 | 4 | HARDWARE | HW Problem | L0B phase and amplitude jitter are ~double what they were yesterday
127541 | 42281.1277 | 1 | LI21 | 41 | 4 | HARDWARE | HW Problem | 21-4 klystron magnet focus supply intermittently tripping on flow switch interlock, need flow & flow switch checked
127542 | 42281.2126 | 2 | LI25 | 41 | 4 | HARDWARE | HW Problem | Klystron 25-4 tripping off the beam constantly when set to accelerate
127545 | 42281.6133 | 1 | LI30 | 11 | 4 | HARDWARE | HW Problem | 30-1: PFN needs tuning. Slope is 1.0
127159 | 42270.4201 | 1 | LI22 | 11 | 4 | HARDWARE | HW Problem | 22-1: Intermittent Klystron Water fault.
127194 | 42270.9481 | 1 | LI22 | 81 | 4 | HARDWARE | HW Problem | 22-8 tripping on klystron magnet flow switch faults.
127234 | 42271.5439 | 0.3 | LI22 | 41 | 4 | HARDWARE | HW Problem | 22-4 MKSU Trigger Enable Fault, no remote reset.
127543 | 42281.2934 | 1 | LI23 | 61 | 4 | HARDWARE | HW Problem | 23-6 high phase jitter.
127585 | 42283.2555 | 1 | LI22 | 41 | 4 | HARDWARE | HW Problem | Klystron 22-4 has considerable beam volts and phase jitter
127587 | 42283.2624 | 1 | LI03 | 31 | 4 | HARDWARE | HW Problem | 3-3 Klystron Filament Voltgage and current meters reads approx 12A and 4V.
127587 | 42283.2624 |  | LI03 | 31 | 4 | HARDWARE | HW Problem | 3-3 Klystron Filament Voltgage and current meters reads approx 12A and 4V.
127599 | 42283.5735 | 1 | LI26 | 11 | 4 | HARDWARE | HW Problem | 26-1 has a KMC fault - no remote reset. KMC = electro-magnet current
127625 | 42284.6189 | 1 | LI28 | 71 | 4 | HARDWARE | HW Problem | 28-7 WG#2 FS fault
127665 | 42285.8813 |  | DR13 | 1 | 4 | HARDWARE | HW Problem | 1-8 PAD looks dead.
127667 | 42286.2185 | 1 | LI15 | 71 | 4 | HARDWARE | HW Problem | 15-7 PIOP is dead and will not IPL
127593 | 42283.4156 | 2 | LI20 | 93 | 4 | HARDWARE | HW Problem | LI20-3C water leak on flow restrictor - Inst tech can fix
127617 | 42284.4686 | 0.1 | LI26 | 81 | 4 | HARDWARE | HW Problem | 26-8 Remove a 22" section of cable from sub drive line to the IPA.
127669 | 42286.225 | 4 | LI18 | 81 | 4 | HARDWARE | HW Problem | 18-8 No green light. Modulator trips CB7 at HV on.
127691 | 42286.6145 |  | LI02 | 41 | 4 | HARDWARE | HW Problem | 2-4 No phase or AMPL reading from PAD but FEPK indicates 57MW.
127711 | 42286.785 | 1 | LI16 | 51 | 4 | HARDWARE | HW Problem | 16-5 waveguide vacuum fault.
128257 | 42307.0662 | 1 | LI13 | 11 | 4 | HARDWARE | HW Problem | Klystron 13-1 MOD faulted and will not reset
128266 | 42307.547 | 0.5 | LI23 | 41 | 4 | HARDWARE | HW Problem | Klys 23-4 AMPL is not regulating well.
128326 | 42311.3712 |  | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4 amplitude/phase jitter, clearly affecting beam.
128341 | 42311.5714 | 1 | LI15 | 81 | 4 | HARDWARE | HW Problem | 15-8 Mod T20 compartment cooling fan is dead
128361 | 42312.034 | 1.5 | LI25 | 81 | 4 | HARDWARE | HW Problem | Klystron LI25-81 is tripped and will not reset remotely.
128366 | 42312.3913 |  | LI28 | 41 | 4 | HARDWARE | HW Problem | Klystron 28-4 (and 22-2) have huge phase jitter (4 deg) and voltage jitter (4kV) at 20 Hz.
127417 | 42277.5451 |  | LI20 | 93 | 4 | HARDWARE | HW Problem | PPS inhibit lights on the Modulator 20-3C does not appear to be working. When we cycle the PPS state where the lights should be inhibited and light up, they do not. I am not sure if that requires modulator power to function, or if it is powered by 24V power.
127588 | 42283.2688 | 1 | LI11 | 81 | 4 | HARDWARE | HW Problem | 11-8 Water flow/fire alarm light is on. The MKSU shows no water faults. The Fire alarm box LED is green.
127700 | 42286.6337 |  | LI10 | 21 | 4 | HARDWARE | HW Problem | 10-2 and 10-3 show bad CAMAC error.
127701 | 42286.6355 |  | LI10 | 41 | 4 | HARDWARE | HW Problem | 10-4 No RF when drive is requested.
127980 | 42296.3959 | 0.5 | LI04 | 51 | 4 | HARDWARE | HW Problem | 4-7: MOD faulted, VVS breaker tripped.
127980 | 42296.3959 | 0.5 | LI04 | 51 | 4 | HARDWARE | HW Problem | 4-7: MOD faulted, VVS breaker tripped.
128524 | 42317.8302 | 1 | LI05 | 71 | 4 | HARDWARE | HW Problem | 5-7 WGV fault plus maybe some other issues
128817 | 42328.7122 | 1 | LI28 | 31 | 4 | HARDWARE | HW Problem | Klys 28-3 tripping constantly HVR. Won't stay on. Cycles to AMPL with very low beamvolts when reset. Unable to get thyratron to fire. Duplicate CATER. See CATER 128844
127938 | 42292.7152 | 0.75 | LI18 | 81 | 4 | HARDWARE | HW Problem | Kly. LI18 - 8 tripped on "MOD" fault and won't reset.
127960 | 42293.9366 | 1 | LI28 | 51 | 4 | HARDWARE | HW Problem | Klystron 28-5 mod faulted and won't reset remotely
128026 | 42297.4868 | 3 | LI28 | 61 | 4 | HARDWARE | HW Problem | 28-6: Station in LOCAL, trips when put in REMOTE.
128026 | 42297.4868 | 0.5 | LI28 | 61 | 4 | HARDWARE | HW Problem | 28-6: Station in LOCAL, trips when put in REMOTE.
128168 | 42303.376 | 0.5 | LI25 | 51 | 4 | HARDWARE | HW Problem | 25-5 HVR. In local, can't reset remotely.
127038 | 42267.1437 | 1 | LI23 | 21 | 4 | HARDWARE | HW Problem | Klys 23-2 Pad -15 Volts value is not within tolerance. The current value reads -13.0 but expected in -15 with 1.0 as the tolerance.
127102 | 42268.9269 |  | LI20 | 51 | 4 | HARDWARE | HW Problem | TCAV0 RFS 20-5 has no drive power, no PAC out on diagnostic PAD monitor. Curiously, when we reboot the PAC, there is some RF. But after reboot is done....not PAC out RF.
127182 | 42270.5935 | 4 | LI20 | 61 | 4 | HARDWARE | HW Problem | 20-6 MOD fault, won't reset on 9/23 Owl Shift.
127422 | 42277.5867 | 1 | LI29 | 81 | 4 | HARDWARE | HW Problem | 29-8 Mod grid trigger chassis needs to be replaced because of bad componant solder joint.
128138 | 42300.5105 | 0.75 | LI30 | 51 | 4 | HARDWARE | HW Problem | 30-5: Seeing some beam volts jitter.
128275 | 42308.1872 | 1 | LI16 | 51 | 4 | HARDWARE | HW Problem | Klystron 16-4 MOD and will not reset
128275 | 42308.1872 | 1 | LI16 | 51 | 4 | HARDWARE | HW Problem | Klystron 16-4 MOD and will not reset
128305 | 42310.4048 | 1 | LI11 | 21 | 4 | HARDWARE | HW Problem | 11-2 Modulator, Check Phase balance.
128364 | 42312.1616 | 1 | LI20 | 71 | 4 | HARDWARE | HW Problem | 20-7 (L0A) BVLT drop outs.
128405 | 42313.3383 |  | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 trips frequently.
128413 | 42313.4082 |  | LI02 | 61 | 4 | HARDWARE | HW Problem | Kly. LI02 - 6 Amplit. came back low after the last POMM. It is showing extreme AMPL jitter as well.
128432 | 42313.7613 | 1 | LI02 | 31 | 4 | HARDWARE | HW Problem | Klys 2-3 has a WGV fault
128460 | 42314.547 | 1 | LI25 | 51 | 4 | HARDWARE | HW Problem | 25-5 has gone to HVR multiple times.
128462 | 42314.567 | 1 | LI23 | 21 | 4 | HARDWARE | HW Problem | 23-2 has high phase jitter.
126889 | 42256.1444 | 1 | LI29 | 11 | 4 | HARDWARE | HW Problem | 29-1 is faulting on Thyratron filament O/C, has high current and very low voltage when trying to bring up NOTE: thyratron was replaced, waiting VVS to complete run-up
126889 | 42256.1444 | 1.5 | LI29 | 11 | 4 | HARDWARE | HW Problem | 29-1 is faulting on Thyratron filament O/C, has high current and very low voltage when trying to bring up NOTE: thyratron was replaced, waiting VVS to complete run-up
127081 | 42268.3246 |  | LI22 | 61 | 4 | HARDWARE | HW Problem | 22-6 has high Klystron Vacuum, high voltage on at 100kv for now
127247 | 42272.0937 | 1 | LI29 | 71 | 4 | HARDWARE | HW Problem | Klystron LI29-71 shows 15 water faults on klystron errors panel.
127968 | 42295.1558 | 1 | LI25 | 61 | 4 | HARDWARE | HW Problem | 25-6 MOD fault, no remote reset
128154 | 42301.563 | 1 | LI23 | 11 | 4 | HARDWARE | HW Problem | Klys 23-1 showing AMPL/PHAS jitter.
128342 | 42311.5736 | 1 | LI16 | 61 | 4 | HARDWARE | HW Problem | 16-6 Mod Thyratron blower assem has a badly corroded bolt that needs to be replaced asap. Mod grid trigger chassis has a bunt wire and needs to be replaced.
128352 | 42311.5915 | 2 | LI23 | 21 | 4 | HARDWARE | HW Problem | 23-2 Forward Energy peak detector reading unphysically high (183.37) on PIOP
128390 | 42312.7819 | 1 | LI20 | 71 | 4 | HARDWARE | HW Problem | L0A has phase and amplitude dropouts
128393 | 42312.867 |  | LI29 | 21 | 4 | HARDWARE | HW Problem | Klys 29-2 tripping on water.
128205 | 42305.0487 | 2 | LI01 | 51 | 4 | HARDWARE | HW Problem | 1-5 thyratron compartment fan is bad, bearing is screeching
128222 | 42305.5962 | 1 | LI02 | 51 | 4 | HARDWARE | HW Problem | 2-5 Thyratron cooling blower is dead
128238 | 42306.2147 | 1 | LI23 | 21 | 4 | HARDWARE | HW Problem | 23-2 MOD fault, needs local reset.
128259 | 42307.3288 |  | LI22 | 31 | 4 | HARDWARE | HW Problem | 22-3 phase jitter
126900 | 42257.2142 | 1 | LI26 | 81 | 4 | HARDWARE | HW Problem | 26-8 Klystron Water fault light intermittently flashes
127115 | 42269.3284 | 1 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3 has an intermittent Klystron Water glitch, needs calibrated and/or adjusted
128482 | 42315.6768 |  | LI02 | 31 | 4 | HARDWARE | HW Problem | dupe see 128308, Klys 2-3 is having AMPL jumps.
128485 | 42315.9181 | 1 | LI14 | 81 | 4 | HARDWARE | HW Problem | Klys 14-8 dropped to KHD without any indication as to why.
128488 | 42316.5185 | 0.5 | LI10 | 41 | 4 | HARDWARE | HW Problem | station 10-4 mod faulted
127207 | 42271.3265 | 0 | LI27 | 81 | 4 | HARDWARE | HW Problem | Klystron LI27-81 is tripped off on Mod fault and will not reset remotely.
127211 | 42271.3628 |  | LI23 | 71 | 4 | HARDWARE | HW Problem | KLYS:LI23:71:WNDW2_TEMP read back reports 137.7 DEG.s
127260 | 42272.4109 | 1 | LI20 | 61 | 4 | HARDWARE | HW Problem | 20-6: Snubber boards need to be recheck after we fixed the miswires pin 7 & 8 of the molex connector P2 that plug to Enerpro board.
127367 | 42276.0355 | 1 | LI10 | 31 | 4 | HARDWARE | HW Problem | 10-3 When Keyed to green light light, the green light cycles on and off.
127586 | 42283.2584 | 2 | LI01 | 31 | 4 | HARDWARE | HW Problem | 1-3 KMC. Klystron Bulk Fiat CB trips. With the Trim p/s off and its load disconnected, CB trips. There is a Klixon red LED on the cb box.
128051 | 42298.0774 | 1.5 | LI25 | 71 | 4 | HARDWARE | HW Problem | Kly LI25-7 has a purple "NAN" on the Phase reading. Tried to IPL PIOP and Reset PIOP but no change was noticed.
128274 | 42308.069 | 1 | LI22 | 41 | 4 | HARDWARE | HW Problem | 22-4 BV jitter.
128276 | 42308.2251 | 1 | LI02 | 31 | 4 | HARDWARE | HW Problem | 2-3 is MOD faulted.
128276 | 42308.2251 | 0.1 | LI02 | 31 | 4 | HARDWARE | HW Problem | 2-3 is MOD faulted.
128284 | 42308.9631 | 1 | LI29 | 61 | 4 | HARDWARE | HW Problem | Klys 29-6 has lots of phase dropouts of around -2.5 degrees
128284 | 42308.9631 | 1 | LI29 | 61 | 4 | HARDWARE | HW Problem | Klys 29-6 has lots of phase dropouts of around -2.5 degrees
128288 | 42309.2785 | 1 | LI14 | 21 | 4 | HARDWARE | HW Problem | 14-2 has a lot of PFN jitter
128291 | 42309.4581 | 0.5 | LI22 | 11 | 4 | HARDWARE | HW Problem | duplicate see 127965.Klystron 22-1 is cycling frequently. It usually comes back on its own, but it trashes the FEL whenever it cycles. Currently on the beam.
128306 | 42310.4157 | 0 | LI22 | 71 | 4 | HARDWARE | HW Problem | 22-7 phase jitter is pretty bad. See it very clearly on energy BPMs in LCLS and on the FEL jitter.
128308 | 42310.4497 |  | LI02 | 31 | 4 | HARDWARE | HW Problem | 2-3 has bouts of no Phase and Amplitude readback.
128327 | 42311.3775 | 0.5 | LI09 | 61 | 4 | HARDWARE | HW Problem | LI09-6 KLY water faults
127080 | 42268.2989 | 2 | LI27 | 81 | 4 | HARDWARE | HW Problem | 27-8 Modulator, Tripping on HVOC.
127439 | 42277.8099 | 1 | LI30 | 11 | 4 | HARDWARE | HW Problem | 30-1 Modulator has DeQuing undercurrent fault.
127865 | 42290.7667 | 2 | LI13 | 81 | 4 | HARDWARE | HW Problem | Kly. LI13 - 8 faulted on MOD fault. VVS and CB7 tripped
127981 | 42296.3966 | 6 | LI10 | 71 | 4 | HARDWARE | HW Problem | station 10-7 KMC fault
128489 | 42316.5975 | 0.5 | LI09 | 11 | 4 | HARDWARE | HW Problem | station 9-1 mod faulted
128497 | 42317.0699 |  | LI02 | 31 | 4 | HARDWARE | HW Problem | duplicate. Klystron LI02-31 shows "phase mean"; does not show any PAD value. Also has amplitude problems. Amplitude jitter has increased dramatically over the weekend and amplitude FTPs show dropouts.
128499 | 42317.1072 | 0.5 | LI02 | 31 | 4 | HARDWARE | HW Problem | Klystron LI02-31 has a bad vacuum gauge controller.
128525 | 42317.9234 | 0.5 | LI23 | 41 | 4 | HARDWARE | HW Problem | 23-4 Tripping (HVR faulting) repeatedly.
128537 | 42318.3855 | 1 | LI26 | 41 | 4 | HARDWARE | HW Problem | 26-4 DQ curr low = 5A. Needs a calibration for linear regulation.
128571 | 42320.0552 | 1 | LI13 | 51 | 4 | HARDWARE | HW Problem | 13-5 has a false VVS status
128573 | 42320.1713 | 0.5 | LI23 | 21 | 4 | HARDWARE | HW Problem | 23-2 tripped HVR no remote reset as station in is local
127258 | 42272.3871 | 1 | LI24 | 41 | 4 | HARDWARE | HW Problem | Seeing Water WG2 fault and klystron water .
127317 | 42275.1525 | 0.5 | LI14 | 11 | 4 | HARDWARE | HW Problem | 14-1 Fire Alarm Latching Box "System Ready" green LED is out.
127352 | 42275.5047 |  | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2 faulty phase/amplitude readback. Doesn't come back with PIOP IPL. **edit tallmike** - PIOP swaps didn't help this, but PEM came out and made it work somehow.
127409 | 42277.1761 | 1 | LI06 | 11 | 4 | HARDWARE | HW Problem | 6-6 Needs VGC replaced. Fault light will not clear.
127437 | 42277.8071 | 1 | LI23 | 81 | 4 | HARDWARE | HW Problem | 23-8 Modulator DeQuing current is zero.
127578 | 42282.9632 | 0 | LI28 | 11 | 4 | HARDWARE | HW Problem | 28-1 frequent HVR faults.
127737 | 42287.4899 |  | LI17 | 51 | 4 | HARDWARE | HW Problem | 17-5 TBR Trips on over voltage. Found core bias at 22.5 amps, klystron filament at 0 volts and 23 amps. Klystron tank connection hot and will not come off. _____With test load, all readings normal.____
127903 | 42291.6668 |  | LI17 | 61 | 4 | HARDWARE | HW Problem | station 17-6 has KMC fault
128340 | 42311.57 | 1 | LI15 | 61 | 4 | HARDWARE | HW Problem | 15-6 Mod PFN compartment fan is dead
128343 | 42311.5752 | 1 | LI16 | 71 | 4 | HARDWARE | HW Problem | 16-7 Mod PFN compartment fan is dead
128344 | 42311.5763 | 1 | LI17 | 11 | 4 | HARDWARE | HW Problem | 17-1 Mod PFN compartment fan is dead
128347 | 42311.5798 | 1 | LI17 | 61 | 4 | HARDWARE | HW Problem | 17-6 EOLC Stack burnt, Also T20 Network resistor is burnt
128347 | 42311.5798 | 1 | LI17 | 61 | 4 | HARDWARE | HW Problem | 17-6 EOLC Stack burnt, Also T20 Network resistor is burnt
128350 | 42311.5877 | 1 | LI18 | 41 | 4 | HARDWARE | HW Problem | 18-4 Mod PFN compartment fan is dead
128351 | 42311.5888 | 1.5 | LI19 | 11 | 4 | HARDWARE | HW Problem | 19-1 Mod PFN compartment fan is dead
128369 | 42312.4377 | 0.2 | LI20 | 81 | 4 | HARDWARE | HW Problem | 20-8 (L0B) WGV fault
128387 | 42312.7232 |  | LI20 | 71 | 4 | HARDWARE | HW Problem | 20-7 (L0A): Still has BV dropouts. PFN tuning may be needed.
128402 | 42313.1873 |  | LI16 | 81 | 4 | HARDWARE | HW Problem | 16-8 has a ACC1 water fault that won't clear
128406 | 42313.339 |  | LI30 | 21 | 4 | HARDWARE | HW Problem | Ll30-2 Klystron water fault
128408 | 42313.3521 |  | LI18 | 51 | 4 | HARDWARE | HW Problem | NTF see 128421. Kly. LI18-5 Amplit. still low. Was 59.15 before shut down. Now at 50.8. Beam volts in low.
127481 | 42278.9716 | 0.5 | LI23 | 11 | 4 | HARDWARE | HW Problem | 23-1 IPA doesn't seem to track well during phase scans.
128279 | 42308.6254 | 1 | LI29 | 21 | 4 | HARDWARE | HW Problem | 29-2 is tripping excessively. The station does reset remotely but the frequency is disruptive.
128283 | 42308.8762 | 0.5 | LI29 | 11 | 4 | HARDWARE | HW Problem | station 29-1 mod faulting repeatedly--has 77 AMPL faultls since this 08:00morning
128325 | 42311.3582 | 0.5 | LI09 | 61 | 4 | HARDWARE | HW Problem | Kly. LI09 - 6 Seeing a "water klystron" fault coming and going.
128339 | 42311.5687 | 1 | LI13 | 61 | 4 | HARDWARE | HW Problem | 13-6 Mod PFN compartment fan is dead
128384 | 42312.7183 | 1 | LI30 | 31 | 4 | HARDWARE | HW Problem | Klys 30-3 has CBL fault.
128590 | 42320.5064 | 1 | LI26 | 61 | 4 | HARDWARE | HW Problem | 26-6 improve PFN regulation
128591 | 42320.5067 | 1 | LI26 | 71 | 4 | HARDWARE | HW Problem | 26-7 improve PFN regulation
128592 | 42320.5071 | 1 | LI26 | 81 | 4 | HARDWARE | HW Problem | 26-8 improve PFN regulation
128595 | 42320.5593 |  | LI11 | 61 | 4 | HARDWARE | HW Problem | 11-6: HV CBT.
127689 | 42286.6105 |  | DR13 | 1 | 4 | HARDWARE | HW Problem | 1-8 new PAD installed, needs poly.
128009 | 42297.3029 | 1 | LI05 | 81 | 4 | HARDWARE | HW Problem | 5-8 Has a water leak
128163 | 42303.2604 | 1 | LI12 | 41 | 4 | HARDWARE | HW Problem | 12-4 WGV flt
128362 | 42312.0941 |  | LI23 | 71 | 4 | HARDWARE | HW Problem | Klystron LI23-71 is tripping frequently. No faults indicated on Errors panel.
128420 | 42313.5443 |  | LI11 | 61 | 4 | HARDWARE | HW Problem | Kly. LI 11 - 6 MKSU lost MKSU voltages. see pic. in PCD Log. Swapped MKSU pcd03271 out, pcd03094 in. symptoms did not change. Swapped piop cables with 11-5 and problem followed. Will transfer cater to hdwr for Piop swap. NIRP test ok.
128461 | 42314.5535 | 1 | LI16 | 51 | 4 | HARDWARE | HW Problem | 16-5 mod fault.
128503 | 42317.3933 | 1 | LI23 | 41 | 4 | HARDWARE | HW Problem | 23-4 needs calibration
128348 | 42311.5816 | 1 | LI17 | 71 | 4 | HARDWARE | HW Problem | 17-7 Thyratron compartment cooling fan is dead
128439 | 42314.0916 | 0.1 | LI12 | 81 | 4 | HARDWARE | HW Problem | 12-8 does not respond normally to phase scans. Unclear whether or not the phase shifter is responding appropriately to trim requests. Phase scans look like this: http://physics-elog.slac.stanford.edu/facetelog/show.jsp?dir=/2015/45/06.11&pos=2015-11-06T02:10:14
127290 | 42273.3274 | 2 | LI09 | 41 | 4 | HARDWARE | HW Problem | 9-4 Stepper relay K13 is not activating when Modulator is keyed to green.
127290 | 42273.3274 | 1 | LI09 | 41 | 4 | HARDWARE | HW Problem | 9-4 Stepper relay K13 is not activating when Modulator is keyed to green.
127319 | 42275.1553 | 0.5 | LI18 | 41 | 4 | HARDWARE | HW Problem | 18-4 has a thyratron filament U/C fault
127536 | 42280.9478 | 0.2 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3 dequing unstable, HV not regulating. Adjusted HV setpoint down from 21.85 to 21-65 with no change to HV readback.
127699 | 42286.6316 | 1 | LI09 | 51 | 4 | HARDWARE | HW Problem | 9-5 is faulted on klystron water.
128087 | 42298.9271 | 1 | LI13 | 31 | 4 | HARDWARE | HW Problem | 13-3 HV meter not indicating
128367 | 42312.3966 | 0.75 | LI22 | 21 | 4 | HARDWARE | HW Problem | klystron 22-2 (and 28-4) has a big phase jitter (2.2 deg rms) and voltage jitter too, mainly spikes down in voltage of 3 kV and corresponding phase spikes up to 10 deg.
128627 | 42321.0804 | 1 | LI03 | 31 | 4 | HARDWARE | HW Problem | 3-3 CBL fault.
128628 | 42321.0809 |  | LI13 | 21 | 4 | HARDWARE | HW Problem | 13-2 MOD fault.
128631 | 42321.2647 | 0.1 | LI21 | 11 | 4 | HARDWARE | HW Problem | L1S 21-1 tripped FLK PEM dispached
128632 | 42321.3458 | 1 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23 - 3 amplit. jitter is 4.45 %
128645 | 42321.5688 | 0.5 | LI24 | 81 | 4 | HARDWARE | HW Problem | Kly Mod 24-8: very low DQing current. Please adjust DQ current up to 20 amps or so.
128654 | 42321.7587 | 0.75 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3 repeatedly MOD faulting.
128655 | 42321.8653 | 2 | LI30 | 21 | 4 | HARDWARE | HW Problem | 30-2 KHD
128657 | 42321.9292 | 1 | LI27 | 41 | 4 | HARDWARE | HW Problem | 27-4 klystron filament supply chassis removed for installation in 30-2. Needs functional replacement.
128659 | 42321.9542 | 0.5 | LI11 | 41 | 4 | HARDWARE | HW Problem | 11-4 needs klystron filament time delay (K2) reset to 55 minutes
128662 | 42322.1201 | 1 | LI23 | 81 | 4 | HARDWARE | HW Problem | 23-8 FLK, won't reset.
128665 | 42322.8065 | 1 | LI02 | 31 | 4 | HARDWARE | HW Problem | 2-3 has erratic phase readback starting at 04:00 this morning (Nov 14), no readback at all since 13:00. PAD Mixer Gain reads 0.00, PAD +15 Volts reads 5.16. PAD Phase Mode on PIOP alternates between TRACK and LOST.
127089 | 42268.5677 | 1 | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2 Modulator, Trips off upon turn-on, with no latched fault. See HVOC on fault log.
127252 | 42272.1037 |  | LI22 | 51 | 4 | HARDWARE | HW Problem | Klystron LI22-51 shows 68 water faults on klystron errors panel. This klystron is currently tripped off and PEM needs water flow adjusted to bring it back on. Urgency changed from low priority" to "immediate"."
127289 | 42273.324 | 1 | LI03 | 61 | 4 | HARDWARE | HW Problem | 3-6 Trim p/s needs replace. Regulation is out of tols
127371 | 42276.1682 | 0.75 | LI06 | 31 | 4 | HARDWARE | HW Problem | 6-3 Fire Alarm Latching Box "System Ready" green LED is out.
127423 | 42277.5882 | 1 | LI28 | 51 | 4 | HARDWARE | HW Problem | 28-5 T20 input snubber wires are burnt
127461 | 42278.3036 | 1 | LI17 | 21 | 4 | HARDWARE | HW Problem | 17-2 Trim P/S needs replace. KMC on the MKSU
127937 | 42292.6478 |  | LI17 | 61 | 4 | HARDWARE | HW Problem | 17-6 is faulted on a KMC, Klystron magnet circuit, error. Klystron magnet klixon tell-tale indicates a klixon fault.
127937 | 42292.6478 | 0.8 | LI17 | 61 | 4 | HARDWARE | HW Problem | 17-6 is faulted on a KMC, Klystron magnet circuit, error. Klystron magnet klixon tell-tale indicates a klixon fault.
128448 | 42314.4307 | 4 | LI22 | 61 | 4 | HARDWARE | HW Problem | Klystron 22-6 is MOD and can't be reset remotely UPDATE: 22-6 apparently wasn't reading back phase measurements. We were having a huge problem with keeping the beam on at > 30 Hz (looked perfect at 30 Hz), which we tried to fix by resyncing PAUs, zeroing offsets on the longitudinal and LTU feedbacks, etc. When 22-6 tripped off, the problem was fixed and we could run at 120. Maxwell had been looking at klystron phase jitter, but 22-6 hadn't been reading back phase, because he thought it was a measurement problem. => maybe 22-6 was a problem?
128470 | 42314.6936 | 0.75 | LI02 | 51 | 4 | HARDWARE | HW Problem | Klystron 2-5 mod faulted and won't reset remotely.
128528 | 42318.2577 | 0.5 | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3 in HVR, no remote reset.
128575 | 42320.1822 | 1 | LI06 | 11 | 4 | HARDWARE | HW Problem | 6-1 PIOP display shows Battery Fault
128598 | 42320.5623 | 1 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3: "SCR Fan Fault".
128451 | 42314.4398 | 1 | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 needs calibration for BV, DQ.
128474 | 42314.8558 | 1 | LI20 | 71 | 4 | HARDWARE | HW Problem | L0A is tripping off repeatedly, it seems to turn back on, then drop out on amplitude mean.
128481 | 42315.4825 | 1 | LI12 | 11 | 4 | HARDWARE | HW Problem | 12-1 WGV fault.
128484 | 42315.7631 | 1 | LI03 | 21 | 4 | HARDWARE | HW Problem | 3-2 HV meter not reading
128486 | 42315.9399 | 0 | LI29 | 81 | 4 | HARDWARE | HW Problem | Klys 29-8 has beam volts jitter.
128493 | 42316.8551 | 0.2 | LI28 | 21 | 4 | HARDWARE | HW Problem | Klys 28-2 cycling constantly
128493 | 42316.8551 | 1 | LI28 | 21 | 4 | HARDWARE | HW Problem | Klys 28-2 cycling constantly
128493 | 42316.8551 | 1.5 | LI28 | 21 | 4 | HARDWARE | HW Problem | Klys 28-2 cycling constantly
128510 | 42317.4553 | 0.5 | LI10 | 21 | 4 | HARDWARE | HW Problem | 10-2 has mod faulted and will not reset remotely.
127045 | 42267.2278 |  | LI30 | 61 | 4 | HARDWARE | HW Problem | 30-6 is faulted on Klystron Water
127095 | 42268.6644 |  | LI26 | 61 | 4 | HARDWARE | HW Problem | 26-6: Trim supply dead (KMC fault).
127331 | 42275.1708 | 1 | LI15 | 71 | 4 | HARDWARE | HW Problem | 15-7 MKSU has "bad status" and won't IPL
127642 | 42285.1132 | 1 | LI02 | 71 | 4 | HARDWARE | HW Problem | 2-7 Has a water leak between Modulator and Klystron.
127880 | 42291.0016 |  | LI02 | 31 | 4 | HARDWARE | HW Problem | duplicate see 127690. Kly. LI02 - 3 PAD "phase' and Amplit. read back is bad.
127930 | 42292.4971 | 1 | LI20 | 41 | 4 | HARDWARE | HW Problem | FACET XTCAV no output. PAC I&Q out good. BV good. No Drive. No FWD power. TWT off?
128088 | 42298.9597 | 1 | LI10 | 41 | 4 | HARDWARE | HW Problem | 10-4 beam current readback has dropped to almost half of normal.
128124 | 42299.6917 | 1 | LI26 | 71 | 4 | HARDWARE | HW Problem | 26-7: Stations trip off when put in Remote.
128417 | 42313.4745 |  | LI12 | 41 | 4 | HARDWARE | HW Problem | Kly. LI 12 - 4 PAD DAC status shows "FALSE".
128666 | 42322.8406 | 1 | LI25 | 51 | 4 | HARDWARE | HW Problem | Klys 25-5 is tripping repeatedly (six times in the past hour)
128667 | 42322.8423 | 1 | LI29 | 61 | 4 | HARDWARE | HW Problem | Klys 29-6 trips frequently (~2x per hour)
128674 | 42323.6129 | 0.5 | LI21 | 31 | 4 | HARDWARE | HW Problem | 21-3: Replace lower PFN strap #4.
128677 | 42323.6441 | 0.5 | LI12 | 41 | 4 | HARDWARE | HW Problem | 12-4 WGV
128696 | 42324.5607 | 0.2 | LI06 | 61 | 4 | HARDWARE | HW Problem | 6-6: "Bad Timing Results", need's it's PFN tuned.
128697 | 42324.5701 | 1.5 | LI23 | 81 | 4 | HARDWARE | HW Problem | Modulators: 23-8, 28-4, and 29-4 have BV OOT. 11/19/15 28-4, and 29-4 were done with BV adjustment.
128700 | 42324.6122 | 1 | LI27 | 21 | 4 | HARDWARE | HW Problem | 27-2 BV dropouts
127239 | 42271.7261 | 1 | LI29 | 11 | 4 | HARDWARE | HW Problem | 29-1 WGV fault
127284 | 42272.9332 | 0.75 | LI23 | 81 | 4 | HARDWARE | HW Problem | 23-8 MOD fault, won't reset
127315 | 42275.1508 | 0.5 | LI13 | 11 | 4 | HARDWARE | HW Problem | 13-1 Fire Alarm Latching Box "System Ready" green LED is out.
127335 | 42275.2926 |  | LI25 | 71 | 4 | HARDWARE | HW Problem | 25-7 - phase from PIOP is NAN
127337 | 42275.3039 | 2 | LI24 | 51 | 4 | HARDWARE | HW Problem | 24-5 high phase and amplitude jitter
127668 | 42286.2222 |  | LI16 | 81 | 4 | HARDWARE | HW Problem | 16-8 Modulator water fault light is on. MKSU summary flt light is on.
127713 | 42286.7885 | 1 | LI17 | 61 | 4 | HARDWARE | HW Problem | 17-6 has a klystron focus magnet fault.
127842 | 42290.3185 |  | LI25 | 51 | 4 | HARDWARE | HW Problem | 25-5 is MOD faulted.
127849 | 42290.4624 |  | LI09 | 51 | 4 | HARDWARE | HW Problem | 9-5 has a persistent Klystron Water fault
127850 | 42290.4901 |  | LI01 | 51 | 4 | HARDWARE | HW Problem | Klystron 1-5 phase is not trimming correctly and does not trim to the expected value.
128600 | 42320.5756 |  | LI20 | 71 | 4 | HARDWARE | HW Problem | L0A (20-7) tripped off on MOD fault. Unable to reset remotely.
128656 | 42321.8939 | 2 | LI11 | 41 | 4 | HARDWARE | HW Problem | Klystron 11-4 MOD faulted and won't reset remotely
128424 | 42313.6139 | 0.5 | LI17 | 41 | 4 | HARDWARE | HW Problem | 17-4 beam volts is very unstable pulse-to-pulse. http://physics-elog.slac.stanford.edu/facetelog/show.jsp?dir=/2015/45/05.11&pos=2015-11-05T14:42:25 Observe on Nov. 06 11:45 RF.Amplit. dropped to 35 to 38.0. At same time beam started to oscillate. Removed 17-4 from beam
127755 | 42287.9411 | 0.5 | LI17 | 41 | 4 | HARDWARE | HW Problem | 17-4 Trips on over voltage and over current
127879 | 42290.9998 |  | LI05 | 81 | 4 | HARDWARE | HW Problem | Kly. LI05 - 8 PAD "phase' and Amplit. read back is now bad.
127901 | 42291.6435 |  | LI17 | 61 | 4 | HARDWARE | HW Problem | 17-6 KMC trip (see yesterday's CATER 127867)
128085 | 42298.9258 | 1 | LI11 | 41 | 4 | HARDWARE | HW Problem | 11-4 HV meter not indicating
128224 | 42305.5986 | 1 | LI06 | 41 | 4 | HARDWARE | HW Problem | 6-4 PFN compartment cooling fan is dead
128720 | 42325.44 | 2 | LI21 | 31 | 4 | HARDWARE | HW Problem | 21-3, Modulator. Install new Thyratron. Coordinate with ACR. This swap was arranged by Phil and Stanek for owl shift. Instabilites on 11/16. Could be due to sagging VVS voltage.
128731 | 42325.633 | 7.5 | LI24 | 81 | 4 | HARDWARE | HW Problem | Kly Mod 24-8: replace DQing divider. Divider no longer outputting signal to regulator board.
128737 | 42326.0808 | 2 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4 BV unstable.
128737 | 42326.0808 | 1 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4 BV unstable.
128737 | 42326.0808 | 1 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4 BV unstable.
128737 | 42326.0808 | 3 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4 BV unstable.
126938 | 42261.5462 | 1 | LI20 | 51 | 4 | HARDWARE | HW Problem | 20-5 (aka TCAV0): The valve is currently open. This CATER is for the TS/R of the WGV no Fault on the MKSU. The waveguide valve is closed but there is no WGV fault on the MKSU so triggers are available for the klystron to run.
127169 | 42270.4808 | 0 | LI25 | 61 | 4 | HARDWARE | HW Problem | 25-6: PIOP Status not correct. Reset/IPL no help.
127195 | 42271.0598 | 0.3 | LI23 | 11 | 4 | HARDWARE | HW Problem | 23-1 is MOD faulting frequently. I saw the water summary fault light several times on the MKSU but didn't see a specific flow switch interlock light
127196 | 42271.06 | 0.3 | LI23 | 21 | 4 | HARDWARE | HW Problem | 23-2 is MOD faulting frequently.
128346 | 42311.5783 | 1 | LI17 | 51 | 4 | HARDWARE | HW Problem | 17-5 Mod PFN compartment fan is dead
127462 | 42278.3093 | 1 | LI12 | 51 | 4 | HARDWARE | HW Problem | 12-5 VGC both flt lights are on.
127799 | 42289.1539 |  | LI13 | 11 | 4 | HARDWARE | HW Problem | 13-1 is MOD faulted - VVS breaker is tripped
128008 | 42297.197 | 3 | LI25 | 41 | 4 | HARDWARE | HW Problem | 25-4 Frequent KLT faults then self resets.
128008 | 42297.197 | 3 | LI25 | 41 | 4 | HARDWARE | HW Problem | 25-4 Frequent KLT faults then self resets.
128099 | 42299.2723 | 0.5 | LI29 | 31 | 4 | HARDWARE | HW Problem | 29-3 has high phase jitter.
128278 | 42308.5796 | 1 | LI27 | 81 | 4 | HARDWARE | HW Problem | As of 13:15, Klystron 27-8 has tripped off every 5 minutes. The station always resets remotely but is becoming exceptionally annoying.
128290 | 42309.4541 | 0.5 | LI23 | 21 | 4 | HARDWARE | HW Problem | KLYS 23-3 is HVR faulted and won't reset remotely. We took this station off the beam for now.
128383 | 42312.7167 | 1.5 | LI22 | 21 | 4 | HARDWARE | HW Problem | Klys 22-2 has really bad AMPL jitter.
128552 | 42318.876 | 0.75 | LI22 | 11 | 4 | HARDWARE | HW Problem | Klystron 22-1 has frequent beam volt drop outs that trip the beam.
128744 | 42326.3532 | 6 | LI21 | 71 | 4 | HARDWARE | HW Problem | 21-7 has mod faulted and will not reset remotely
128754 | 42326.5463 | 0.2 | LI28 | 11 | 4 | HARDWARE | HW Problem | 28-1 tripped to HVR no remote reset as station is in LOCAL.
128757 | 42326.6301 | 0.5 | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 trips when placed on the beam. TS&R. For evidence see: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=814952 2/15/16 Engineering taking over...
128757 | 42326.6301 | 2 | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 trips when placed on the beam. TS&R. For evidence see: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=814952 2/15/16 Engineering taking over...
128757 | 42326.6301 |  | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 trips when placed on the beam. TS&R. For evidence see: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=814952 2/15/16 Engineering taking over...
128757 | 42326.6301 | 6 | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 trips when placed on the beam. TS&R. For evidence see: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=814952 2/15/16 Engineering taking over...
128761 | 42327.2065 | 1 | LI16 | 51 | 4 | HARDWARE | HW Problem | 16-5 has a false MOD fault status
128764 | 42327.3781 | 0.5 | LI24 | 81 | 4 | HARDWARE | HW Problem | 24-8 DQing high
128764 | 42327.3781 | 0.2 | LI24 | 81 | 4 | HARDWARE | HW Problem | 24-8 DQing high
128767 | 42327.4018 | 2 | LI11 | 21 | 4 | HARDWARE | HW Problem | 11-2: Door #7 damaged, unable to latch securely.
128768 | 42327.4055 | 0.5 | LI11 | 81 | 4 | HARDWARE | HW Problem | 11-8: Ground hook holder on thyratron door missing.
128769 | 42327.4073 | 1 | LI12 | 21 | 4 | HARDWARE | HW Problem | 12-2: PFN strap #4 lower needs replacing.
128770 | 42327.409 | 1 | LI12 | 51 | 4 | HARDWARE | HW Problem | 12-5: Replace PFN strap #5 upper.
128771 | 42327.4111 | 1 | LI13 | 61 | 4 | HARDWARE | HW Problem | 13-6: Correct strap to cap connection on #2 lower cap.
128774 | 42327.4156 | 1 | LI14 | 21 | 4 | HARDWARE | HW Problem | 14-2: Replace PFN strap #5 upper.
128775 | 42327.4172 | 0.5 | LI15 | 41 | 4 | HARDWARE | HW Problem | 15-4: Replace PFN strap #2 upper.
128777 | 42327.4192 | 1 | LI15 | 51 | 4 | HARDWARE | HW Problem | 15-5: Replace upper PFN strap #6.
128778 | 42327.421 | 1 | LI15 | 61 | 4 | HARDWARE | HW Problem | 15-6: PFN strap #2 upper needs replacing.
128779 | 42327.4227 | 1 | LI15 | 81 | 4 | HARDWARE | HW Problem | 15-8: PFN strap #7 upper needs replacing.
127424 | 42277.5895 | 1 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 EOLC stack needs to be replaced
127608 | 42283.8931 | 1 | LI23 | 61 | 4 | HARDWARE | HW Problem | 23-6 trips off frequently, at least once per hour.
127751 | 42287.7466 |  | LI02 | 61 | 4 | HARDWARE | HW Problem | Duplicate see 125253. Klystron 2-6 has erroneous PAD temperature readback. It claims to be -77 C.
127756 | 42287.9414 | 3 | LI01 | 51 | 4 | HARDWARE | HW Problem | 1-5 BV intermittently becomes unstable. May have bad dequing circuit connection. The D'Qing chassis or connector reseating seems to have fixed the problem. /phill
127758 | 42287.9431 | 0.5 | LI17 | 21 | 4 | HARDWARE | HW Problem | 17-2 low beam voltage.
127798 | 42289.1526 | 0.5 | LI13 | 61 | 4 | HARDWARE | HW Problem | 13-6 is MOD
127374 | 42276.1715 | 1 | LI06 | 61 | 4 | HARDWARE | HW Problem | 6-6 has a Waveguide Vacuum Fault
127707 | 42286.7429 | 1 | LI28 | 11 | 4 | HARDWARE | HW Problem | 28-1 core bias unstable when mod is on
128249 | 42306.476 | 1 | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7 needs TS from local into "REMOTE"
128249 | 42306.476 | 1.5 | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7 needs TS from local into "REMOTE"
128249 | 42306.476 | 0.2 | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7 needs TS from local into "REMOTE"
128375 | 42312.5316 | 0.5 | LI30 | 71 | 4 | HARDWARE | HW Problem | 30-7 HVR; unable to reset remotely.
128400 | 42313.1644 | 0.5 | LI17 | 41 | 4 | HARDWARE | HW Problem | 17-4 output unstable when turned on
128750 | 42326.4159 | 1 | LI25 | 51 | 4 | HARDWARE | HW Problem | 25-5 HVR no remote reset since the station is in LOCAL
128780 | 42327.4295 | 1 | LI18 | 31 | 4 | HARDWARE | HW Problem | 18-3: Replace PFN straps and clips. See job for details.
128781 | 42327.4314 | 1 | LI18 | 51 | 4 | HARDWARE | HW Problem | 18-5: PFN straps and door latch need replacing. See job for details.
128782 | 42327.4331 | 1 | LI18 | 71 | 4 | HARDWARE | HW Problem | 18-7: Upper PFN strap #4 needs replacing.
128783 | 42327.4347 | 1 | LI19 | 41 | 4 | HARDWARE | HW Problem | 19-4: PFN strap #4 upper needs replacing.
128784 | 42327.4368 | 1 | LI20 | 93 | 4 | HARDWARE | HW Problem | 20-3C: PFN strap #2 upper needs replacing.
128788 | 42327.4873 |  | LI09 | 21 | 4 | HARDWARE | HW Problem | Klystron LI09-21 is tripped off and will not reset remotely.
128791 | 42327.5631 | 1 | LI26 | 61 | 4 | HARDWARE | HW Problem | 26-6 VVS trips
128800 | 42328.1208 | 4 | LI27 | 11 | 4 | HARDWARE | HW Problem | 27-1 MOD faulted due to thyratron aging.
128800 | 42328.1208 | 0.5 | LI27 | 11 | 4 | HARDWARE | HW Problem | 27-1 MOD faulted due to thyratron aging.
128800 | 42328.1208 | 2 | LI27 | 11 | 4 | HARDWARE | HW Problem | 27-1 MOD faulted due to thyratron aging.
127050 | 42267.6144 | 1 | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7: Klystron Water faults (intermittent) on MKSU.
127696 | 42286.6219 | 1 | LI05 | 71 | 4 | HARDWARE | HW Problem | 5-7 MKSU CBL fault.
127703 | 42286.6458 | 1 | LI11 | 41 | 4 | HARDWARE | HW Problem | 11-4 EMC fault
127795 | 42289.1462 | 2 | LI14 | 51 | 4 | HARDWARE | HW Problem | 14-5 is FLK
127797 | 42289.1498 | 2 | LI13 | 81 | 4 | HARDWARE | HW Problem | 13-8 is MOD faulted
127982 | 42296.3975 | 0.5 | LI12 | 41 | 4 | HARDWARE | HW Problem | station 12-4 WGV fault
128478 | 42315.2834 | 0.75 | LI15 |  | 4 | HARDWARE | HW Problem | 15-8 MOD fault
127020 | 42266.2848 | 1 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 Loose wires needs securing
127051 | 42267.6154 | 1 | LI24 | 41 | 4 | HARDWARE | HW Problem | 24-4: Klystron Water fault (solid) on MKSU.
127318 | 42275.1534 | 0.5 | LI15 | 71 | 4 | HARDWARE | HW Problem | 15-7 Fire Alarm Latching Box "System Ready" green LED is out.
127373 | 42276.1698 | 0.75 | LI02 | 71 | 4 | HARDWARE | HW Problem | 2-7 Fire Alarm Latching Box "System Ready" green LED is out.
127459 | 42278.2141 |  | LI15 | 71 | 4 | HARDWARE | HW Problem | 15-7 PIOP needs to be put back in the crate or replace.
127537 | 42280.9581 | 1 | LI25 | 11 | 4 | HARDWARE | HW Problem | Getting beam volts jitter from Klys 25-1
127572 | 42282.6819 | 0.5 | LI26 | 11 | 4 | HARDWARE | HW Problem | Klys 26-1 KMC fault.
127760 | 42287.9552 |  | LI09 | 81 | 4 | HARDWARE | HW Problem | 9-8 has no RF drive.
127762 | 42288.0877 | 1 | LI09 | 71 | 4 | HARDWARE | HW Problem | 9-7 has intermittent ACC1 water faults
127940 | 42292.7916 | 1 | LI10 | 21 | 4 | HARDWARE | HW Problem | Kly. LI10 - 2 tripping on "MOD" fault, resets but then trips again. Won't stay on.
128059 | 42298.2809 | 3 | LI25 | 41 | 4 | HARDWARE | HW Problem | Kly. LI 25 - 4 has started tripping on " DELTA KLY TEMP
128070 | 42298.4986 | 1 | LI12 | 51 | 4 | HARDWARE | HW Problem | Klys 12-5 has pretty bad AMPL jitter
128669 | 42323.2958 | 1 | LI17 | 61 | 4 | HARDWARE | HW Problem | 17-6 has high amplitude jitter.
128802 | 42328.3157 | 0.5 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3 MOD faulted. Phillip checked this station. Now, Phillip is still monitoring thru the day..
128805 | 42328.381 | 0.5 | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8: only trips when placed on the beam, CATER 128757
128810 | 42328.4659 | 0.5 | LI22 | 51 | 4 | HARDWARE | HW Problem | 22-5 MOD fault not reset.
128812 | 42328.5963 | 1 | LI21 | 71 | 4 | HARDWARE | HW Problem | 21-7 MOD fault no reset.
128813 | 42328.6433 | 1 | LI20 | 93 | 4 | HARDWARE | HW Problem | 20-3C failed PPS test. Turns on H V with key Coordinate with Mike Stanek for PPS key on CB7 may need rad safety work form
128813 | 42328.6433 | 1 | LI20 | 93 | 4 | HARDWARE | HW Problem | 20-3C failed PPS test. Turns on H V with key Coordinate with Mike Stanek for PPS key on CB7 may need rad safety work form
128816 | 42328.7074 |  | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2 PAD phase readback does not update/change when KPHR changes. 11/29 edit: disabled phase trim for this station.
126902 | 42257.2187 | 1 | LI22 | 51 | 4 | HARDWARE | HW Problem | 22-5 WGV flt
126995 | 42264.6359 | 3 | DR13 | 1 | 4 | HARDWARE | HW Problem | Water leak from KF01 8A 33 dripping into camac power supply
126995 | 42264.6359 | 1 | DR13 | 1 | 4 | HARDWARE | HW Problem | Water leak from KF01 8A 33 dripping into camac power supply
127438 | 42277.8088 | 1 | LI26 | 81 | 4 | HARDWARE | HW Problem | 26-8 Modulator has DeQuing undercurrent fault.
127453 | 42278.0764 |  | LI20 | 93 | 4 | HARDWARE | HW Problem | Unable to completely verify RF permit for 20-3C modulator. Chain B PPS inhibit light comes on, but Chain A does not.
128025 | 42297.4839 | 1 | LI29 | 31 | 4 | HARDWARE | HW Problem | 29-3: Flashing PFN fan fault.
128038 | 42297.5822 |  | LI25 | 11 | 4 | HARDWARE | HW Problem | The water flow at 25-1 is below tolerance at 12.8 GPM.
128429 | 42313.7095 | 0.5 | LI21 | 31 | 4 | HARDWARE | HW Problem | Klys 21-3 MKT fault.
128551 | 42318.7318 | 0.5 | LI29 | 61 | 4 | HARDWARE | HW Problem | 29-6 MOD fault, won't reset
128558 | 42319.4578 |  | LI21 | 31 | 4 | HARDWARE | HW Problem | Klys 21-3 MKT fault. Focus supply in EPICS indicates off but current and voltage read normal, faultless fault?
128566 | 42319.71 | 1 | LI22 | 81 | 4 | HARDWARE | HW Problem | 22-8: DQing current is low and unstable.
128567 | 42319.7724 | 2 | LI29 | 31 | 4 | HARDWARE | HW Problem | Klystron 29-3 mod faulted and won't reset remotely
128594 | 42320.5549 | 0.1 | LI29 | 11 | 4 | HARDWARE | HW Problem | Time accounting- vacuum burst in LI29 caused 0.2h LCLS downtime. Related to TBR work PEM is doing on 29-1.
128597 | 42320.5608 |  | LI12 | 21 | 4 | HARDWARE | HW Problem | 12-2: HV CBT.
128604 | 42320.6239 | 1 | LI20 | 71 | 4 | HARDWARE | HW Problem | L0A Beam volts dropping out a lot. Clear effect on LCLS performance.
127142 | 42269.934 | 1 | LI30 | 21 | 4 | HARDWARE | HW Problem | 30-2 tripping on ACOC
127492 | 42279.3843 | 1 | LI12 | 21 | 4 | HARDWARE | HW Problem | ACC water fault
127511 | 42279.6049 | 1 | LI20 | 81 | 4 | HARDWARE | HW Problem | Kly Mod 20-8, L0B: after SLED de-tuned, attempted to raise BV up to 350KV. @ 310 KV or so (21.5 KV HVPS), began to get many EOLC faults. TS/R EOLC faults: possible tuning issue or something wrong in mod with ckt or pulse cable issue. Assign as PAMM job.
128487 | 42316.1531 | 1 | LI09 | 31 | 4 | HARDWARE | HW Problem | 9-3 trips to HVR. I can reset it remotely and it comes on okay each time, but then trips off again with about 3 minutes.
128487 | 42316.1531 | 1 | LI09 | 31 | 4 | HARDWARE | HW Problem | 9-3 trips to HVR. I can reset it remotely and it comes on okay each time, but then trips off again with about 3 minutes.
128487 | 42316.1531 | 1 | LI09 | 31 | 4 | HARDWARE | HW Problem | 9-3 trips to HVR. I can reset it remotely and it comes on okay each time, but then trips off again with about 3 minutes.
128487 | 42316.1531 | 1 | LI09 | 31 | 4 | HARDWARE | HW Problem | 9-3 trips to HVR. I can reset it remotely and it comes on okay each time, but then trips off again with about 3 minutes.
128722 | 42325.4606 | 1.5 | LI23 | 11 | 4 | HARDWARE | HW Problem | 23-1 getting calibrations done. Cater per Phillip.
128722 | 42325.4606 | 0.2 | LI23 | 11 | 4 | HARDWARE | HW Problem | 23-1 getting calibrations done. Cater per Phillip.
128728 | 42325.5816 | 0.2 | LI28 | 11 | 4 | HARDWARE | HW Problem | 28-1 has tripped and had to be reset remotely. Also needs at DQing vs Beam volts Calibration.
128728 | 42325.5816 | 0.5 | LI28 | 11 | 4 | HARDWARE | HW Problem | 28-1 has tripped and had to be reset remotely. Also needs at DQing vs Beam volts Calibration.
128808 | 42328.4094 | 1 | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 trips every minute or so. Maybe only happens at accelerate. At one point I saw our ACR CUD flash with the "EMC" message.
128818 | 42328.8463 | 1 | LI11 | 61 | 4 | HARDWARE | HW Problem | 11-6 MOD fault would not reset
128819 | 42328.8843 | 1 | LI11 | 61 | 4 | HARDWARE | HW Problem | 11-6 T20 compartment cooling fan not working
128825 | 42329.0772 | 1 | LI21 | 41 | 4 | HARDWARE | HW Problem | Klystron 21-4 has tripped three times in the past hour and a half. Not sure if the job in CATER #128676 is meant to address this - if it is, then sorry for the duplicate!
128825 | 42329.0772 | 1 | LI21 | 41 | 4 | HARDWARE | HW Problem | Klystron 21-4 has tripped three times in the past hour and a half. Not sure if the job in CATER #128676 is meant to address this - if it is, then sorry for the duplicate!
128830 | 42329.7825 | 3 | LI29 | 71 | 4 | HARDWARE | HW Problem | 29-7 tripped approximately 5 time in about 3 mins.
128831 | 42329.7983 | 3 | LI22 | 61 | 4 | HARDWARE | HW Problem | 22-6 has mod faulted and Security got a supervisory alarm.
128838 | 42329.9871 | 3 | LI22 | 61 | 4 | HARDWARE | HW Problem | Klys 22-6 showing beam volts jitter.
128842 | 42330.1799 |  | LI25 | 21 | 4 | HARDWARE | HW Problem | Klys 25-2 has an intermittent CAM fault
128845 | 42330.4521 | 0.5 | LI17 | 21 | 4 | HARDWARE | HW Problem | Klystron 17-2 Fault Lockout won't reset.
128847 | 42330.6214 | 2 | LI27 | 21 | 4 | HARDWARE | HW Problem | Klys 27-2 showing beam volts jitter.
126947 | 42262.2527 | 1.5 | LI23 | 61 | 4 | HARDWARE | HW Problem | 23-6 core bias connection at A8TB2 pin E in the Klystron Junction Box is overheated
126964 | 42263.2867 | 1 | LI27 | 21 | 4 | HARDWARE | HW Problem | 27-2 Klystron Filament is faulted on O/C
126985 | 42264.2704 | 1.5 | LI22 | 61 | 4 | HARDWARE | HW Problem | 22-6 2ea. PFN cap needs replace. LHS 1st from top. Coil and straps side of caps. the other, RHS 2nd down from top. Coil and straps side of caps.
127055 | 42267.6195 | 1 | LI30 | 71 | 4 | HARDWARE | HW Problem | 30-7: WVGC has no filament activation indication.
127154 | 42270.385 | 0.5 | LI20 | 71 | 4 | HARDWARE | HW Problem | Kly. LI 20 - 7 shows DeQuing Under Current reads "Fault". DeQuing Amp. shows 0.0 AMP.s / DeQuing Volt.s = 7.12 volts.
127637 | 42284.9264 | 1 | LI20 | 93 | 4 | HARDWARE | HW Problem | 20-3C turns on at key on (automatically comes on when Ch-A PPS permit restored; does not wait to be told to turn on); HV off does not turn off mod. when PPS-chain-B permit lost.
127965 | 42294.677 | 1 | LI22 | 11 | 4 | HARDWARE | HW Problem | Klystron 22-1 has frequent reflected energy faults.
127970 | 42295.3466 |  | LI24 | 11 | 4 | HARDWARE | HW Problem | 24-1 No amplitude. No RF drive output.
127991 | 42296.513 | 0.25 | LI23 | 31 | 4 | HARDWARE | HW Problem | Klystron 23-3 has tripped on HVR and will not reset remotely
128021 | 42297.4526 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | This is a follow up to the cater "127895" and "127622". Three chassis in RFHUT were replaced from the troubleshooting of L1X. They are 2856MHz to 11.424GHz 4X Multiplier. X-Band PAD. LO Generator. It is a complex problem. All the removed chassis need to be inspected and repaired and retested. It's very possible that some new components and PCBs are needed during the repairing.
128047 | 42297.6899 | 0.75 | LI10 | 71 | 4 | HARDWARE | HW Problem | klystron 10-7 MOD faulted and won't reset remotely
126869 | 42255.2994 | 2 | LI22 | 41 | 4 | HARDWARE | HW Problem | 22-4 Modulator, EOLC fault won't reset
128338 | 42311.567 | 1 | LI12 | 81 | 4 | HARDWARE | HW Problem | 12-8 Mod Thyratron compartment fan is dead
128476 | 42315.1923 | 0.75 | LI25 | 51 | 4 | HARDWARE | HW Problem | 24-5 MOD fault, no reset (in Local)
128611 | 42320.7681 | 1 | LI16 | 51 | 4 | HARDWARE | HW Problem | 16-5 MOD fault, won't reset.
128625 | 42321.0565 | 1 | LI09 | 31 | 4 | HARDWARE | HW Problem | 9-3 MOD fault, doesn't reset.
128625 | 42321.0565 | 2 | LI09 | 31 | 4 | HARDWARE | HW Problem | 9-3 MOD fault, doesn't reset.
128630 | 42321.259 |  | LI16 | 51 | 4 | HARDWARE | HW Problem | 16-5 MOD fault.
128643 | 42321.5489 | 3 | LI29 | 31 | 4 | HARDWARE | HW Problem | 29-3: Touch Panel flashing "PFN fan" fault, but the actual is "T20 fan" fault. Needs to check and correct the wiring issue.
127490 | 42279.3508 | 1 | LI23 | 81 | 4 | HARDWARE | HW Problem | Klystron water switch fault
127741 | 42287.5604 | 1 | LI15 |  | 4 | HARDWARE | HW Problem | 15-1, 15-2 and 15-3 have low beam voltage.
128480 | 42315.3408 | 1 | LI15 | 41 | 4 | HARDWARE | HW Problem | 15-4 mod fault, no reset.
128853 | 42330.9801 | 2 | LI21 | 31 | 4 | HARDWARE | HW Problem | Klys 21-3 showing beam volts jitter. Not that bad, but it hurts us at low energy.
128854 | 42331.046 | 1 | LI25 | 51 | 4 | HARDWARE | HW Problem | Klys 25-5 keeps tripping Note: Phillip is still monitoring this station thru the day. PN
128854 | 42331.046 | 1 | LI25 | 51 | 4 | HARDWARE | HW Problem | Klys 25-5 keeps tripping Note: Phillip is still monitoring this station thru the day. PN
128857 | 42331.3089 | 1 | LI23 | 11 | 4 | HARDWARE | HW Problem | 23-1 FLK, showing multiple water faults in error log.
128874 | 42331.8063 |  | DR13 | 1 | 4 | HARDWARE | HW Problem | Kly. DR13 kly. is having water faults, looks like early this day
127090 | 42268.5774 | 1 | LI28 | 61 | 4 | HARDWARE | HW Problem | 28-6 Modulator, HV set low, about half power.
127287 | 42273.3162 | 1 | LI25 | 11 | 4 | HARDWARE | HW Problem | 25-1 Station drops to HVR with an occasional ACOC.
127612 | 42284.058 |  | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 cycles constantly.
128052 | 42298.0881 | 3 | LI10 | 51 | 4 | HARDWARE | HW Problem | Kly. LIN 10 - 4 is tripped "MOD" fault and won't reset.
128052 | 42298.0881 | 1 | LI10 | 51 | 4 | HARDWARE | HW Problem | Kly. LIN 10 - 4 is tripped "MOD" fault and won't reset.
128052 | 42298.0881 | 1 | LI10 | 51 | 4 | HARDWARE | HW Problem | Kly. LIN 10 - 4 is tripped "MOD" fault and won't reset.
128053 | 42298.1322 | 1 | LI25 | 11 | 4 | HARDWARE | HW Problem | Kly. LI25 - 1 tripped on "MOD" fault and won't reset.
128092 | 42299.0376 |  | LI12 | 61 | 4 | HARDWARE | HW Problem | PAD Ground = 5.0 PAD +15.0 Volts = +19.9 volts PAD - 15.0 Volts = +19.9 volts PAD +/_ 6 volt balance = +5.0volts PAS DAC Status = FALSE
128156 | 42301.6084 | 0.5 | LI25 | 51 | 4 | HARDWARE | HW Problem | Klys 25-5 cycling frequently on AMPL in error logger.
128661 | 42322.035 | 2 | LI30 | 51 | 4 | HARDWARE | HW Problem | 30-5 exhibiting BV dropouts.
128668 | 42323.2948 | 1 | LI18 | 51 | 4 | HARDWARE | HW Problem | 18-5 has high amplitude jitter.
128670 | 42323.297 | 1 | LI16 | 21 | 4 | HARDWARE | HW Problem | 16-2 high phase jitter.
128670 | 42323.297 | 1 | LI16 | 21 | 4 | HARDWARE | HW Problem | 16-2 high phase jitter.
128671 | 42323.2972 | 1 | LI16 | 71 | 4 | HARDWARE | HW Problem | 16-7 high phase jitter.
128672 | 42323.6096 | 0.5 | LI20 | 61 | 4 | HARDWARE | HW Problem | 20-6: Replace lower strap #8 on PFN.
128675 | 42323.6147 | 0.5 | LI21 | 51 | 4 | HARDWARE | HW Problem | 21-5: Replace lower PFN strap #1.
128675 | 42323.6147 | 0.5 | LI21 | 51 | 4 | HARDWARE | HW Problem | 21-5: Replace lower PFN strap #1.
128676 | 42323.6162 | 1.5 | LI21 | 41 | 4 | HARDWARE | HW Problem | 21-4: Replace 4 PFN straps. See job for details.
128680 | 42324.3588 |  | LI21 | 71 | 4 | HARDWARE | HW Problem | 21-7 klystron perveance low; below 1.65.
126917 | 42258.1484 | 2 | LI21 | 31 | 4 | HARDWARE | HW Problem | 21-3 MKSU has a Klystron Magnet flt that does not clear.
127295 | 42273.7008 | 1 | LI27 | 41 | 4 | HARDWARE | HW Problem | station 27-4 shows 45 water faults since 08:00 this morning (9 hours)
127697 | 42286.6253 |  | LI05 | 81 | 4 | HARDWARE | HW Problem | 5-8 no RF power to klystron when drive is applied.
127743 | 42287.6292 | 0.8 | LI25 | 51 | 4 | HARDWARE | HW Problem | Klystron 25-5 HVR faults frequently, which negatively affects FEL performance. Taking this station off of the beam for now. SEE CATER #128091
127767 | 42288.2497 | 1 | LI15 | 81 | 4 | HARDWARE | HW Problem | 15-8 has a WGV fault
127786 | 42288.5391 | 1 | LI16 |  | 4 | HARDWARE | HW Problem | PP-YY re-sync faults with odd trigger rates in sector 16. FIDO?
127868 | 42290.7749 |  | LI19 | 11 | 4 | HARDWARE | HW Problem | Kly. LI19 - 1 faulted on ELECTRO_MAGNET MOD fault.
127948 | 42293.3954 | 1 | LI23 | 31 | 4 | HARDWARE | HW Problem | Li23 modulator 23-3 has 30 Hz (1 out of 4) time slot jitter, and many single phase jitter spikes.
128349 | 42311.5842 | 1 | LI18 | 11 | 4 | HARDWARE | HW Problem | 18-1 Mod PFN compartment fan is dead
128569 | 42319.875 | 1.5 | LI20 | 71 | 4 | HARDWARE | HW Problem | L0A (20-7) mod fault
128596 | 42320.5601 | 1 | LI12 | 11 | 4 | HARDWARE | HW Problem | 12-1: Red HV On lamp not lit with MOD on.
128861 | 42331.395 | 0.5 | LI09 | 11 | 4 | HARDWARE | HW Problem | 9-1 has tripped off and will not reset remotely.
128878 | 42332.22 | 1 | LI12 | 41 | 4 | HARDWARE | HW Problem | 12-4 WGV flt
128879 | 42332.2245 | 1 | LI17 | 41 | 4 | HARDWARE | HW Problem | 17-4 Mod flt
127052 | 42267.6166 | 1.5 | LI26 | 41 | 4 | HARDWARE | HW Problem | 26-4: Delta Temp fault. Solenoid body hot to touch. Powered down.
127094 | 42268.6635 | 1 | LI24 | 11 | 4 | HARDWARE | HW Problem | 24-1: Waveguide 2 water fault.
127098 | 42268.7538 |  | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7 no RF drive but drive request normal please check SSSBST
127112 | 42269.2617 | 0.5 | LI21 | 71 | 4 | HARDWARE | HW Problem | 21-7: The HV ON light is not working. It needs to replace.
127124 | 42269.5535 | 1 | LI25 | 21 | 4 | HARDWARE | HW Problem | 25-2: ACOC faults.
127242 | 42271.7761 | 3 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4 has oil leak from modulator cabinet.
127245 | 42271.8213 | 2 | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 tripping off on HVR often. Resets each time, but then trips again a short time later.
128293 | 42309.722 | 1 | LI29 | 21 | 4 | HARDWARE | HW Problem | Klystron 29-2 HVR and won't reset remotely.
128345 | 42311.5773 | 1 | LI17 | 41 | 4 | HARDWARE | HW Problem | 17-4 Mod PFN compartment fan is dead
128418 | 42313.4814 | 1 | LI03 | 61 | 4 | HARDWARE | HW Problem | Kly. LI 03 - 6 Phase jitter is 0.5deg.s
128500 | 42317.1756 |  | LI02 | 31 | 4 | HARDWARE | HW Problem | 2-3 has a cable fault - MCC tried swapping the PIOP with no change. I cycled power to the MKSU then the PAD with no change. With PAD substitution box installed, MKSU is happy, station is on - Need PAD checked / swapped?
128664 | 42322.644 | 1 | LI24 | 41 | 4 | HARDWARE | HW Problem | Klys 24-4 showing 24 V fault.
126891 | 42256.1479 |  | LI30 | 71 | 4 | HARDWARE | HW Problem | 30-7 VGC is faulted LO" but no fault on the MKSU or modulator
127248 | 42272.0961 | 1 | LI27 | 51 | 4 | HARDWARE | HW Problem | Klystron LI27-51 shows 214 accumulated water faults on klystron errors panel.
127622 | 42284.5447 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X feedback PAD channel 0 (L1X In) broken. Its amplitude changes by 40% during typical phase scan. Channel 1 (L1X Out) seems fine and tracks the changes the amplitude set point does during scan.
127622 | 42284.5447 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X feedback PAD channel 0 (L1X In) broken. Its amplitude changes by 40% during typical phase scan. Channel 1 (L1X Out) seems fine and tracks the changes the amplitude set point does during scan.
127674 | 42286.4289 | 1 | LI19 | 11 | 4 | HARDWARE | HW Problem | 19-1 - Klystron vacuum is rising with spikes up to 7.7E-7. Also has kly focus magnet problem (open?) and a leaky window.
127976 | 42296.1541 | 4 | LI21 | 41 | 4 | HARDWARE | HW Problem | 21-4 has amplitude jitter and has tripped off twice in the last 5 minutes.
127976 | 42296.1541 | 1 | LI21 | 41 | 4 | HARDWARE | HW Problem | 21-4 has amplitude jitter and has tripped off twice in the last 5 minutes.
127987 | 42296.4717 | 3 | LI20 | 71 | 4 | HARDWARE | HW Problem | L0A tripped (HVR or MOD, don't remember) and only BVLTs came back after remote reset. Amplitude still showed zero. This briefly held off the beam.
128353 | 42311.5936 | 1 | LI12 | 61 | 4 | HARDWARE | HW Problem | 12-6 Thyratron compartment fan needs to be replaced making loud noise and possibly tripping breaker
128404 | 42313.3306 |  | LI24 | 21 | 4 | HARDWARE | HW Problem | Ll24-2 klystron water fault
128533 | 42318.3045 |  | LI29 | 11 | 4 | HARDWARE | HW Problem | LI29 KLYS10, TBR-WGV Leaks--MFDV & KTL to work together to add large pumping system to pump on sector while WGV is closed and KTL does the tube swap. is faulted on Klystron Filament U/C and won't reset, reading 175 volts and zero current (3.29 amps set point) Suspected waveguide valve leak. TBR drop delayed until PAMM.
128773 | 42327.414 | 1 | LI14 | 11 | 4 | HARDWARE | HW Problem | 14-1: Replace clip to upper coil #1.
128773 | 42327.414 | 1 | LI14 | 11 | 4 | HARDWARE | HW Problem | 14-1: Replace clip to upper coil #1.
128828 | 42329.3762 | 1.5 | LI21 | 71 | 4 | HARDWARE | HW Problem | Kly Mod 21-7: Mod fault, no reset.
128871 | 42331.6209 | 2 | LI22 | 21 | 4 | HARDWARE | HW Problem | Please range 22-2 Thyratron stopped firing
128873 | 42331.7777 | 1.5 | LI23 | 71 | 4 | HARDWARE | HW Problem | Klys 23-7 tripped, no reset.
126984 | 42264.2651 | 1 | LI22 | 31 | 4 | HARDWARE | HW Problem | 22-3 PFN cap needs replace. RHS 4th up from bottom. Coil and straps side of caps
126996 | 42265.2786 | 1 | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7 PFN cap needs replace. LHS 4th down.
127096 | 42268.6654 | 1 | LI30 | 71 | 4 | HARDWARE | HW Problem | 30-7: Station "motorboating". Needs Enerpro/HV Regulator checked out.
127291 | 42273.3298 | 0.5 | LI09 | 51 | 4 | HARDWARE | HW Problem | 9-5 Faulty VGC needs replace
127297 | 42273.8927 | 1 | LI22 | 61 | 4 | HARDWARE | HW Problem | station 22-6 tripped on MKSU trig fault
128721 | 42325.4509 | 2 | LI11 | 11 | 4 | HARDWARE | HW Problem | 11-7 while opening the Magnet Circuit Breaker to TBR the klystron, noticed a odor coming from the Circuit Breaker.
128732 | 42325.6736 | 1 | LI20 | 61 | 4 | HARDWARE | HW Problem | 20-6 improve PFN votage regulation.
128732 | 42325.6736 | 1.5 | LI20 | 61 | 4 | HARDWARE | HW Problem | 20-6 improve PFN votage regulation.
128734 | 42325.6827 | 0.5 | LI20 | 81 | 4 | HARDWARE | HW Problem | 20-8 improve PFN votage regulation
128734 | 42325.6827 | 0.1 | LI20 | 81 | 4 | HARDWARE | HW Problem | 20-8 improve PFN votage regulation
128734 | 42325.6827 | 0.25 | LI20 | 81 | 4 | HARDWARE | HW Problem | 20-8 improve PFN votage regulation
126873 | 42255.3095 | 0.5 | LI27 | 21 | 4 | HARDWARE | HW Problem | 27-2 Modulator, No klystron filament power.
128007 | 42297.0256 | 1 | LI13 | 11 | 4 | HARDWARE | HW Problem | 13-1 MOD fault, won't reset.
128020 | 42297.4489 | 0.5 | LI03 | 21 | 4 | HARDWARE | HW Problem | Water leak above the FIAT rack at the steel I beam water is running over to the modulator and onto the FIAT rack.
128056 | 42298.2413 | 1 | LI26 | 61 | 4 | HARDWARE | HW Problem | 26-6 and 26-7 phase shifters are sticky, preventing accurate phase scans and leaving the stations in a phase quagmire.
128057 | 42298.2613 |  | LI28 | 41 | 4 | HARDWARE | HW Problem | Kly. LI 28 - 4 has Phase jitter = to 3.18 and Beam volt jitter of 1.38.
128061 | 42298.3084 | 0.25 | LI25 | 41 | 4 | HARDWARE | HW Problem | Kly. LI 25 - 4 has tripped off again. Won't reset. Informed PEM.
128133 | 42300.3726 | 6 | LI26 | 71 | 4 | HARDWARE | HW Problem | 26-7 Modulator, PFN capacitor failed, damaging the PFN rack and Modulator doors.
128133 | 42300.3726 | 3 | LI26 | 71 | 4 | HARDWARE | HW Problem | 26-7 Modulator, PFN capacitor failed, damaging the PFN rack and Modulator doors.
128133 | 42300.3726 | 1 | LI26 | 71 | 4 | HARDWARE | HW Problem | 26-7 Modulator, PFN capacitor failed, damaging the PFN rack and Modulator doors.
128133 | 42300.3726 | 3 | LI26 | 71 | 4 | HARDWARE | HW Problem | 26-7 Modulator, PFN capacitor failed, damaging the PFN rack and Modulator doors.
128133 | 42300.3726 | 2 | LI26 | 71 | 4 | HARDWARE | HW Problem | 26-7 Modulator, PFN capacitor failed, damaging the PFN rack and Modulator doors.
128135 | 42300.4399 | 0.5 | LI23 | 21 | 4 | HARDWARE | HW Problem | 23-1: Step changes in RF AMPL?
128673 | 42323.6114 | 0.5 | LI21 | 11 | 4 | HARDWARE | HW Problem | 21-1: Replace lower PFN clip on #1 coil.
128733 | 42325.68 | 0.2 | LI20 | 71 | 4 | HARDWARE | HW Problem | 20-7 improve PFN votage regulation.
128733 | 42325.68 | 0.5 | LI20 | 71 | 4 | HARDWARE | HW Problem | 20-7 improve PFN votage regulation.
128772 | 42327.4121 | 0.5 | LI23 | 21 | 4 | HARDWARE | HW Problem | DQing low on 23-2
128893 | 42332.6719 | 0.5 | LI25 | 21 | 4 | HARDWARE | HW Problem | 25-2: Upper PFN strap #1 needs replacing.
128894 | 42332.6742 | 1 | LI25 | 31 | 4 | HARDWARE | HW Problem | 25-3: Lower PFN strap #7 and lower PFN coil clip need to be replaced.
128895 | 42332.6757 | 0.5 | LI25 | 71 | 4 | HARDWARE | HW Problem | 25-7: Upper strap #1 needs replacing.
128896 | 42332.6781 | 0.5 | LI26 | 11 | 4 | HARDWARE | HW Problem | 26-1: Lower PFN clip needs to be replaced.
128898 | 42332.6968 | 0.5 | LI26 | 21 | 4 | HARDWARE | HW Problem | 26-2: Upper PFN strap #6 needs replacing.
128899 | 42332.6987 | 0.5 | LI26 | 41 | 4 | HARDWARE | HW Problem | 26-4: Lower PFN clip needs to be replaced.
128904 | 42332.8595 |  | LI20 | 61 | 4 | HARDWARE | HW Problem | 20-6 Bad high power attenuator. Drive does not go to zero.
127249 | 42272.0986 | 1 | LI25 | 61 | 4 | HARDWARE | HW Problem | Klystron LI25-61 has accumulated 609 water faults as shown on klystron errors panel.
127521 | 42279.7118 | 0.75 | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3 mod faulted and won't reset remotely
127614 | 42284.2866 | 1 | LI24 | 51 | 4 | HARDWARE | HW Problem | 24-5 Motor-Boating Loudly
127643 | 42285.1182 | 1 | LI03 | 11 | 4 | HARDWARE | HW Problem | 3-1 Has a water leak between Fiat rack and klystron.
127712 | 42286.7859 |  | LI16 | 81 | 4 | HARDWARE | HW Problem | 16-8 has water summary fault.
127757 | 42287.942 | 0.5 | LI17 | 31 | 4 | HARDWARE | HW Problem | 17-3 Trips on over voltage and over current
128483 | 42315.739 | 0.75 | LI22 | 21 | 4 | HARDWARE | HW Problem | Klys 22-2 showing beam volts jitter.
128751 | 42326.461 | 1 | LI29 | 71 | 4 | HARDWARE | HW Problem | 29-7 Elevated klystron vacuum.
128753 | 42326.5399 | 1 | LI28 | 61 | 4 | HARDWARE | HW Problem | 28-6 trips -> HVR something like roughly every twelve minutes
128760 | 42327.1749 | 1 | LI16 | 51 | 4 | HARDWARE | HW Problem | 16-5 Mod has HV on Thyratron not firing.
127189 | 42270.6945 | 1 | LI23 | 11 | 4 | HARDWARE | HW Problem | 23-1: Accelerator 1 water faults (intermittent).
128211 | 42305.2533 |  | LI23 | 11 | 4 | HARDWARE | HW Problem | 23-1 phase jitter.
128226 | 42305.6011 | 0.5 | LI19 | 11 | 4 | HARDWARE | HW Problem | 19-1 Has water leaking at the union of the Phase Reference Line drop-out coupler to the PAD hose.
128227 | 42305.6018 | 1 | LI10 | 31 | 4 | HARDWARE | HW Problem | 10-3 PFN compartment cooling fan IS dead
128909 | 42332.8977 | 0.5 | LI20 | 71 | 4 | HARDWARE | HW Problem | Klys 20-8 (L0B) showing beam volts dropouts.
128912 | 42332.9919 | 1.5 | LI20 | 81 | 4 | HARDWARE | HW Problem | L0B (20-8) showing a lot of beam volts drop outs
128912 | 42332.9919 | 1.5 | LI20 | 81 | 4 | HARDWARE | HW Problem | L0B (20-8) showing a lot of beam volts drop outs
128912 | 42332.9919 | 1 | LI20 | 81 | 4 | HARDWARE | HW Problem | L0B (20-8) showing a lot of beam volts drop outs
128945 | 42334.8421 | 0.75 | LI23 | 51 | 4 | HARDWARE | HW Problem | Kly. LI23 - 5 tripped and won't reset.
128945 | 42334.8421 | 0.75 | LI23 | 51 | 4 | HARDWARE | HW Problem | Kly. LI23 - 5 tripped and won't reset.
129711 | 42384.677 | 1 | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3 HVR, no remote reset.
130028 | 42399.4344 | 3.5 | LI09 | 11 | 4 | HARDWARE | HW Problem | Klystron LI09-11 is tripped off and will not reset remotely.
129905 | 42394.7106 | 1 | LI28 | 31 | 4 | HARDWARE | HW Problem | Klystron 28-3 showing bad cable status. We think it may be due to a bad PIOP or MKSU Tried IPL and reset, but did not clear faut.
129967 | 42397.1295 | 1 | LI20 | 71 | 4 | HARDWARE | HW Problem | L0A having lots of beam volts jitter
129993 | 42398.1696 | 0.5 | LI09 | 11 | 4 | HARDWARE | HW Problem | Klystron LI09-11 is tripped off on Mod Fault and will not reset remotely.
129997 | 42398.1995 | 0.5 | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 Tripped on HVR without remote reset.
130035 | 42400.3458 | 1 | LI02 | 81 | 4 | HARDWARE | HW Problem | Klystron LI02-81 has a WG vacuum fault and will not reset remotely. Check WG VGC calibration before transferring to another group.
130035 | 42400.3458 | 0.5 | LI02 | 81 | 4 | HARDWARE | HW Problem | Klystron LI02-81 has a WG vacuum fault and will not reset remotely. Check WG VGC calibration before transferring to another group.
130066 | 42401.8086 |  | LI25 | 61 | 4 | HARDWARE | HW Problem | TBR - 25-6 tripping on klystron vacuum,
128921 | 42333.4359 | 0.1 | LI29 | 61 | 4 | HARDWARE | HW Problem | 29-6 MOD fault, won't reset.
129154 | 42346.4265 | 4 | LI20 | 61 | 4 | HARDWARE | HW Problem | 20-6 (aka Gun): Per FJD widen PFN pulse for a "good" flat-top of 2.2 uS.
130116 | 42404.0633 | 1 | LI20 | 61 | 4 | HARDWARE | HW Problem | LCLS Gun 20-6 tripping on HVR. Resets remotely but causes intermittent loss of delivery.
130116 | 42404.0633 | 1 | LI20 | 61 | 4 | HARDWARE | HW Problem | LCLS Gun 20-6 tripping on HVR. Resets remotely but causes intermittent loss of delivery.
130196 | 42407.71 | 1.5 | LI11 | 81 | 4 | HARDWARE | HW Problem | Klys 11-8 showing MKT.
130865 | 42443.6779 | 1 | LI28 | 11 | 4 | HARDWARE | HW Problem | Klystron 28-1 MOD faulted and can't be reset remotely. It's in local mode
130896 | 42445.5931 | 1 | LI23 | 41 | 4 | HARDWARE | HW Problem | 23-4 Mod D Q ing cap in thy comp has a badly burnt connection
130904 | 42445.7329 |  | LI21 | 71 | 4 | HARDWARE | HW Problem | Klystron 21-7 Focus Supply communication problem between genesys ps and EPICS. Also PLC shows FAULT but modulator is ON. Steele 3/16/16 21:00: Modulator is NOT on. Klystron was replaced this morning and station has not been set up or turned on yet. thyratron is off, core bias is off, etc. Has loto been removed from focus supply breaker?
129987 | 42397.837 | 1 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 is HVR and will not reset remotely
130000 | 42398.3712 |  | LI03 | 71 | 4 | HARDWARE | HW Problem | 3-7 MOD fault.
130019 | 42398.7305 | 0.75 | LI14 | 81 | 4 | HARDWARE | HW Problem | Klys 14-8 MOD faulted and won't reset remotely
130032 | 42399.8244 | 0.75 | LI13 | 11 | 4 | HARDWARE | HW Problem | station 13-1 mod faulted
130046 | 42401.3652 | 2 | LI26 | 51 | 4 | HARDWARE | HW Problem | 26-5: 600Vac, the AC current displays on the tounch panel is 66.6 Amps. it should be 90 to 100A.
130046 | 42401.3652 | 2 | LI26 | 51 | 4 | HARDWARE | HW Problem | 26-5: 600Vac, the AC current displays on the tounch panel is 66.6 Amps. it should be 90 to 100A.
130046 | 42401.3652 | 3 | LI26 | 51 | 4 | HARDWARE | HW Problem | 26-5: 600Vac, the AC current displays on the tounch panel is 66.6 Amps. it should be 90 to 100A.
130061 | 42401.6132 | 3 | LI28 | 51 | 4 | HARDWARE | HW Problem | 28-4: PFN has unreasonable slope (1.1), unable to fix-up timing.
130108 | 42403.6621 | 0.8 | LI28 | 51 | 4 | HARDWARE | HW Problem | 28-5 FLK external faults. RE is higher than other sector stations. Please check.
130111 | 42403.725 | 2 | LI28 | 11 | 4 | HARDWARE | HW Problem | 28-1 mod faulted and trips almost immediately after reset.
130115 | 42403.9753 | 2 | LI20 | 71 | 4 | HARDWARE | HW Problem | Klystron 20-7 showing periodic beam volts dropouts.
128966 | 42336.4077 | 1 | LI29 | 31 | 4 | HARDWARE | HW Problem | 29-3 BV jitters
128983 | 42337.8544 | 1 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 trips repeatedly, especially when it is cold out. 26 trips owl shift 11/29, starting to trip frequently again as of 8pm 11/29.
129850 | 42391.4258 | 1 | LI20 | 93 | 4 | HARDWARE | HW Problem | 20-3C is tripped, doesn't reset. VP-400 interlock is faulted.
129850 | 42391.4258 | 4 | LI20 | 93 | 4 | HARDWARE | HW Problem | 20-3C is tripped, doesn't reset. VP-400 interlock is faulted.
130005 | 42398.4314 | 0.5 | LI20 | 41 | 4 | HARDWARE | HW Problem | 20-4B (XTCAV) has tripped off and does not reset.
130020 | 42398.7455 | 1 | LI02 | 81 | 4 | HARDWARE | HW Problem | Klystron 2-8 WGV fault
130092 | 42403.4127 | 1 | LI30 | 11 | 4 | HARDWARE | HW Problem | LCLS Modulators: 27-5,29-6,30-1,30-3,30-4 - BVs need to adjust to spec.
130212 | 42408.5593 | 1 | LI13 | 41 | 4 | HARDWARE | HW Problem | 13-4: Seeing some beam volts jitter.
130233 | 42409.5736 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X no amplitude (21-2)
129026 | 42339.5555 | 1 | LI29 | 31 | 4 | HARDWARE | HW Problem | 29-3: Seeing some beam volts jitter.
129204 | 42348.5348 | 0.5 | LI17 | 41 | 4 | HARDWARE | HW Problem | 17-4: "KLV" fault.
129299 | 42353.5193 | 2 | LI20 | 93 | 4 | HARDWARE | HW Problem | 20-3C: found at "FLK". Reset and on to constant RE's.
129357 | 42368.5531 | 1 | LI28 | 81 | 4 | HARDWARE | HW Problem | TBR LI 28-8, Review tube for water leak at hole between red mag & tank.
129357 | 42368.5531 | 1 | LI28 | 81 | 4 | HARDWARE | HW Problem | TBR LI 28-8, Review tube for water leak at hole between red mag & tank.
129444 | 42374.7206 |  | LI23 | 11 | 4 | HARDWARE | HW Problem | 23-1 klystron output dropped by about 50% over the last 12 hours. Modulator Beam Volts and current look normal. Drive setting at 99%. Something weird with input RF?
129446 | 42374.8162 | 0.5 | LI27 | 11 | 4 | HARDWARE | HW Problem | 27-1 PAD does not read any phase or amplitude. FEPK indicates normal power level. See 129512.
129449 | 42374.8303 | 0.8 | LI25 | 11 | 4 | HARDWARE | HW Problem | 25-1 intermittent DeQuing under current faults. Current fluctuates between 10 and 9.5 amps.
129864 | 42391.6387 | 1 | LI25 | 51 | 4 | HARDWARE | HW Problem | Kly Mod 25-5: tripped on ACOC/HVOC.
129866 | 42391.7561 | 1 | LI28 | 21 | 4 | HARDWARE | HW Problem | Klystron 28-2 MOD faulting incessantly
130144 | 42404.9985 | 1 | LI21 | 21 | 4 | HARDWARE | HW Problem | station 21-2 WGV won't clear after PAC reboot
130339 | 42415.9052 | 2 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 tripped, can't reset remotely.
130340 | 42415.9146 | 2 | LI06 | 11 | 4 | HARDWARE | HW Problem | Station 6-1 tripping repeatedly. RE faults
130417 | 42418.8235 | 1 | LI17 | 21 | 4 | HARDWARE | HW Problem | 17-2 HV meter not reading
130149 | 42405.4101 | 0.5 | LI11 | 11 | 4 | HARDWARE | HW Problem | 11-1 MOD fault.
130195 | 42407.6917 | 0 | LI27 | 11 | 4 | HARDWARE | HW Problem | Klys 27-1 tripping frequently.
130226 | 42409.4134 | 0.5 | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3: MOD fault
130226 | 42409.4134 | 1 | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3: MOD fault
130226 | 42409.4134 | 1 | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3: MOD fault
130235 | 42409.602 |  | LI05 | 11 | 4 | HARDWARE | HW Problem | 5-1 KMC fault.
130318 | 42413.3477 | 3 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3 FLK This station has multiple faults. One shows the Thyr Res current jumping from 22A to 0. The supply voltage controller to it the Green LED flashes on and off. The Thyratron HTR timing keeps reseting never timing out. The Doors flt counts door flts but the Fault monitor is green ok for doors.Looks like this started at 5:30 . Turned over to days.TS&R.
130320 | 42413.4143 | 0.75 | LI17 | 71 | 4 | HARDWARE | HW Problem | 17-7 MOD fault no reseet
130373 | 42417.2761 | 1 | LI04 | 11 | 4 | HARDWARE | HW Problem | 4-1 Water leak in front of FIAT rack
130390 | 42417.77 | 1 | LI20 | 51 | 4 | HARDWARE | HW Problem | ACR request for PEM to recover 20-5 through 20-8 after PPS interruption from PAMM.
130399 | 42418.2866 | 1 | LI09 | 11 | 4 | HARDWARE | HW Problem | 9-1 WGV fault.
129062 | 42341.0188 |  | LI06 | 21 | 4 | HARDWARE | HW Problem | 6-2 WGV
129119 | 42344.4854 | 7 | LI22 | 31 | 4 | HARDWARE | HW Problem | Kly. LI 22 - 3 faulted on "MOD VVS" Voltage fault.
129215 | 42349.1895 | 0.5 | LI04 | 41 | 4 | HARDWARE | HW Problem | Klystron 4-4 MOD faulted with no remote reset
129215 | 42349.1895 | 1 | LI04 | 41 | 4 | HARDWARE | HW Problem | Klystron 4-4 MOD faulted with no remote reset
129418 | 42374.2841 | 8 | LI28 | 51 | 4 | HARDWARE | HW Problem | 28-5 T20 began humming and CB7 tripped at 18kv when ramping up I inspected main breaker pan / contactors, SCR pan and T20 but didn't see any signs of damaged components.
129450 | 42374.8364 | 0 | LI21 | 31 | 4 | HARDWARE | HW Problem | 21-3 epics shows modulator has 0 Amps DeQuing.
130335 | 42415.3313 | 0.5 | LI30 | 41 | 4 | HARDWARE | HW Problem | 30-4: DQ curr is low = 7.5A
130335 | 42415.3313 | 0.1 | LI30 | 41 | 4 | HARDWARE | HW Problem | 30-4: DQ curr is low = 7.5A
130395 | 42417.8573 | 1 | LI23 | 41 | 4 | HARDWARE | HW Problem | Klystron 23-4 MOD faulting frequently. We are able to reset it, but then trips off after a few minutes.
130401 | 42418.2941 |  | LI12 | 11 | 4 | HARDWARE | HW Problem | 12-1 VVS breaker is tripped
130444 | 42419.8263 | 1.2 | LI15 | 41 | 4 | HARDWARE | HW Problem | 15-4 tripped on WGV Wave Guide Vacuum, can't recover remotely. WGV problem>
130448 | 42420.2744 | 3 | LI29 | 61 | 4 | HARDWARE | HW Problem | Klys 29-6 trips off every 20-30 minutes. Resets okay, but is a nuance. Replace the thyratron with a new one and prepare it for weekend high energy beam.
130458 | 42421.0998 | 3 | LI20 | 71 | 4 | HARDWARE | HW Problem | L0A (20-7) beam volt dropouts, clearly seen on downstream BPM and overall LCLS performance.
130458 | 42421.0998 | 2 | LI20 | 71 | 4 | HARDWARE | HW Problem | L0A (20-7) beam volt dropouts, clearly seen on downstream BPM and overall LCLS performance.
129296 | 42353.4512 | 1 | LI09 | 11 | 4 | HARDWARE | HW Problem | Kly. LI 09 - 1 tripped on "MOD" Fault and won't reset.
129676 | 42383.3117 | 1 | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3 BV jitter sigma 0.3kV
129820 | 42390.4564 |  | LI30 | 41 | 4 | HARDWARE | HW Problem | Sector 30 - The Klystron Cooling System water flows appears to be the lowest flow at 30-4. That station has been tripping on KMC, which could be improved by better flow.
129861 | 42391.5553 | 2 | LI13 | 21 | 4 | HARDWARE | HW Problem | 13-2 needs troubleshooting -- trips VVS breaker
129994 | 42398.1706 |  | LI24 | 11 | 4 | HARDWARE | HW Problem | 24-1 tripping on reflected energy fairly frequently. Amplitude needs to be slightly reduced.
129996 | 42398.1939 | 0.5 | LI02 | 61 | 4 | HARDWARE | HW Problem | Klystron LI02-61 is faulted on WG Vacuum and will not reset remotely.
130494 | 42423.5126 | 0.5 | LI09 | 41 | 4 | HARDWARE | HW Problem | 9-4: MOD fault.
130495 | 42423.5139 |  | LI12 | 11 | 4 | HARDWARE | HW Problem | 12-1: MOD fault.
130495 | 42423.5139 | 1.5 | LI12 | 11 | 4 | HARDWARE | HW Problem | 12-1: MOD fault.
130495 | 42423.5139 | 1 | LI12 | 11 | 4 | HARDWARE | HW Problem | 12-1: MOD fault.
130495 | 42423.5139 | 2 | LI12 | 11 | 4 | HARDWARE | HW Problem | 12-1: MOD fault.
130499 | 42423.53 | 0.75 | LI26 | 11 | 4 | HARDWARE | HW Problem | 26-1: RF AMPL High OOT due to beam volts high.
130506 | 42423.7658 | 0.75 | LI05 | 31 | 4 | HARDWARE | HW Problem | Klystron 5-3 WGV fault
129116 | 42344.3095 | 0 | LI22 | 61 | 4 | HARDWARE | HW Problem | Klys 22-6 is tripping frequently
129251 | 42351.1932 | 1 | LI11 | 61 | 4 | HARDWARE | HW Problem | Klystron 11-6 MOD faulted with no remote reset
129447 | 42374.8215 | 2 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 MKSU not communicating with PIOP.
129559 | 42377.7483 | 6 | LI09 | 61 | 4 | HARDWARE | HW Problem | 9-6 supervisory alarm. Latching box indicating fire alarm, reset and it comes back. Keyed off mod and alarm clears.
129626 | 42381.5399 | 0.5 | LI30 | 41 | 4 | HARDWARE | HW Problem | 30-4 is showing Electro-Magnet Current fault and is not contributing to the beam.
129725 | 42385.5572 |  | LI30 | 41 | 4 | HARDWARE | HW Problem | 30-4 is showing KMC Electro-Magnet Current fault.
129725 | 42385.5572 | 1.5 | LI30 | 41 | 4 | HARDWARE | HW Problem | 30-4 is showing KMC Electro-Magnet Current fault.
129810 | 42389.624 | 1 | LI30 | 41 | 4 | HARDWARE | HW Problem | 30-4 tripped on electromagnet current.
129909 | 42395.3395 | 2 | LI20 | 61 | 4 | HARDWARE | HW Problem | Station gauge faults on LCLS Gun, L0A, L0B. All stations went WGV.
130483 | 42423.2004 | 0.5 | LI26 | 31 | 4 | HARDWARE | HW Problem | 26-3 HVR no remote reset
130498 | 42423.5255 | 1 | LI28 | 81 | 4 | HARDWARE | HW Problem | 28-8: Seeing some beam volts jitter.
130535 | 42425.4503 | 1.5 | LI27 | 51 | 4 | HARDWARE | HW Problem | 27-5 low uk.
130535 | 42425.4503 | 1.5 | LI27 | 51 | 4 | HARDWARE | HW Problem | 27-5 low uk.
130543 | 42425.6164 | 2.5 | LI27 | 61 | 4 | HARDWARE | HW Problem | 27-6 trips off immediately upon being brought on.
130576 | 42427.1524 | 0.2 | LI23 | 81 | 4 | HARDWARE | HW Problem | Klys 23-8 phase not tracking well during phase scans. See: http://physics-elog.slac.stanford.edu/lclselog/show.jsp?dir=/2016/08/27.02&pos=2016-02-27T01:43:57
130583 | 42427.5106 | 2 | LI05 | 71 | 4 | HARDWARE | HW Problem | Kly Mod 5-7: no triggers present, PEM will replace MKSU.
130588 | 42428.3239 | 1 | LI02 | 81 | 4 | HARDWARE | HW Problem | 2-8 WGV fault
130589 | 42428.3654 | 0.8 | LI24 | 31 | 4 | HARDWARE | HW Problem | Kly. LI24 - 3 tripping on "EVO" faults.
130590 | 42428.4576 | 0 | LI05 | 11 | 4 | HARDWARE | HW Problem | 5-1 trips frequently and has 13 link faults.
129008 | 42338.7058 | 1 | LI28 | 11 | 4 | HARDWARE | HW Problem | Klystron 28-1 HVR and will not reset remotely. It is in local.
129092 | 42342.3804 | 4 | LI11 | 61 | 4 | HARDWARE | HW Problem | 11-6 has tripped CB7 several times in the last week. FACET is complaining about the disruption.
129102 | 42342.5654 |  | LI20 | 71 | 4 | HARDWARE | HW Problem | L0A slow modulator HV ramp up. The modulators typically ramp up to full voltage in 2 to 3 seconds, L0A is taking about 6 seconds. This slow ramp up leads to station operating at low power and subsequent multi-pacting and vacuum activity and vacuum trips.
129102 | 42342.5654 | 1 | LI20 | 71 | 4 | HARDWARE | HW Problem | L0A slow modulator HV ramp up. The modulators typically ramp up to full voltage in 2 to 3 seconds, L0A is taking about 6 seconds. This slow ramp up leads to station operating at low power and subsequent multi-pacting and vacuum activity and vacuum trips.
129115 | 42343.6213 | 1 | LI21 | 71 | 4 | HARDWARE | HW Problem | Klys 21-7 showing beam volts jitter
129250 | 42350.9886 |  | LI22 | 61 | 4 | HARDWARE | HW Problem | Book keeping entry 22-2 tripped no remote reset due to LOCAL control PEM switched to REMOTE control (0.1 hours).
129690 | 42383.9073 | 1 | LI02 | 81 | 4 | HARDWARE | HW Problem | Klys 2-8 tripping frequently on TOC according to kdisp.
130031 | 42399.61 | 1 | LI28 | 51 | 4 | HARDWARE | HW Problem | Kly Mod 28-5: FLK'd off on amplitude faults.
130548 | 42425.8753 | 0.5 | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7 HVR. Trips immediately on reset.
130593 | 42429.2688 |  | LI05 | 71 | 4 | HARDWARE | HW Problem | 5-7 CABLE fault, swapping PIOP cables with 5-8, problem moves to 5-8, looks like the PIOP for 5-7 is bad
130604 | 42429.7688 | 2 | LI30 | 71 | 4 | HARDWARE | HW Problem | 30-7: Water, presumably ACS LCW, dripping from SLED cavity. Water appears to be coming from the copper tubing around the SLED cavity. See attached picture.
130604 | 42429.7688 | 1 | LI30 | 71 | 4 | HARDWARE | HW Problem | 30-7: Water, presumably ACS LCW, dripping from SLED cavity. Water appears to be coming from the copper tubing around the SLED cavity. See attached picture.
130616 | 42430.4099 | 4 | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7 has had six vacuum faults in the last 1.5 hours.
130617 | 42430.4103 | 0.5 | LI28 | 51 | 4 | HARDWARE | HW Problem | 28-5 is MOD faulted.
130619 | 42430.4186 | 0 | LI28 | 51 | 4 | HARDWARE | HW Problem | Klys 28-5 is MOD fault wont reset
130637 | 42430.7849 | 1.5 | LI16 | 41 | 4 | HARDWARE | HW Problem | Klystron LI16-41 is tripped off and will not reset remotely.
130025 | 42399.1161 | 1 | LI13 | 11 | 4 | HARDWARE | HW Problem | Klystron LI13-11 is tripped off and will not reset remotely.
130062 | 42401.6341 | 0 | LI29 | 61 | 4 | HARDWARE | HW Problem | 29-6 has had 30 klystron over-voltage faults in the last 24 hours.
130073 | 42402.4108 |  | LI18 | 51 | 4 | HARDWARE | HW Problem | Kly. LI 18 - 5 AMPLIT. drooped from 2200 Feb. 30 / 60.6 to today which reads 53.70 .
130128 | 42404.3758 |  | DR13 | 1 | 4 | HARDWARE | HW Problem | Kly. DR 13 - 1 PAD Trigger (PADO) is out of tol. Reads -1.00 but the exp. is -0.20.
130620 | 42430.4193 | 0 | LI23 | 71 | 4 | HARDWARE | HW Problem | Klys 23-7 is HVR, will reset but trips off immediately.
130634 | 42430.7419 | 1 | LI28 | 11 | 4 | HARDWARE | HW Problem | Klystron 28-1 HVR and won't reset since it is in local
130651 | 42431.5439 | 1 | LI28 | 11 | 4 | HARDWARE | HW Problem | 28-1: Replace PFN clip on upper coil.
130654 | 42431.6012 | 1 | LI20 | 81 | 4 | HARDWARE | HW Problem | 20-8 Has a EOLC stack that needs to be replaced
130663 | 42431.6802 | 1 | LI28 | 51 | 4 | HARDWARE | HW Problem | 28-5 sick, showing intermittent MOD faults and very low amplitude when MOD on.
130663 | 42431.6802 | 0.2 | LI28 | 51 | 4 | HARDWARE | HW Problem | 28-5 sick, showing intermittent MOD faults and very low amplitude when MOD on.
130668 | 42431.8492 | 2 | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7 mod fault
130673 | 42432.061 | 0.5 | LI23 | 41 | 4 | HARDWARE | HW Problem | 23-4 HVR, no remote reset. Station is in local.
128959 | 42335.9245 | 1 | LI23 | 11 | 4 | HARDWARE | HW Problem | Klystron 23-1 MOD faulting quite frequently. Mod faulting on water faults
129319 | 42354.5744 | 1 | LI27 | 21 | 4 | HARDWARE | HW Problem | Klys 27-2 tripping frequently. Thyratron reservoir at 7.5 V
130722 | 42433.7374 | 0.25 | LI09 | 11 | 4 | HARDWARE | HW Problem | Klystron LI09-11 is tripped off on WGV and will not reset remotely.
129304 | 42353.5952 | 0.5 | LI09 | 11 | 4 | HARDWARE | HW Problem | 9-1, Modulator; tripped on "MOD" fault and won't reset.
129729 | 42386.3417 | 1 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 tripping frequently
130048 | 42401.3743 | 0.5 | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2: 600Vac, on the touch panel the AC curr displays 29.9 Amps. It should be from 89 to 100 A.
130052 | 42401.4546 | 0.3 | DR01 | 1 | 4 | HARDWARE | HW Problem | 1-7B WGV fault, does not pass NIRP
130191 | 42407.0865 | 0.5 | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7 HVR no reset since the modulator is in LOCAL
130263 | 42410.4525 | 0.5 | LI11 | 71 | 4 | HARDWARE | HW Problem | 11-7 Has a KMC fault
130686 | 42432.4287 | 0.75 | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2 has tripped excessively over the last day. It then tripped off ~10 times in 5 mins.
130703 | 42432.9522 | 1 | DR03 | 1 | 4 | HARDWARE | HW Problem | SRTL compressor tripped, can't reset.
130723 | 42434.6266 | 1 | LI27 | 81 | 4 | HARDWARE | HW Problem | station 27-8 will not stay on
130733 | 42435.7449 | 3 | LI02 | 11 | 4 | HARDWARE | HW Problem | 2-1 went into heater delay at 17:50 on 3/6 without any cause apparent from the ACR.
130733 | 42435.7449 | 1 | LI02 | 11 | 4 | HARDWARE | HW Problem | 2-1 went into heater delay at 17:50 on 3/6 without any cause apparent from the ACR.
129486 | 42375.7809 | 0.75 | LI28 | 51 | 4 | HARDWARE | HW Problem | Klystron 28-5 MOD faulted and won't reset remotely
129520 | 42376.6739 | 1 | LI20 |  | 4 | HARDWARE | HW Problem | 20-6, 20-7, 20-8 tripped off from PPS.
130113 | 42403.8445 | 0.3 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4 AMJ. BV oscillation
130143 | 42404.9876 | 1 | LI28 | 71 | 4 | HARDWARE | HW Problem | station 28-7 beamvolts dropped a lot in last half hour
130276 | 42410.7532 | 1 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 is MOD faulted.
130773 | 42437.7285 | 1.5 | LI25 | 51 | 4 | HARDWARE | HW Problem | Klystron 25-5 MOD faults frequently. We are able to reset it, but it trips off again about 1 minute later.
130841 | 42442.6338 |  | LI20 | 93 | 4 | HARDWARE | HW Problem | station 20-3-c Mod faulted
130841 | 42442.6338 | 0.25 | LI20 | 93 | 4 | HARDWARE | HW Problem | station 20-3-c Mod faulted
129178 | 42347.529 | 0.25 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 prepare for on beam operation without HPRF. Moved to cater 127376
129452 | 42374.8913 | 1 | LI10 | 31 | 4 | HARDWARE | HW Problem | 10-3 has KMC fault
129471 | 42375.5629 | 0.5 | LI29 | 31 | 4 | HARDWARE | HW Problem | 29-3: Door interlock was broken on modulator door #5.
129522 | 42376.7195 | 1.5 | LI05 | 21 | 4 | HARDWARE | HW Problem | 5-2 has KMC fault
129672 | 42383.0348 |  | LI03 | 71 | 4 | HARDWARE | HW Problem | LI 03 - 7 has a "PAD" status = false.
129697 | 42384.3991 | 1 | LI13 | 81 | 4 | HARDWARE | HW Problem | 13-8 bad amplitude jitter.
130312 | 42412.5756 | 5 | LI23 | 81 | 4 | HARDWARE | HW Problem | 23-8: HV OFF and HV RESET commands on the PLC Interface board are not work.
130312 | 42412.5756 | 2 | LI23 | 81 | 4 | HARDWARE | HW Problem | 23-8: HV OFF and HV RESET commands on the PLC Interface board are not work.
130312 | 42412.5756 | 3 | LI23 | 81 | 4 | HARDWARE | HW Problem | 23-8: HV OFF and HV RESET commands on the PLC Interface board are not work.
130312 | 42412.5756 | 3 | LI23 | 81 | 4 | HARDWARE | HW Problem | 23-8: HV OFF and HV RESET commands on the PLC Interface board are not work.
130330 | 42414.5827 | 0.8 | LI17 | 41 | 4 | HARDWARE | HW Problem | 17-4 phenomenally large BV jitter (every fourth shot is zero) PEM dispatched.
130776 | 42437.8953 | 1 | LI27 | 81 | 4 | HARDWARE | HW Problem | Klystron 27-8 HVR with no remote reset
129436 | 42374.6593 | 1 | LI30 | 81 | 4 | HARDWARE | HW Problem | 30-8: WGV fault
129454 | 42374.9014 | 1 | LI28 | 61 | 4 | HARDWARE | HW Problem | 28-6 MOD fault no reset.
129455 | 42374.9198 | 1 | LI29 | 21 | 4 | HARDWARE | HW Problem | 29-2 mod faulted and won't reset remotely
129620 | 42381.3211 |  | LI18 | 61 | 4 | HARDWARE | HW Problem | DUPLICATE of 129563. 18-6 PIOP needs replace.
129646 | 42382.4054 | 0.5 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 HVR, won't reset remotely.
129884 | 42392.8322 | 1 | LI30 | 51 | 4 | HARDWARE | HW Problem | Klystron LI30-51 is tripped off and will not reset remotely.
130877 | 42444.6214 | 0.5 | LI28 | 31 | 4 | HARDWARE | HW Problem | Klys 28-3 MOD fault.
131068 | 42454.4339 | 1 | LI20 | 51 | 4 | HARDWARE | HW Problem | 20-5 TCAV0 RF amplitude & phase unstable. Amplitude jitter ~10%, Phase jitter ~5 degrees. Modulator pulses look stable on Klys Diagnostic PAD and PIOP FTP.
131511 | 42488.528 |  | LI24 | 11 | 4 | HARDWARE | HW Problem | Klys 24-1 receiving reflected energy faults.
131591 | 42495.7713 | 1 | LI30 | 51 | 4 | HARDWARE | HW Problem | LI 30 - 5 has 1.2 to 2.0 jitter. ( See MCC log book for pic.s )
128975 | 42337.1675 | 1 | LI04 | 21 | 4 | HARDWARE | HW Problem | station 4-2 is losing triggers almost continuously ('lost triggers in 1 min.' field is almost always red with typical numbers ranging from 5 to 15). MOD HVS 11/29/15 13:00 After much CAMAC investigation discovered that the cable between the PAD and the MKSU at the MKSU side is damaged...the plasitc connector housing is cracked, mechanically not sound, and needs to be replaced.
128989 | 42338.2013 | 0.5 | LI11 | 61 | 4 | HARDWARE | HW Problem | 11-6 MOD fault and won't reset
129297 | 42353.454 | 1 | LI11 | 61 | 4 | HARDWARE | HW Problem | Kly. LI 11 - 6 tripped on "MOD" fault and won't reset.
129768 | 42388.4756 | 1 | LI30 | 51 | 4 | HARDWARE | HW Problem | 30-5 trips frequently on over-voltage. Update 2/1/2016 - 30-5 tripping frequently showing MKT, MOD, and HVR faults, as well as occasional OV faults.
129956 | 42396.539 | 4 | LI21 | 41 | 4 | HARDWARE | HW Problem | LCLS Modulators: BV and DQ OOT adjustment at 21-4, 24-6, 26-5, 27-1, and 27-2.
130060 | 42401.5988 | 1 | LI28 |  | 4 | HARDWARE | HW Problem | Sectoer 28 water temperature is high causing the SLEDS to de-Tune.
130324 | 42413.8577 | 1 | LI20 | 71 | 4 | HARDWARE | HW Problem | L0A (20-7) jitter affecting FEL. Seeing phase, amplitude and beam volts jitter.
130356 | 42416.577 |  | DR13 | 1 | 4 | HARDWARE | HW Problem | WGV fault on DR13 compressor klystron
130438 | 42419.544 | 1.5 | LI25 | 21 | 4 | HARDWARE | HW Problem | 25-2: Seeing some pretty big beam volts jitter. Let 25-2 running until the thyratron goes dead. After that, put the system on maintenance mode.
130438 | 42419.544 | 3 | LI25 | 21 | 4 | HARDWARE | HW Problem | 25-2: Seeing some pretty big beam volts jitter. Let 25-2 running until the thyratron goes dead. After that, put the system on maintenance mode.
130875 | 42444.5861 | 1 | LI29 | 11 | 4 | HARDWARE | HW Problem | 29-1: Need to fix the AC current display for 600VAC.
130901 | 42445.6876 | 1 | LI29 | 41 | 4 | HARDWARE | HW Problem | station 29-4 won't stay on--mod faults repeatedly
128920 | 42333.4345 | 1 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 PEM Calibration
128920 | 42333.4345 | 0.5 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 PEM Calibration
128937 | 42334.1456 | 1 | LI22 | 31 | 4 | HARDWARE | HW Problem | 22-3 has repeated, intermittent Klystron Water faults as seen by blinking red light on MKSU chassis
128956 | 42335.7696 |  | LI02 | 81 | 4 | HARDWARE | HW Problem | Kly. LI 02 - 08 keeps tripping on "MOD" fault. Reset but it faults again between 1 - 2 min.s Error disp. shows 25 Ampl. faults. Mod water flow/fire alarm light occasionally flashing.
129286 | 42352.8108 | 2 | LI29 | 11 | 4 | HARDWARE | HW Problem | 29-1 tripping of every 5 minutes or so. Goes MOD fault and will reset. Pulled off the beam for now.
129286 | 42352.8108 | 1 | LI29 | 11 | 4 | HARDWARE | HW Problem | 29-1 tripping of every 5 minutes or so. Goes MOD fault and will reset. Pulled off the beam for now.
129286 | 42352.8108 | 0.5 | LI29 | 11 | 4 | HARDWARE | HW Problem | 29-1 tripping of every 5 minutes or so. Goes MOD fault and will reset. Pulled off the beam for now.
129286 | 42352.8108 |  | LI29 | 11 | 4 | HARDWARE | HW Problem | 29-1 tripping of every 5 minutes or so. Goes MOD fault and will reset. Pulled off the beam for now.
129286 | 42352.8108 | 1 | LI29 | 11 | 4 | HARDWARE | HW Problem | 29-1 tripping of every 5 minutes or so. Goes MOD fault and will reset. Pulled off the beam for now.
129286 | 42352.8108 | 1 | LI29 | 11 | 4 | HARDWARE | HW Problem | 29-1 tripping of every 5 minutes or so. Goes MOD fault and will reset. Pulled off the beam for now.
129286 | 42352.8108 | 1 | LI29 | 11 | 4 | HARDWARE | HW Problem | 29-1 tripping of every 5 minutes or so. Goes MOD fault and will reset. Pulled off the beam for now.
129887 | 42393.7996 | 1 | LI13 | 11 | 4 | HARDWARE | HW Problem | Klystron LI13-11 is tripped off on Mod Fault and will not reset remotely.
131180 | 42460.1265 |  | LI28 | 61 | 4 | HARDWARE | HW Problem | 28-6 trips frequently. Might be related to being on accelerate.
131436 | 42482.3746 |  | LI10 | 81 | 4 | HARDWARE | HW Problem | 10-8 1" pump out valve needs replacement
131480 | 42486.6029 | 1 | LI23 | 51 | 4 | HARDWARE | HW Problem | 23-5: Replace loose strap on lower PFN.
131481 | 42486.6046 | 1 | LI23 | 61 | 4 | HARDWARE | HW Problem | 23-6: Replace upper PFN strap.
129135 | 42345.5501 |  | LI12 | 31 | 4 | HARDWARE | HW Problem | 12-3 WGV fault, no remote recovery.
129524 | 42376.7205 | 1 | LI15 | 61 | 4 | HARDWARE | HW Problem | 15-6 has KMC fault
129689 | 42383.8011 |  | LI24 | 21 | 4 | HARDWARE | HW Problem | 24-2 showing phase jitter on FTP. This doesn't look real, don't see anything on AMPL or BVLT.
130026 | 42399.2252 |  | LI13 | 21 | 4 | HARDWARE | HW Problem | 13-2 MOD fault, See129861.
130064 | 42401.6808 | 3 | LI21 | 51 | 4 | HARDWARE | HW Problem | 21-5: PFN has unreasonable slope (-1.2), unable to fix-up timing.
130154 | 42405.4366 | 0.5 | LI21 | 21 | 4 | HARDWARE | HW Problem | L1-X MKT fault. Can't reset from ACR.
130497 | 42423.5161 | 0.3 | LI15 | 81 | 4 | HARDWARE | HW Problem | 15-8: RF AMPL low due to BV low. VVS low at 575.
130902 | 42445.6918 | 1 | LI30 | 71 | 4 | HARDWARE | HW Problem | can't turn 30-7 on remotely
130903 | 42445.6926 | 3 | LI24 | 61 | 4 | HARDWARE | HW Problem | Klystron 24-6 FLK fault and won't reset remotely
130907 | 42445.8649 | 1 | LI21 | 21 | 4 | HARDWARE | HW Problem | Klystron 21-3 MOD faulted and won't reset remotely
131101 | 42456.6975 | 1 | LI02 | 61 | 4 | HARDWARE | HW Problem | 2-6 WGV
131130 | 42457.6481 | 1 | LI28 | 61 | 4 | HARDWARE | HW Problem | Klystron LI28-61 is tripped off and does not reset remotely.
131276 | 42466.6511 | 1 | LI25 | 51 | 4 | HARDWARE | HW Problem | 25-5 modulator Beam Volts is very unstable, +/- 5kV pulse-to-pulse. This started at 14:45 on 4/6/16. Coincident with this, the DeQing current jumped up from 16 to 44 Amps. Klystron will not be needed for beam until Friday, so PEM should be able to take it to work on. Check with ACR.
131297 | 42469.2427 |  | LI26 | 31 | 4 | HARDWARE | HW Problem | Klys 26-3 tripped, no remote reset. Ops headed out to try and reset locally
129117 | 42344.3629 | 2 | LI21 | 71 | 4 | HARDWARE | HW Problem | 21-7 BVLT dropous
129151 | 42346.0788 | 1 | LI20 | 93 | 4 | HARDWARE | HW Problem | Kly. 20 - 3C tripped on WGV
129437 | 42374.6731 | 1 | LI29 | 71 | 4 | HARDWARE | HW Problem | 29-7: Perform BV & DQ calibration
129535 | 42377.2684 | 1 | LI05 | 41 | 4 | HARDWARE | HW Problem | 5-4 KMC flt
129599 | 42380.4781 | 0.25 | LI17 | 21 | 4 | HARDWARE | HW Problem | 17-2 MOD Fault, won't reset
129749 | 42388.407 | 0.25 | LI25 | 51 | 4 | HARDWARE | HW Problem | 25-5 has Ext Intlks and ACOC follows after. Need TS for problems.
130948 | 42447.5956 | 1 | LI11 | 81 | 4 | HARDWARE | HW Problem | 11-8 Loud arcing on turn on (eyeballed, looks like the Despiking Coil ) OBTW .
131217 | 42463.965 | 1 | LI20 | 93 | 4 | HARDWARE | HW Problem | 20-3C WGV. No remote reset
131301 | 42469.9071 |  | LI24 | 61 | 4 | HARDWARE | HW Problem | Klystron LI24-61 has frequent trigger faults and mod faults that self recover. Klystron is too unstable for regular use. Stanek note 4/19/16. Observed station in field with Accelerate triggers assigned. Every time Modulator lost triggers, it coincided with an MKSU 'Klys Focus Fault', which shows up as 'External Fault' on modulator PLC.
128968 | 42336.4943 | 0.75 | LI09 | 11 | 4 | HARDWARE | HW Problem | 9-1 MOD fault no remote reset.
129176 | 42347.5206 | 0.1 | LI29 | 81 | 4 | HARDWARE | HW Problem | KLYS:LI29:81:DL_WG_TEMP is 111.9 Deg. F
129372 | 42372.4123 |  | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3: Oil leak (slow and containable) under klystron.
130089 | 42403.3286 | 1 | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7 EOLCs TS/R
130316 | 42413.061 | 2.5 | LI30 | 41 | 4 | HARDWARE | HW Problem | station 30-4 Mod faulted
130361 | 42416.6139 | 0.75 | LI04 | 31 | 4 | HARDWARE | HW Problem | Klys 4-3 WGV with no reset, not recovering
130930 | 42446.7346 | 1 | LI23 | 21 | 4 | HARDWARE | HW Problem | station 23-2 WGV fault
130956 | 42448.47 | 1 | LI15 | 71 | 4 | HARDWARE | HW Problem | Kly Mod 15-7: mod faulted, no reset.
131053 | 42453.2902 | 1 | LI30 | 51 | 4 | HARDWARE | HW Problem | station 30-5 mod faults repeatedly
131053 | 42453.2902 | 1.5 | LI30 | 51 | 4 | HARDWARE | HW Problem | station 30-5 mod faults repeatedly
131053 | 42453.2902 |  | LI30 | 51 | 4 | HARDWARE | HW Problem | station 30-5 mod faults repeatedly
131053 | 42453.2902 | 1.4 | LI30 | 51 | 4 | HARDWARE | HW Problem | station 30-5 mod faults repeatedly
131084 | 42454.9961 | 1.6 | LI01 | 51 | 4 | HARDWARE | HW Problem | Klystron 1-5 mod faulted and won't reset remotely
131197 | 42460.718 |  | LI18 | 51 | 4 | HARDWARE | HW Problem | Kly. LI 18 - 5 show a Amplt. jitter 0.4%. ( see pic. in MCC logbook )
129090 | 42342.2217 | 1 | LI11 | 61 | 4 | HARDWARE | HW Problem | Kly. LI 11 - 6 is tripped on a "MOD" fault and won't reset.
129460 | 42375.0631 | 1 | LI28 | 51 | 4 | HARDWARE | HW Problem | 28-5 MOD faulted again. It's in local.
129512 | 42376.5295 |  | LI27 | 11 | 4 | HARDWARE | HW Problem | 27-1 amplitude is reading 0. The Beam Volts is reading a normal value and responds to off reset on commands but the amplitude always reads 0.
129614 | 42381.1912 | 1 | LI15 | 61 | 4 | HARDWARE | HW Problem | 15-6 The Plus5 and the Ref bot drift out of tols
129669 | 42382.9185 |  | LI09 | 21 | 4 | HARDWARE | HW Problem | Klys 9-2 PHAS is wandering.
129789 | 42388.9161 | 1 | LI23 | 41 | 4 | HARDWARE | HW Problem | 23-4 tripping often
129833 | 42390.5537 | 0.1 | LI12 | 21 | 4 | HARDWARE | HW Problem | station 12-2 has WGV fault
129847 | 42391.3928 |  | LI16 | 51 | 4 | HARDWARE | HW Problem | 16-5 BVLTs low, can't make up amplitude
129870 | 42391.952 | 1 | LI13 | 41 | 4 | HARDWARE | HW Problem | 13-4 T20 compartment fan not working
129877 | 42392.4066 | 1 | LI30 | 51 | 4 | HARDWARE | HW Problem | Klystron 30-5 MOD faulting every two or three minutes.
129878 | 42392.419 | 2 | LI18 | 81 | 4 | HARDWARE | HW Problem | Kly Mod 18-8: off with hard KLT fault. (Klystron delta temp.)
129878 | 42392.419 | 2 | LI18 | 81 | 4 | HARDWARE | HW Problem | Kly Mod 18-8: off with hard KLT fault. (Klystron delta temp.)
130970 | 42449.4403 | 0.3 | LI21 | 31 | 4 | HARDWARE | HW Problem | 21-3: Needs to replace snubber board (phase A). Measured 30-ohm.
131342 | 42474.3614 | 1 | LI27 | 81 | 4 | HARDWARE | HW Problem | 27-8 mod faulting constantly. Resets but trips again shortly thereafter. Many "ACOC" on the fault log.
131342 | 42474.3614 | 1 | LI27 | 81 | 4 | HARDWARE | HW Problem | 27-8 mod faulting constantly. Resets but trips again shortly thereafter. Many "ACOC" on the fault log.
130273 | 42410.6086 |  | LI21 | 11 | 4 | HARDWARE | HW Problem | L1S phase jitter is creeping up since 3:30 am, Feb 09. Was 0.03%, oscillating as high as 0.55% now and can be seen on the final beam energy. Physics log posts: http://physics-elog.slac.stanford.edu/lclselog/show.jsp?dir=/2016/06/10.02&pos=2016-02-10T14:13:26 http://physics-elog.slac.stanford.edu/lclselog/show.jsp?dir=/2016/06/10.02&pos=2016-02-10T14:21:24
130363 | 42416.6781 | 1.5 | LI28 | 31 | 4 | HARDWARE | HW Problem | Klystron 28-3 MOD faulted and won't reset remotely
130363 | 42416.6781 | 2 | LI28 | 31 | 4 | HARDWARE | HW Problem | Klystron 28-3 MOD faulted and won't reset remotely
130460 | 42421.3565 | 1 | LI29 | 11 | 4 | HARDWARE | HW Problem | 29-1 persistent heater delay
130511 | 42424.0323 | 1.5 | LI21 | 31 | 4 | HARDWARE | HW Problem | 21-3 tripping EOLC status PEM dispatched.
131211 | 42462.1708 | 1 | LI02 | 71 | 4 | HARDWARE | HW Problem | 2-7 WGV
131246 | 42465.4156 | 1 | LI21 | 11 | 4 | HARDWARE | HW Problem | 21-1 (L1-S) has mod faulted and will not reset remotely.
131334 | 42473.6181 | 0.5 | LI20 | 81 | 4 | HARDWARE | HW Problem | L0B has tripped four times in thirty minutes after recovering from 4/13/16 PAMM (with T20 replacement).
131335 | 42473.6309 | 1 | LI20 | 81 | 4 | HARDWARE | HW Problem | EPICS readout for RFW 20-8 L0B is flaky reports nonphysical vavle. Please see MCC ELOG entry: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=840181
131347 | 42474.5306 | 2.5 | LI23 | 81 | 4 | HARDWARE | HW Problem | Kly Mod 23-8: remove thyatron S/N 101058 from FACET kly mod 9-3 and install in modulator 23-8. Energize 23-8 and check for pulse stability.
131384 | 42478.586 | 1 | LI30 | 61 | 4 | HARDWARE | HW Problem | LI30-6 ACS #2 water faults
131400 | 42480.2811 |  | LI28 | 31 | 4 | HARDWARE | HW Problem | LI28-3 water faults last 24 hrs
131407 | 42480.5241 | 1 | LI29 | 31 | 4 | HARDWARE | HW Problem | 29-3 intermittent false "MKSU Trigger Enable Fault" status.
131417 | 42480.9603 | 4 | LI28 | 11 | 4 | HARDWARE | HW Problem | Klys 28-1 trips frequently, and now won't reset remotely. Ops headed out to try and reset locally. Update: core bias undercurrent fault
131439 | 42482.5213 |  | LI10 | 81 | 4 | HARDWARE | HW Problem | 10-8 klystron yoke needs replacement.
131440 | 42482.5234 |  | LI10 | 11 | 4 | HARDWARE | HW Problem | 10-1 station waveguide needs support bracket
131478 | 42486.5849 | 0.5 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 Mod EOLC stack needs to be replaced
131479 | 42486.5878 | 1.5 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 Mod HV filter cap in T20 compartment needs evaluated by equipment manager due to noticeable swelling of inside cap.
128925 | 42333.5371 | 1 | LI25 | 21 | 4 | HARDWARE | HW Problem | 25-2: Perform calibration and adjust DQing current
129276 | 42352.4557 | 1 | LI25 | 51 | 4 | HARDWARE | HW Problem | KLY. LI 28 - 5 MOD fault tripped and won't reset.
129459 | 42375.05 | 0.5 | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3 is MOD faulting frequently.
130916 | 42446.3186 |  | LI20 | 81 | 4 | HARDWARE | HW Problem | LCLS Inector Vault Water System - L0A (20-7), L0B (20-8), TCav. (20-5) panel: L0B FLOW = -2.35 and is red. ARCH. Strip chart shows the signal dropped to 7.0 at 3 -16 -16 : 06:03 and then dropped to -2.35 at 3 - 16 - 16 : 07:03 Note: changed low warning limit from 5.5 to -5.5 GPM and low alarm limit from 4.7 to -6 GPM to green up LCLS Home. Warning & alarm limits need to be restored once the problem is fixed.
131071 | 42454.4812 | 3 | LI24 | 31 | 4 | HARDWARE | HW Problem | Move the thyratron at Station 23-8 and install it at Station 24-3.
131105 | 42457.0202 |  | DR13 | 1 | 4 | HARDWARE | HW Problem | DR13 compressor phase has been making 30 degree drops for no known reason. See MccElog entry on 3/28 at 00:20 for more info.
131247 | 42465.4769 |  | LI30 | 71 | 4 | HARDWARE | HW Problem | 30-7: tripping on reflected energy followed by MOD ACOC.
131386 | 42478.59 |  | LI25 | 61 | 4 | HARDWARE | HW Problem | 25-6 is having frequent trip. Some of the faults that showed AMM, OV MKT. 132125 was closed.
129238 | 42349.8555 | 1 | LI22 | 61 | 4 | HARDWARE | HW Problem | 22-6 BV jitters Thyratron EOL but asked PEM to eek a little bit more out if they can.
129364 | 42372.2891 | 1 | LI18 | 11 | 4 | HARDWARE | HW Problem | 18-1 modulator klystron filament current meter face is broken off
129727 | 42385.7564 | 0.75 | LI18 | 61 | 4 | HARDWARE | HW Problem | Klystron LI18-61 is tripped on WG Vac and will not reset remotely. Klystron tripped on same fault 2 hrs. ago. No vacuum activity observed.
129828 | 42390.5432 | 0.75 | LI04 | 71 | 4 | HARDWARE | HW Problem | 4-7: RF AMPL low OOT due to beam volts low.
129873 | 42392.0804 | 1 | LI13 | 21 | 4 | HARDWARE | HW Problem | 13-2 MOD fault
129885 | 42392.9262 | 1 | LI09 | 11 | 4 | HARDWARE | HW Problem | 9-1 HV meter not reading
130024 | 42398.9059 | 0.75 | LI02 | 61 | 4 | HARDWARE | HW Problem | Klystron 2-6 WGV faulted
130232 | 42409.5271 |  | LI22 | 71 | 4 | HARDWARE | HW Problem | 22-7: RF AMPL low OOT. BV, SAT and timing look OK.
131128 | 42457.6086 |  | LI23 | 41 | 4 | HARDWARE | HW Problem | 23-4 Low uK.
131133 | 42458.1126 |  | LI00 | 99 | 4 | HARDWARE | HW Problem | LI00 Klys99 (a.k.a. SBUNCH) amplitude has been dropping out periodically. See MccElog entry on Mar 29 at 02:45 for more info.
131370 | 42477.684 |  | LI26 | 61 | 4 | HARDWARE | HW Problem | 26-6 is faulted on "Water Waveguide 1".
131370 | 42477.684 |  | LI26 | 61 | 4 | HARDWARE | HW Problem | 26-6 is faulted on "Water Waveguide 1".
129002 | 42338.5839 | 1 | LI27 | 81 | 4 | HARDWARE | HW Problem | 27-8: calibration, BV and DQing adjustment.
129478 | 42375.7195 | 0.75 | LI21 | 31 | 4 | HARDWARE | HW Problem | Klystron 21-3 MOD faulted and won't reset remotely
129523 | 42376.72 | 1.5 | LI15 | 21 | 4 | HARDWARE | HW Problem | 15-2 has KMC fault
129832 | 42390.5511 | 0.75 | LI14 | 51 | 4 | HARDWARE | HW Problem | 14-5: RF AMPL low OOT due to beam volts low.
131349 | 42474.5997 | 0.5 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 has mod faulted and will not reset remotely.
131443 | 42483.2888 | 1 | LI29 | 61 | 4 | HARDWARE | HW Problem | 29-6 trips about every twenty minutes.
131443 | 42483.2888 | 1 | LI29 | 61 | 4 | HARDWARE | HW Problem | 29-6 trips about every twenty minutes.
131443 | 42483.2888 | 3 | LI29 | 61 | 4 | HARDWARE | HW Problem | 29-6 trips about every twenty minutes.
131547 | 42493.0933 | 3 | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X amplitude is 0. EOIC diagnosed issue to TWT power supply problem.
131575 | 42494.76 | 0 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 has Modulator Interlock Lockout. Will not reset remotely.
131575 | 42494.76 | 1.5 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 has Modulator Interlock Lockout. Will not reset remotely.
128907 | 42332.8645 |  | LI25 | 21 | 4 | HARDWARE | HW Problem | 25-2 sled fault
128954 | 42335.6296 | 1 | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7 MOD fault, won't reset.
128957 | 42335.7869 |  | LI02 | 81 | 4 | HARDWARE | HW Problem | Kly. LI 02 - 8 keeps tripping and won't stay on. Has 40 AMPLT. trips on water
129109 | 42342.7862 | 0.75 | LI09 | 11 | 4 | HARDWARE | HW Problem | Klystron LI09-11 is tripped off and will not reset remotely.
129248 | 42350.8703 | 0.75 | LI17 | 51 | 4 | HARDWARE | HW Problem | 17-5 MOD fault no reset
129367 | 42372.2934 | 0.5 | LI28 | 71 | 4 | HARDWARE | HW Problem | 28-7 SOLA PS2 24VDC power supply has no output
129584 | 42380.2638 |  | LI00 | 12 | 4 | HARDWARE | HW Problem | Station 00-1B has low amplitude. Right now it is about 33 where it was ~65 prior to shutdown.
131456 | 42485.4551 | 2 | LI24 | 61 | 4 | HARDWARE | HW Problem | Kly. LI 24 - 6 faulting on MKSU faults and tripping off on "MOD" fault.
129052 | 42340.6001 | 0.2 | LI28 | 81 | 4 | HARDWARE | HW Problem | 28-8 has mod faulted, is currently HVR and will not reset remotely.
129305 | 42353.6294 | 2 | LI09 | 11 | 4 | HARDWARE | HW Problem | Kly Mod 9-1 (Feedback): CB 7 tripped.
129366 | 42372.2922 | 0.2 | LI26 | 51 | 4 | HARDWARE | HW Problem | 26-5 AIP modulator PLC CPU has no stored values, may need to be re-programmed
129683 | 42383.574 | 1.5 | LI22 | 21 | 4 | HARDWARE | HW Problem | 22-2 needs to check on klixon components to make sure that they are still OK.
129722 | 42385.185 | 0.5 | LI09 | 11 | 4 | HARDWARE | HW Problem | Kly. LI 09 - 1 tripped on "MOD" fault and won't reset
129924 | 42395.4117 | 1 | LI25 | 21 | 4 | HARDWARE | HW Problem | Modulator 25-2 has frequent large Beam Voltage spikes. Thyratron reservoir at 6.9 Volts.
129929 | 42395.5001 | 1.5 | LI28 | 71 | 4 | HARDWARE | HW Problem | KYS:LI28:7 tripping off. Error log shows it's tripped off 25 times in the last 4 hours.
130909 | 42445.9236 | 1 | LI24 | 41 | 4 | HARDWARE | HW Problem | Klystron 24-4 mod faulted and won't reset remotely
131012 | 42451.4926 | 2 | LI22 | 41 | 4 | HARDWARE | HW Problem | 22-4: RF AMPL drop-out spikes with no evidence of PHJ, BVJ or BCJ.
131014 | 42451.5213 |  | LI21 | 31 | 4 | HARDWARE | HW Problem | 21-3 has tripped three times this shift. Eventually resets remotely, but only after cycling MKSU, Klystron over-voltage, and MOD fault lockout a few times.
131179 | 42459.788 | 4 | LI00 | 1 | 4 | HARDWARE | HW Problem | TBR - K01 MOD fault lockout. Klystron IKOCs with vacuum bursts.
131179 | 42459.788 | 1.5 | LI00 | 1 | 4 | HARDWARE | HW Problem | TBR - K01 MOD fault lockout. Klystron IKOCs with vacuum bursts.
131352 | 42475.2564 |  | LI23 | 81 | 4 | HARDWARE | HW Problem | 23-8 WGV fault is from the plunger switch. Manually actuated it to clear but it would soon fault again.
131451 | 42485.4183 |  | LI20 | 61 | 4 | HARDWARE | HW Problem | 20-6 (Gun): Intermittent MKOV/MKOC trips coinciding with HVPS excursions.
131523 | 42490.8879 |  | LI21 | 81 | 4 | HARDWARE | HW Problem | Klys 21-8 and 22-8 getting BVLT dropouts on the same pulse. Would CATER each individual klystron, but the problem is common to both of them as their BVLT dropouts are always on the same pulse. Makes it seem less like a thyratron problem, but unclear. This is too weird, not sure what could cause that. MOD 5/1/16 HVS this is a LLRF timing problem with LI22 FIDO. LI22 FIDO RF levels need to be checked. Timing of PDUs which use this FIDO are jumping ~500nS w.r.t. the rest of the timing system, causing these apparent" dropouts.
128935 | 42333.9238 | 1 | LI01 | 61 | 4 | HARDWARE | HW Problem | 1-6 mod fault
128981 | 42337.4909 | 1.5 | LI09 | 11 | 4 | HARDWARE | HW Problem | 9-1 MOD fault no reset
128984 | 42337.9341 | 1.5 | LI02 | 71 | 4 | HARDWARE | HW Problem | Klystron 2-7 is in extended heater delay. It went into KHD over an hour ago for no clear reason and still can't reset.
129288 | 42352.8454 | 1 | LI19 | 41 | 4 | HARDWARE | HW Problem | Klystron 19-4 mod faulted and won't reset remotely
129481 | 42375.7241 | 0.25 | LI29 | 31 | 4 | HARDWARE | HW Problem | Klystron 29-3 MOD faulted but cannot reset remotely since it is in LOCAL
129483 | 42375.7278 | 0.5 | LI20 | 71 | 4 | HARDWARE | HW Problem | 20-7 MOD fault no reset.
129521 | 42376.7189 | 1.5 | LI02 | 81 | 4 | HARDWARE | HW Problem | 2-8 has KMC fault
129562 | 42377.9128 | 1 | LI21 | 51 | 4 | HARDWARE | HW Problem | Kly. LI 21 - 5 tripping off. Faults are "MOD EVO".
130911 | 42445.9489 |  | LI24 | 61 | 4 | HARDWARE | HW Problem | time accounting (-3H) 24-6 tripped several times MKT. PEM to remove ferrite cores that were inserted today. Not sure what this will do.
130941 | 42447.4679 |  | LI15 | 71 | 4 | HARDWARE | HW Problem | 15-7 is not getting accelerate triggers. This behavior follows a PIOP swap that was executed to clear up an "IPL" fault.
131199 | 42461.4384 | 1 | LI02 | 71 | 4 | HARDWARE | HW Problem | 2-7 WGV, probably needs gauge to be juiced.
131259 | 42466.2723 | 1 | LI21 | 71 | 4 | HARDWARE | HW Problem | LI21-7 showing water faults last 24 hours. ACR says offline, will try to check later this morning
131325 | 42473.2702 | 0.8 | LI27 | 81 | 4 | HARDWARE | HW Problem | Klys 27-8 started tripping frequently around 05:00 on 4/13. Left off the beam.
131326 | 42473.2707 | 1 | LI27 | 51 | 4 | HARDWARE | HW Problem | Klys 27-5 went FLK. No reset.
131474 | 42486.5782 | 0.5 | LI30 | 71 | 4 | HARDWARE | HW Problem | 30-7 PFN strap burnt ( upper, LHS, # 1 )
129023 | 42339.5353 | 0.1 | LI20 | 81 | 4 | HARDWARE | HW Problem | 20-8 (aka L0B): FMON high OOT. Was stepped up from 14 to 15 amps on 11/25 at ~9:00. Who, when, where & why?
129118 | 42344.3731 | 2 | LI28 | 31 | 4 | HARDWARE | HW Problem | Kly. LI 28 - 3 keeps tripping off.
129233 | 42349.57 | 0.5 | LI18 | 21 | 4 | HARDWARE | HW Problem | 18-2 tripped on klystron filament undercurrent
129233 | 42349.57 | 2 | LI18 | 21 | 4 | HARDWARE | HW Problem | 18-2 tripped on klystron filament undercurrent
129247 | 42350.7489 | 4 | LI21 | 11 | 4 | HARDWARE | HW Problem | 21-1 (aka L1S): Per FJD widen PFN pulse for a "good" flat-top of 1.9 uS.
129342 | 42356.3262 |  | LI19 | 41 | 4 | HARDWARE | HW Problem | 19-4 MOD fault.
129528 | 42376.9291 | 2 | LI00 | 98 | 4 | HARDWARE | HW Problem | LI00 SHB #2 shows "MOD" fault and zero AMPLIT. Won't reset.
129573 | 42379.4196 | 0.3 | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3: wENT hvr, NO REMOTE RESET.
131067 | 42454.4056 | 2 | LI15 | 41 | 4 | HARDWARE | HW Problem | 15-4: Check the WG VGC with the calibration load.
131076 | 42454.5572 | 0.7 | LI01 | 51 | 4 | HARDWARE | HW Problem | Klystron LI01-51 is tripped off and will not reset remotely. Klystron shows three voltage faults on summary.
131076 | 42454.5572 | 2 | LI01 | 51 | 4 | HARDWARE | HW Problem | Klystron LI01-51 is tripped off and will not reset remotely. Klystron shows three voltage faults on summary.
131395 | 42479.482 |  | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3: Filament button not working on vacuum gauge.
131502 | 42487.9362 | 0.5 | LI28 | 31 | 4 | HARDWARE | HW Problem | station 28-3 Mod faulted
131526 | 42491.7279 | 1 | LI23 | 81 | 4 | HARDWARE | HW Problem | Klys 23-8 FLK, no reset.
131566 | 42493.9558 |  | LI19 | 81 | 4 | HARDWARE | HW Problem | 19-8 Klystron leaks water from the drain hole.
129726 | 42385.6722 | 0.75 | LI18 | 61 | 4 | HARDWARE | HW Problem | Klystron LI18-61 is faulted on WG Vac and will not reset remotely.
129830 | 42390.5486 | 0.75 | LI13 | 41 | 4 | HARDWARE | HW Problem | 13-4: RF AMPL low OOT due to beam volts low.
131138 | 42458.2894 |  | LI12 | 31 | 4 | HARDWARE | HW Problem | Lots of water summary faults from 12-3.
129010 | 42338.8217 | 1 | LI20 | 81 | 4 | HARDWARE | HW Problem | L0B WGV fault, won't reset locally. Station gauge is faulted.
129013 | 42339.0823 | 1 | LI30 | 51 | 4 | HARDWARE | HW Problem | 30-5 trips repeatedly on MKT.
129013 | 42339.0823 |  | LI30 | 51 | 4 | HARDWARE | HW Problem | 30-5 trips repeatedly on MKT.
129056 | 42340.7806 | 0.5 | LI28 | 21 | 4 | HARDWARE | HW Problem | Klystron 28-2 mod faulted and won't reset remotely.
129070 | 42341.2201 | 1 | LI20 | 81 | 4 | HARDWARE | HW Problem | L0B WGV fault.
129232 | 42349.5158 | 0.15 | LI25 | 21 | 4 | HARDWARE | HW Problem | Kly. LI 25 - 2 tripped on WGV. Won't reset.
129255 | 42351.2907 | 8 | LI11 | 81 | 4 | HARDWARE | HW Problem | 11-8 MOD fault. At the station found the modulator with a blue lock and the tag said that there is a burning smell coming from it. Did a visual and did not see nothing obvious . Safer to have to people with HV on .
129312 | 42354.271 | 2 | LI23 | 61 | 4 | HARDWARE | HW Problem | LI23-6 water summary faults overnight
129566 | 42378.366 | 1 | LI28 | 31 | 4 | HARDWARE | HW Problem | Klys 28-3 HVR in local.
129657 | 42382.5531 | 2 | LI13 | 61 | 4 | HARDWARE | HW Problem | 13-6: Core bias fault - meter reads 0A.
130925 | 42446.5672 | 2 | LI30 | 81 | 4 | HARDWARE | HW Problem | 30-8 MOD fault, no remote reset.
131198 | 42461.3088 | 2 | LI15 | 61 | 4 | HARDWARE | HW Problem | 15-6 WGV
131258 | 42466.1196 | 1 | LI24 | 51 | 4 | HARDWARE | HW Problem | 24-5 has high phase and amplitude jitter.
131258 | 42466.1196 |  | LI24 | 51 | 4 | HARDWARE | HW Problem | 24-5 has high phase and amplitude jitter.
131290 | 42468.1609 | 1.5 | LI26 | 31 | 4 | HARDWARE | HW Problem | Klys 26-3 tripped, no remote reset
131486 | 42486.7023 | 1 | LI21 | 61 | 4 | HARDWARE | HW Problem | Klys Focus supply readbacks for 21-6 have purple (disconnected) status for Resistance, Voltage, & Current Readback. Klystron focus as read by PIOP is fine (14.97 A) & station is on & running.
131576 | 42494.8873 | 0.75 | LI23 | 51 | 4 | HARDWARE | HW Problem | 23-5 HVR, doesn't come up at remote command.
131587 | 42495.6083 | 2 | LI23 | 31 | 4 | HARDWARE | HW Problem | Klystron LI23-31 is tripped off and will not reset remotely.
131590 | 42495.7508 | 1 | LI24 | 21 | 4 | HARDWARE | HW Problem | 24-2 MOD fault, doesn't reset.
128908 | 42332.8655 | 3 | LI27 | 11 | 4 | HARDWARE | HW Problem | 27-1 thyratron heater frequently faulting on over voltage
128908 | 42332.8655 | 1 | LI27 | 11 | 4 | HARDWARE | HW Problem | 27-1 thyratron heater frequently faulting on over voltage
129114 | 42343.5308 | 1.5 | LI20 | 61 | 4 | HARDWARE | HW Problem | 20-6 LCLS gun beam volts jitter, clearly visible on BC1.
129440 | 42374.6975 | 1 | LI20 |  | 4 | HARDWARE | HW Problem | After PPS trip, LCLS injector vault stations will not turn on 20-5 20-6 20-7 20-8. PEM dispatched to reset this condition
129479 | 42375.7199 | 0.25 | LI28 | 61 | 4 | HARDWARE | HW Problem | Klystron 28-6 MOD faulted and won't reset remotely
129673 | 42383.0508 | 1 | LI18 | 51 | 4 | HARDWARE | HW Problem | Kly. LI 18 - 5 has a bad "PAD = false".
130953 | 42447.8942 | 0.75 | LI03 | 31 | 4 | HARDWARE | HW Problem | station 3-7 mod faulted
131025 | 42451.8054 | 0.75 | LI13 | 11 | 4 | HARDWARE | HW Problem | Klys 13-1 is MOD faulted, wont reset.
131098 | 42456.5641 | 1 | LI28 | 11 | 4 | HARDWARE | HW Problem | Station 28-1 HVR. Resets but trips off immediately after. Continuously faulting.
131251 | 42465.6251 | 2 | LI21 | 71 | 4 | HARDWARE | HW Problem | Kly Mod 21-7: HV readback on touch screen approx. 1/4 of set point. Actual BV matches set point of 21.6KV. Either readback board (attached to DC filter caps above T20) is bad or PLC requires reprogramming. comment by p.stiles: Beings the BV matches the setpoint of 21.6 kV, the Enerpro board is OK. Also, the part of the HV regulator board involved in regulating the 21.6 kV is also OK. The problem is in the part of the HV regulator board associated with sending a DC signal to the PLC analog in module OR a wiring issue (not so likely) OR a faulty PLC analog input (AI2_CH1). The HV divider (connected to HV regulator board) divides the DC HV by about 2527:1 so the voltage at TP25 on the HV reg. board should be about 8.55 volts. This is buffered and filtered by U15A and goes out, via RN2, to PLC AI2_CH1. See e-log entry 845040 for some troubleshooting measurements and ideas.
131251 | 42465.6251 | 3 | LI21 | 71 | 4 | HARDWARE | HW Problem | Kly Mod 21-7: HV readback on touch screen approx. 1/4 of set point. Actual BV matches set point of 21.6KV. Either readback board (attached to DC filter caps above T20) is bad or PLC requires reprogramming. comment by p.stiles: Beings the BV matches the setpoint of 21.6 kV, the Enerpro board is OK. Also, the part of the HV regulator board involved in regulating the 21.6 kV is also OK. The problem is in the part of the HV regulator board associated with sending a DC signal to the PLC analog in module OR a wiring issue (not so likely) OR a faulty PLC analog input (AI2_CH1). The HV divider (connected to HV regulator board) divides the DC HV by about 2527:1 so the voltage at TP25 on the HV reg. board should be about 8.55 volts. This is buffered and filtered by U15A and goes out, via RN2, to PLC AI2_CH1. See e-log entry 845040 for some troubleshooting measurements and ideas.
131251 | 42465.6251 |  | LI21 | 71 | 4 | HARDWARE | HW Problem | Kly Mod 21-7: HV readback on touch screen approx. 1/4 of set point. Actual BV matches set point of 21.6KV. Either readback board (attached to DC filter caps above T20) is bad or PLC requires reprogramming. comment by p.stiles: Beings the BV matches the setpoint of 21.6 kV, the Enerpro board is OK. Also, the part of the HV regulator board involved in regulating the 21.6 kV is also OK. The problem is in the part of the HV regulator board associated with sending a DC signal to the PLC analog in module OR a wiring issue (not so likely) OR a faulty PLC analog input (AI2_CH1). The HV divider (connected to HV regulator board) divides the DC HV by about 2527:1 so the voltage at TP25 on the HV reg. board should be about 8.55 volts. This is buffered and filtered by U15A and goes out, via RN2, to PLC AI2_CH1. See e-log entry 845040 for some troubleshooting measurements and ideas.
131251 | 42465.6251 | 2 | LI21 | 71 | 4 | HARDWARE | HW Problem | Kly Mod 21-7: HV readback on touch screen approx. 1/4 of set point. Actual BV matches set point of 21.6KV. Either readback board (attached to DC filter caps above T20) is bad or PLC requires reprogramming. comment by p.stiles: Beings the BV matches the setpoint of 21.6 kV, the Enerpro board is OK. Also, the part of the HV regulator board involved in regulating the 21.6 kV is also OK. The problem is in the part of the HV regulator board associated with sending a DC signal to the PLC analog in module OR a wiring issue (not so likely) OR a faulty PLC analog input (AI2_CH1). The HV divider (connected to HV regulator board) divides the DC HV by about 2527:1 so the voltage at TP25 on the HV reg. board should be about 8.55 volts. This is buffered and filtered by U15A and goes out, via RN2, to PLC AI2_CH1. See e-log entry 845040 for some troubleshooting measurements and ideas.
130094 | 42403.4265 | 0.5 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 BV was lower due to thyratron aging. Needs to adjust to spec after thyratron replacement.
130123 | 42404.3062 |  | LI20 | 71 | 4 | HARDWARE | HW Problem | 20-7 Needs PFN tunning for Jitter and BV drops. If problem remains after tuning, recommend Thryratron change
130126 | 42404.3479 | 0.1 | LI04 | 51 | 4 | HARDWARE | HW Problem | Kly. LI 04 - 5 PAD DAC "false" status.
129054 | 42340.7087 | 0 | LI22 | 61 | 4 | HARDWARE | HW Problem | Klys 22-6 showing some beam volts jitter.
129064 | 42341.1448 | 0.5 | LI06 | 81 | 4 | HARDWARE | HW Problem | 6-8 phase readback is bad.
129068 | 42341.1822 | 1 | LI04 | 51 | 4 | HARDWARE | HW Problem | 4-5 bad phase jitter.
129257 | 42351.9384 | 1 | LI22 | 21 | 4 | HARDWARE | HW Problem | 22-2 MOD fault no reset
129844 | 42391.2339 | 1 | LI04 | 11 | 4 | HARDWARE | HW Problem | 4-1 MOD fault, no remote reset.
129989 | 42397.9363 | 1 | LI22 | 51 | 4 | HARDWARE | HW Problem | 22-5 is showing Waveguide vacuum
130040 | 42401.0991 | 0.5 | LI17 | 71 | 4 | HARDWARE | HW Problem | Klystron LI17-71 is in FLK and will not reset remotely.
130095 | 42403.4454 | 1 | LI24 | 41 | 4 | HARDWARE | HW Problem | 24-4 Perform BV & DQ calibration
130097 | 42403.4467 | 0.5 | LI27 | 41 | 4 | HARDWARE | HW Problem | 27-4 Perform BV & DQ calibration
130109 | 42403.71 | 2.5 | LI30 | 51 | 4 | HARDWARE | HW Problem | 30-5 FLK and won't reset remotely.
130112 | 42403.8149 | 2 | LI24 | 61 | 4 | HARDWARE | HW Problem | Klystron 24-6 cycling constantly
130134 | 42404.4414 | 0.5 | LI05 | 81 | 4 | HARDWARE | HW Problem | 5-8 KLV fault.
129362 | 42372.2866 | 1 | LI00 | 1 | 4 | HARDWARE | HW Problem | MOD 0-01 Fire Alarm Ready green LED is out
129458 | 42375.0444 |  | LI20 | 51 | 4 | HARDWARE | HW Problem | TCAV0 Has low amplitude. Too low for LLRF feedbacks to regulate.
129534 | 42377.2616 | 1 | LI02 | 11 | 4 | HARDWARE | HW Problem | 2-1 CBL flt.
129586 | 42380.29 | 1 | LI15 | 21 | 4 | HARDWARE | HW Problem | 15-2 high voltage meter is inperative
130096 | 42403.4461 | 0.5 | LI27 | 31 | 4 | HARDWARE | HW Problem | 27-3 Perform BV & DQ calibration
130114 | 42403.8785 | 2 | LI25 | 81 | 4 | HARDWARE | HW Problem | Klys 25-8 has been cycling frequently.
130125 | 42404.3457 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X has extremely high phase jitter - so much so that it is clearly reflected in the shot-to-shot FEL pulse energy.
130125 | 42404.3457 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X has extremely high phase jitter - so much so that it is clearly reflected in the shot-to-shot FEL pulse energy.
130179 | 42405.9858 | 0.25 | LI15 | 21 | 4 | HARDWARE | HW Problem | Klys 15-2 has BVLT jitter
130260 | 42410.3685 | 4 | LI28 | 11 | 4 | HARDWARE | HW Problem | 28-1 need pulse cable maintenance
130299 | 42411.7521 | 0.5 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 is MOD faulted.
130371 | 42416.9132 | 0.5 | LI20 | 94 | 4 | HARDWARE | HW Problem | Kly. LI 20 - 4A faulted on WGV. Not resetting.
128914 | 42333.1447 | 0.5 | LI29 | 61 | 4 | HARDWARE | HW Problem | 29-6 MOD fault, no reset (in local)
129558 | 42377.7134 | 0.5 | LI25 | 51 | 4 | HARDWARE | HW Problem | Klys 25-5 keeps going HVR. Resets each time, but has happened ~5 times in the past hour
129558 | 42377.7134 | 1 | LI25 | 51 | 4 | HARDWARE | HW Problem | Klys 25-5 keeps going HVR. Resets each time, but has happened ~5 times in the past hour
129564 | 42378.0547 | 1 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 tripped off HVR, is in local control needs local reset.
129594 | 42380.421 | 4 | LI21 | 61 | 4 | HARDWARE | HW Problem | 21-6 trips immediately when turned on.
129594 | 42380.421 | 4 | LI21 | 61 | 4 | HARDWARE | HW Problem | 21-6 trips immediately when turned on.
129594 | 42380.421 | 2 | LI21 | 61 | 4 | HARDWARE | HW Problem | 21-6 trips immediately when turned on.
129594 | 42380.421 | 3 | LI21 | 61 | 4 | HARDWARE | HW Problem | 21-6 trips immediately when turned on.
129685 | 42383.6286 | 0.25 | LI18 | 51 | 4 | HARDWARE | HW Problem | station 18-5 PAD not reading back
129869 | 42391.9332 | 1 | LI17 | 71 | 4 | HARDWARE | HW Problem | Klystron Li17-71 is tripped off and will not reset remotely.
130018 | 42398.7104 | 0.75 | LI30 | 51 | 4 | HARDWARE | HW Problem | 30-5 tripped excessively in a short time frame and now is it in Fault Lockout.
130176 | 42405.9113 | 1 | LI11 | 71 | 4 | HARDWARE | HW Problem | 11-7 T20 compartment fan not working
130178 | 42405.9829 |  | LI02 | 61 | 4 | HARDWARE | HW Problem | Klys 2-6 has lots of AMPL jitter. BVLTs seem ok.
130184 | 42406.3467 | 4 | LI11 | 71 | 4 | HARDWARE | HW Problem | 11-6 has a false KMC on PIOP
130184 | 42406.3467 | 2 | LI11 | 71 | 4 | HARDWARE | HW Problem | 11-6 has a false KMC on PIOP
130187 | 42406.8337 | 1 | LI25 | 11 | 4 | HARDWARE | HW Problem | Klys 25-1 tripped, no reset.
130278 | 42410.8004 | 1 | LI09 | 41 | 4 | HARDWARE | HW Problem | 9-4 has EOLC fault
130337 | 42415.7711 | 1 | LI15 | 41 | 4 | HARDWARE | HW Problem | Klystron 15-4 WGV faulted
129509 | 42376.4858 | 3 | LI22 | 21 | 4 | HARDWARE | HW Problem | Kly Mod 22-2: replaced burned CB 7 breaker.
130486 | 42423.4009 | 2 | LI21 | 61 | 4 | HARDWARE | HW Problem | 21-6 has PAD phase error and PAD +/- 15 V errors. May be able to work on this during operation if Ops can deactivate it from beam. Call ACR and ask.
130549 | 42425.9022 | 0.5 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 MOD fault. Station in local, can't reset.
130577 | 42427.3049 | 1 | LI28 | 51 | 4 | HARDWARE | HW Problem | 28-5 is in fault lock out and will not reset remotely.
130581 | 42427.4795 | 1 | LI02 | 71 | 4 | HARDWARE | HW Problem | Kly. LI 02 - 7 tripped on WGV and it won't reset.
130582 | 42427.4949 |  | LI15 | 31 | 4 | HARDWARE | HW Problem | Klystron LI15-31 is tripped off and will not reset remotely.
130658 | 42431.6058 | 1 | LI20 | 71 | 4 | HARDWARE | HW Problem | 20-7 Mod has a old type HV rect stack that needs to be replaced with a metal type that is less prone to catching fire.
130750 | 42436.6621 | 1.5 | LI20 | 93 | 4 | HARDWARE | HW Problem | Kly. LI20 - 3A tripped off. won't reset. VP400/401 is tripped.
130760 | 42437.4115 | 1 | LI25 | 21 | 4 | HARDWARE | HW Problem | 25-2: HV regulator board has resistor R111=2K, it should be 100K.
130818 | 42440.4519 | 1 | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8: Seeing some beam volts jitter.
129163 | 42346.6993 | 2 | LI27 | 81 | 4 | HARDWARE | HW Problem | Klystron 27-8 MOD faults constantly. We are able to reset it successfully, but it trips almost immediately afterwards.
129163 | 42346.6993 | 1.5 | LI27 | 81 | 4 | HARDWARE | HW Problem | Klystron 27-8 MOD faults constantly. We are able to reset it successfully, but it trips almost immediately afterwards.
129227 | 42349.4888 |  | LI09 | 11 | 4 | HARDWARE | HW Problem | 9-1 has mod faulted and will not reset remotely.
129769 | 42388.4831 | 2 | LI28 | 61 | 4 | HARDWARE | HW Problem | 28-6 tripped off, doesn't reset remotely.
129813 | 42390.0184 | 1 | LI20 | 81 | 4 | HARDWARE | HW Problem | L0B (20-8) PAD RF Amplitude reading 4.81. Expected value 45.0.
130201 | 42408.254 |  | LI09 | 41 | 4 | HARDWARE | HW Problem | Klystron 9-4 has unacceptable jitter problems.
130209 | 42408.5284 | 0 | LI29 | 31 | 4 | HARDWARE | HW Problem | 29-3 repeated MKSU trigger enable faults
130262 | 42410.3996 |  | LI20 | 51 | 4 | HARDWARE | HW Problem | 20-5/TCAV0 water waveguide #1 fault. UPdate 2/12/16: Re-opened CATER. Flow switch faulted again. -jwarren
130779 | 42438.076 | 4 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3 tripped doesn't reset.
130098 | 42403.4734 | 2 | LI28 | 11 | 4 | HARDWARE | HW Problem | 28-1 tripping HVR frequently. Resets but trips again within a minute
130174 | 42405.873 | 1 | LI04 | 51 | 4 | HARDWARE | HW Problem | Klys 4-5 WGV
130181 | 42405.9877 |  | LI15 | 41 | 4 | HARDWARE | HW Problem | Klys 15-4 has BVLT jitter.
130969 | 42449.4164 | 1 | LI25 | 41 | 4 | HARDWARE | HW Problem | 25-4 FTPs show extremely large phase two-state, suspect this is a PAD issue of some sort.
131050 | 42453.0251 |  | LI30 | 51 | 4 | HARDWARE | HW Problem | station 30-5 mod faulted repeatedly. Duplicate cater 131053.
131070 | 42454.4646 | 0.5 | DR13 | 1 | 4 | HARDWARE | HW Problem | DR13 Klystron 1 (NRTL compressor) has tripped off and will not reset remotely.
131073 | 42454.5093 | 0.5 | LI26 | 11 | 4 | HARDWARE | HW Problem | 26-1: RF AMPL low OOT due to beam volts very low.
131096 | 42456.3706 | 2 | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 HVR with no remote reset. Station is in local mode.
128944 | 42334.7838 | 1 | LI17 | 71 | 4 | HARDWARE | HW Problem | Klystron 17-7 fault lockout and won't reset remotely
129843 | 42391.0073 | 2 | LI15 | 31 | 4 | HARDWARE | HW Problem | 15-3 MKT
130079 | 42402.4777 | 1 | LI02 | 11 | 4 | HARDWARE | HW Problem | Kly. LI 02 - 1 went out of phase. When checked, noted that the Focus Monitor was reading 17.0 and was expected to read 14.8. Informed AMRF.
130832 | 42441.2595 | 1 | LI12 | 81 | 4 | HARDWARE | HW Problem | 12-8 WGV flt.
130846 | 42443.1565 | 2 | LI21 | 71 | 4 | HARDWARE | HW Problem | 21-7 trips repeatedly. TBR- kly at end of life.
130846 | 42443.1565 | 2 | LI21 | 71 | 4 | HARDWARE | HW Problem | 21-7 trips repeatedly. TBR- kly at end of life.
130846 | 42443.1565 | 2 | LI21 | 71 | 4 | HARDWARE | HW Problem | 21-7 trips repeatedly. TBR- kly at end of life.
130992 | 42450.5951 | 0.1 | LI30 | 51 | 4 | HARDWARE | HW Problem | Klystron LI30-51 is tripped off on WGV and can't be reset remotely.
129516 | 42376.5915 |  | LI04 | 41 | 4 | HARDWARE | HW Problem | Kly Mod 4-4: requires T20 replacement. Appears the secondary bushing closest to T20 compartment door has burned oil on it. Transformer is burning oil and probably has at least one phase shorted. Station trips CB 7 after short run time and emits strong burned oil smell.
129616 | 42381.2109 | 1 | LI11 | 51 | 4 | HARDWARE | HW Problem | 11-5 T20 compartment fan is dead
129706 | 42384.5792 | 0.5 | LI25 | 21 | 4 | HARDWARE | HW Problem | 25-2: Needs BV and DQ adjustment
130211 | 42408.5572 | 1 | LI03 | 61 | 4 | HARDWARE | HW Problem | 3-6: Seeing some pretty bad beam volts jitter.
130313 | 42412.6638 | 2 | LI12 | 11 | 4 | HARDWARE | HW Problem | Klystron LI12-11 is tripped off and will not reset remotely.
131145 | 42458.3951 |  | LI03 | 21 | 4 | HARDWARE | HW Problem | RFS 3-2 numerous RE trips. EOIC lowered drive. (DT OS 3/29/16 -0.1 hr)
131305 | 42470.3667 | 0.75 | LI28 | 51 | 4 | HARDWARE | HW Problem | 28-5 tripped; can't reset remotely.
130609 | 42430.2701 |  | LI05 | 71 | 4 | HARDWARE | HW Problem | DUPLICATE of 130507 - 5-7 CABLE fault, swapping PIOP cables with 5-8, problem moves to 5-8, looks like the PIOP for 5-7 is bad
130611 | 42430.2843 |  | LI20 | 71 | 4 | HARDWARE | HW Problem | Kly. LI 20 - 7 is cycling. Looks like like the Amplitude is dropping out.
130393 | 42417.7984 | 2 | LI29 | 21 | 4 | HARDWARE | HW Problem | Klystron 29-2 MOD faulting frequently. We are able to reset, but it immediately trips off afterwards.
131042 | 42452.576 | 0.5 | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2: AC curr at the TP's "setpoint and readback" displays =30.3A. It should be double.
131048 | 42452.6853 | 1 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 mod fault, can't reset remotely.
131078 | 42454.65 |  | LI24 | 31 | 4 | HARDWARE | HW Problem | Modulator 24-3: HV current reading problem: Reading is 2.7x amps, should be about 3.7x amps.
131091 | 42455.7595 | 1.5 | LI25 | 81 | 4 | HARDWARE | HW Problem | Kly. LI25 - 8 tripped. Won't reset.
129011 | 42338.8289 | 2 | LI11 | 61 | 4 | HARDWARE | HW Problem | Klystron 11-6 mod faulted and won't reset remotely
129555 | 42377.5541 | 1 | LI26 | 61 | 4 | HARDWARE | HW Problem | Klys 26-6 KMC fault.
129577 | 42379.8274 | 1 | DR13 | 1 | 4 | HARDWARE | HW Problem | The north compressor modulator (KLYS:DR13:1) has a waveguide vacuum fault
129734 | 42387.3854 | 1.5 | LI24 | 11 | 4 | HARDWARE | HW Problem | 24-1 trips immediately when turned on.
129734 | 42387.3854 | 0.8 | LI24 | 11 | 4 | HARDWARE | HW Problem | 24-1 trips immediately when turned on.
129734 | 42387.3854 | 5 | LI24 | 11 | 4 | HARDWARE | HW Problem | 24-1 trips immediately when turned on.
129734 | 42387.3854 | 1.5 | LI24 | 11 | 4 | HARDWARE | HW Problem | 24-1 trips immediately when turned on.
129734 | 42387.3854 | 1 | LI24 | 11 | 4 | HARDWARE | HW Problem | 24-1 trips immediately when turned on.
129874 | 42392.1587 | 1 | LI17 | 71 | 4 | HARDWARE | HW Problem | 17-7 MOD fault, no remote reset.
130059 | 42401.5649 |  | LI24 | 21 | 4 | HARDWARE | HW Problem | 24-2 Klystron forward power (FEPK) drooped to 2MW.
130180 | 42405.9867 | 0.25 | LI15 | 31 | 4 | HARDWARE | HW Problem | Klys 15-3 has BVLT jitter.
130183 | 42406.3219 | 1 | LI02 | 61 | 4 | HARDWARE | HW Problem | 2-6 WGV fault. Vacuum looks fine.
131579 | 42495.2987 | 2 | LI30 | 51 | 4 | HARDWARE | HW Problem | Beam volts jitter on 30-5
131585 | 42495.5529 | 1 | LI28 | 61 | 4 | HARDWARE | HW Problem | 28-6 tripping frequently. This has been the case in the past (and seems to be related to being on accelerate).
128930 | 42333.6615 | 0.75 | LI29 | 51 | 4 | HARDWARE | HW Problem | 29-5: Perform calibration and adjust BV, DQing current to spec
128972 | 42336.7417 |  | LI20 | 61 | 4 | HARDWARE | HW Problem | Gun RF peak power trip, injector vacuum valves went in. -0.1 hr DT SS 11/28.
129420 | 42374.3073 |  | LI27 | 71 | 4 | HARDWARE | HW Problem | 27-7 MKSU shows 000 for rate.
129451 | 42374.8548 | 1 | LI23 | 21 | 4 | HARDWARE | HW Problem | 23-2 Trips off when placed on the beam with in 1 min.
129527 | 42376.9259 | 2 | LI00 | 97 | 4 | HARDWARE | HW Problem | LI00 SHB #1 shows "MOD" fault and zero AMPLIT. Won't reset.
129598 | 42380.4759 | 0.3 | LI15 | 31 | 4 | HARDWARE | HW Problem | KLYS 15-3 Modulator OFF, PAD looks ill.
129598 | 42380.4759 |  | LI15 | 31 | 4 | HARDWARE | HW Problem | KLYS 15-3 Modulator OFF, PAD looks ill.
129718 | 42384.9755 | 0.75 | DR13 | 1 | 4 | HARDWARE | HW Problem | DR13 Klystron 1 is tripped off and will not reset remotely.
130546 | 42425.7779 |  | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7 Modify modulator to prevent the ghost mod off signal.
130546 | 42425.7779 | 1 | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7 Modify modulator to prevent the ghost mod off signal.
130950 | 42447.6211 | 2 | LI30 | 81 | 4 | HARDWARE | HW Problem | 30-8 repeated MOD/HVR faults. Faults immediately on reset.
131137 | 42458.2758 |  | LI18 | 71 | 4 | HARDWARE | HW Problem | Lots of water summary faults from 18-7.
131165 | 42459.4172 | 1 | LI27 | 71 | 4 | HARDWARE | HW Problem | 27-7 low klystron uk.
131168 | 42459.4216 | 0.5 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4 has a CAMAC error. Does not respond to IPL. Shows PIOP Undefined Message error, and Processor Crashed.
131181 | 42460.2135 |  | LI02 | 51 | 4 | HARDWARE | HW Problem | Klys 2-5 is IPL faulted, probably a dead PIOP.
131231 | 42464.8588 | 1 | LI25 | 11 | 4 | HARDWARE | HW Problem | Klys 25-1 tripping frequently
131231 | 42464.8588 | 1 | LI25 | 11 | 4 | HARDWARE | HW Problem | Klys 25-1 tripping frequently
129044 | 42340.5268 |  | LI21 | 71 | 4 | HARDWARE | HW Problem | 21-7: Seeing some beam volts jitter.
129733 | 42387.2498 | 0.75 | LI11 | 61 | 4 | HARDWARE | HW Problem | 11-6 MOD fault and won't reset
129893 | 42394.4457 | 1 | LI04 | 11 | 4 | HARDWARE | HW Problem | klys 4-1 MOD fault, no reset.
130302 | 42412.3461 | 2 | LI28 | 31 | 4 | HARDWARE | HW Problem | Vac. gauge for 28-3 needs to be replaced.
130302 | 42412.3461 | 0.5 | LI28 | 31 | 4 | HARDWARE | HW Problem | Vac. gauge for 28-3 needs to be replaced.
130302 | 42412.3461 | 0.5 | LI28 | 31 | 4 | HARDWARE | HW Problem | Vac. gauge for 28-3 needs to be replaced.
130302 | 42412.3461 |  | LI28 | 31 | 4 | HARDWARE | HW Problem | Vac. gauge for 28-3 needs to be replaced.
130314 | 42412.7316 | 1 | LI29 | 71 | 4 | HARDWARE | HW Problem | Klystron 29-7 tripping repeatedly on KLV. Documented leaky wndow (JC).
130338 | 42415.8456 | 1 | LI28 | 11 | 4 | HARDWARE | HW Problem | 28-1 tripping repeatedly
130338 | 42415.8456 |  | LI28 | 11 | 4 | HARDWARE | HW Problem | 28-1 tripping repeatedly
130341 | 42415.9719 | 1 | LI20 | 71 | 4 | HARDWARE | HW Problem | 20-7 showing beam volts jitter, can see effect on beam.
130420 | 42419.1851 | 1 | LI05 | 81 | 4 | HARDWARE | HW Problem | 5-8 KLV
131250 | 42465.611 | 2 | LI21 | 71 | 4 | HARDWARE | HW Problem | Kly Mod 21-7: mod faulted, no reset.
131369 | 42476.6188 |  | LI29 | 31 | 4 | HARDWARE | HW Problem | Klys 29-3 has an MKSU trigger enable fault, but appears to be correctly actuating on the beam. See 130927. Frequently tripping on MKT now
131437 | 42482.4026 | 0.5 | LI23 | 21 | 4 | HARDWARE | HW Problem | 23-2: DQing current high at ~38 amps.
129050 | 42340.5603 | 0.5 | LI27 | 81 | 4 | HARDWARE | HW Problem | 27-8: Lower clip on PFN coil #1 should be replaced.
129784 | 42388.6461 |  | LI24 | 11 | 4 | HARDWARE | HW Problem | 24-1 Replaced T20 with oil, need to inspect A,B and C input bushing for oil leaks
130463 | 42421.9184 | 1.5 | LI28 | 31 | 4 | HARDWARE | HW Problem | Klys 28-3 is HVR and won't reset remotely
130944 | 42447.544 | 1 | LI02 | 51 | 4 | HARDWARE | HW Problem | 2-5: Lots of beam volts jitter.
131411 | 42480.6016 | 5 | LI28 | 31 | 4 | HARDWARE | HW Problem | Klystron 28-3 MOD faulted and won't reset remotely
131600 | 42496.4235 | 0.2 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 no remote reset since the station is in LOCAL control
130037 | 42400.7589 | 1 | LI02 | 71 | 4 | HARDWARE | HW Problem | station 2-7 has WGV fault
130521 | 42424.5533 | 0 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 MOD faulted with no remote reset
130591 | 42428.7276 | 1 | LI23 | 41 | 4 | HARDWARE | HW Problem | 23-4 mod fault, won't reset remotely
130685 | 42432.4256 | 1 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 is mod faulted and will not reset remotely.
131233 | 42464.9809 | 2 | LI26 | 21 | 4 | HARDWARE | HW Problem | 26-2 Mod fault. Thyratron reservoir interlock.
131289 | 42468.0145 |  | LI25 | 21 | 4 | HARDWARE | HW Problem | Kly. LI 25 - 2 Amplitude. dropped to 0.0 at 13:16 03/07/16 - think this is a readback problem. Klystron is definitely contributing energy to the beam.
131300 | 42469.9029 | 1 | LI24 | 51 | 4 | HARDWARE | HW Problem | Klystron LI24-51 trips off frequently on Mod EVOC fault.
131475 | 42486.5801 | 1 | LI30 | 41 | 4 | HARDWARE | HW Problem | 30-4 Mod has a burnt/broken Carborundum resistor
131485 | 42486.6297 | 2 | LI27 | 11 | 4 | HARDWARE | HW Problem | 27-1 - TBR Core bias is unstable when epy > 200KV.
129632 | 42381.679 | 4 | LI09 | 61 | 4 | HARDWARE | HW Problem | 9-6: Station drops control power when door #1 opened (intermittent). Still in MNT from fire alarm issue.
129732 | 42386.8742 | 1 | LI02 | 41 | 4 | HARDWARE | HW Problem | Klystron LI02-41 is tripped off on Mod Fault and will not reset remotely.
129875 | 42392.2132 | 1 | LI13 | 11 | 4 | HARDWARE | HW Problem | 13-1 MOD fault, no remote reset.
129876 | 42392.3467 | 1 | LI02 | 61 | 4 | HARDWARE | HW Problem | Kly Mod 2-6: faulted off due to WGV fault.
130127 | 42404.3701 | 4 | LI10 | 21 | 4 | HARDWARE | HW Problem | Kly. LI 10 - 2 keeps tripping off. Can reset but it trips off again in a few min.s .
130127 | 42404.3701 | 2 | LI10 | 21 | 4 | HARDWARE | HW Problem | Kly. LI 10 - 2 keeps tripping off. Can reset but it trips off again in a few min.s .
131230 | 42464.8142 | 1 | LI21 | 11 | 4 | HARDWARE | HW Problem | Klystron 21-1 MOD faulted and won't reset remotely
131595 | 42496.2833 |  | LI28 | 31 | 4 | HARDWARE | HW Problem | LI28-3 Water faults last night
129164 | 42346.7067 | 1 | LI20 | 71 | 4 | HARDWARE | HW Problem | The beam volts readback on Klys 20-7 (L0A) is seeing 3 kV spikes upward. They do not appear to be real as they don't seem to affect the beam.
129169 | 42347.0368 | 1 | LI28 | 11 | 4 | HARDWARE | HW Problem | 28-1 HVR, no remote reset.
129421 | 42374.3082 |  | LI27 | 81 | 4 | HARDWARE | HW Problem | 27-8 MKSU shows 000 for rate.
130971 | 42449.444 | 1 | LI21 | 31 | 4 | HARDWARE | HW Problem | 21-3: DQ current is low = 1A. It needs to calibrate.
131200 | 42461.4468 | 1.5 | LI29 | 71 | 4 | HARDWARE | HW Problem | 29-7 is Mod faulted and will not reset remotely.
131208 | 42461.9537 | 2 | LI22 | 21 | 4 | HARDWARE | HW Problem | Kly. LI 22 - 2 Phase is jumping from - 4.25 to +2.7. Do not see any change in Des. Phase or AMPLITUDE . ( see plot in MCC log. )
131320 | 42472.5174 | 0.75 | LI28 | 81 | 4 | HARDWARE | HW Problem | 28-8: Seeing some beam volts jitter.
131687 | 42505.4428 | 1 | LI28 | 51 | 4 | HARDWARE | HW Problem | 28-5 is FLK
129020 | 42339.4936 | 1.5 | LI20 | 81 | 4 | HARDWARE | HW Problem | Kly Mod 20-8 (L0B): re-install old style grid divider across thyratron to monitor anode/pfn volts waveform at front of mod BNC jackfield.
129120 | 42344.5593 |  | LI22 | 31 | 4 | HARDWARE | HW Problem | 22-3: Lock off 22-3 VVS breaker.
129823 | 42390.5059 |  | LI17 | 11 | 4 | HARDWARE | HW Problem | station 17-1 mod faulted
130414 | 42418.6428 | 0.75 | LI12 | 61 | 4 | HARDWARE | HW Problem | Klystron LI12-61 is tripped off on WGV and will not reset remotely.
131315 | 42472.0083 |  | LI22 | 81 | 4 | HARDWARE | HW Problem | Klys 22-8 started having multiple AMPL faults, seems to come and go. Maybe a watch and wait.
131498 | 42487.4985 |  | LI22 | 41 | 4 | HARDWARE | HW Problem | Klys 22-4 has high reflected RF.
131521 | 42489.9815 | 2 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3 MOD fault, doesn't reset remotely.
131539 | 42492.4575 | 2 | LI23 | 81 | 4 | HARDWARE | HW Problem | 23-8 FLK fault and won't reset remotely. When it is reset it cycles frequently.
131550 | 42493.3819 | 4 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3 FLK
131550 | 42493.3819 | 2 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3 FLK
129321 | 42354.6127 | 5 | LI21 | 61 | 4 | HARDWARE | HW Problem | Modulator 21-6: replace T20. Present T20 (air cooled delta-Y) causes ACOC faults when powered up (even at low HV setpoints such as 0.2kV) Removing all secondary connections and with no primary snubber, the problem persists.
129482 | 42375.725 | 0.5 | LI30 | 51 | 4 | HARDWARE | HW Problem | Klystron 30-5 MOD faulted and won't reset remotely
130342 | 42416.0577 | 1 | LI02 | 61 | 4 | HARDWARE | HW Problem | stations 2-6,16-7 and 2-3 have large phase jitter
130342 | 42416.0577 | 1 | LI02 | 61 | 4 | HARDWARE | HW Problem | stations 2-6,16-7 and 2-3 have large phase jitter
130342 | 42416.0577 | 1 | LI02 | 61 | 4 | HARDWARE | HW Problem | stations 2-6,16-7 and 2-3 have large phase jitter
130413 | 42418.5874 | 1 | LI22 | 11 | 4 | HARDWARE | HW Problem | 22-1 The "DEQ-SCR TRIG" LED on the modulator's front panel was out, see picture. The modulator itself is running, but the LED was not lit. It resides on HVPS regulator board. LED or the HVPS regulator board itself need to be replaced.
130416 | 42418.7038 | 1 | LI29 | 61 | 4 | HARDWARE | HW Problem | Klystron 29-6 HVR and won't reset remotely
130451 | 42420.4903 | 1 | LI12 | 21 | 4 | HARDWARE | HW Problem | Station 12-2 MOD fault with not remote reset.
131257 | 42466.0592 |  | LI30 | 81 | 4 | HARDWARE | HW Problem | 30-8 has no amplitude.
131455 | 42485.4544 |  | LI25 | 81 | 4 | HARDWARE | HW Problem | Klystron 25-8 mod faulting frequently. Refer to duplicate cater 128757
131455 | 42485.4544 | 2 | LI25 | 81 | 4 | HARDWARE | HW Problem | Klystron 25-8 mod faulting frequently. Refer to duplicate cater 128757
131455 | 42485.4544 | 2 | LI25 | 81 | 4 | HARDWARE | HW Problem | Klystron 25-8 mod faulting frequently. Refer to duplicate cater 128757
131469 | 42486.456 |  | LI27 | 11 | 4 | HARDWARE | HW Problem | 27-1 The BNC connectors of coaxial cables for the beam volt and beam current monitoring from the klystron tank are loose.
129177 | 42347.5236 | 1 | LI21 | 71 | 4 | HARDWARE | HW Problem | 21-7 needs to have PIOP ratio checked, since the MKSU was changed yesterday.
129779 | 42388.6126 |  | LI20 | 81 | 4 | HARDWARE | HW Problem | L0B no longer able to make up required RF amplitude. LLRF feedback setpoint abruptly increased at ~14:00 and then maxed itself out at 14:20. Klystron output is now too low to keep us on energy in DL1. MOD HVS 1/19/16 please see MCC ELOG: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=824873 Please find opportunity to check SBST operation there is some indication SBST might be saturating. Also there is a gain measurement in the LCLS physics ELOG.
130185 | 42406.461 | 1.5 | LI09 | 51 | 4 | HARDWARE | HW Problem | Station 9-5 MOD faulted no remote reset.
130720 | 42433.5966 | 3 | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7: The LED (D28) for HV On Stat MKSU Pwr of the PLC Interface board is not lit when station is running .
130959 | 42448.5786 | 1 | LI29 | 31 | 4 | HARDWARE | HW Problem | 29-3 is having frequent MKSU trigger enable faults which causes beam drop outs.
131018 | 42451.5525 | 0.2 | LI20 | 51 | 4 | HARDWARE | HW Problem | TCAV0 (20-5) mod fault, can't reset remotely. Station is needed for MD Wednesday, owl shift.
131611 | 42498.4974 | 0 | LI30 | 51 | 4 | HARDWARE | HW Problem | 30-5 multiple instances of thyratron klixon trip. Install klixon tell-tale.
128955 | 42335.7629 | 1 | LI05 | 11 | 4 | HARDWARE | HW Problem | Kly. LI 05 - 01 faulted on KMC Won't reset. Klystron magnet current intermittently jumping up to 17 amps from 15 amps.
129696 | 42384.3099 | 0.8 | LI25 | 11 | 4 | HARDWARE | HW Problem | Kly. LI 25 - 1 tripped on WGV" and won't reset. Station came back but AMRF and PEM found the vacuum gauge is reportedly jumpy and may need to be replaced.
129696 | 42384.3099 |  | LI25 | 11 | 4 | HARDWARE | HW Problem | Kly. LI 25 - 1 tripped on WGV" and won't reset. Station came back but AMRF and PEM found the vacuum gauge is reportedly jumpy and may need to be replaced.
129793 | 42389.147 | 1 | LI22 | 61 | 4 | HARDWARE | HW Problem | 22-6 very high Beam Volt Jiter
129899 | 42394.531 | 0.5 | LI24 | 31 | 4 | HARDWARE | HW Problem | Time accounting: klys 24-3 tripped off, but was stuck in local so we could not reset remotely
130496 | 42423.5151 | 1 | LI15 | 21 | 4 | HARDWARE | HW Problem | 15-2: Bad BVJ.
130569 | 42426.6357 |  | LI09 | 11 | 4 | HARDWARE | HW Problem | 9-1 WGV faulted, no remote reset
130626 | 42430.6532 | 0.75 | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7 is tripped off on HVR and will not reset remotely.
131303 | 42470.0153 | 1 | LI23 | 31 | 4 | HARDWARE | HW Problem | Klys 23-3 is MOD and won't reset remotely. Operator headed out to the field to try and reset remotely - reported that the station is in KHD, which the remote status does not show. Out of spares in L2 that don't trip constantly (I'm looking at you, 24-6), so this is causing downtime.
131358 | 42475.4559 |  | LI25 | 21 | 4 | HARDWARE | HW Problem | 25-2 Water faults intermittent. PIOP messages show ACC1 and WG2.
131476 | 42486.5814 | 1 | LI30 | 31 | 4 | HARDWARE | HW Problem | 30-3 Mod has a burnt PFN strap ( Upper, RHS, # 2 )
131515 | 42489.4034 | 0.1 | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3 phase and amplitude readback is bad (purple & NAN).
129365 | 42372.2907 | 1 | LI26 | 31 | 4 | HARDWARE | HW Problem | 26-3 has a thyratron filament Over Current fault that won't clear
129610 | 42380.692 | 1 | LI30 | 41 | 4 | HARDWARE | HW Problem | Klystron 30-4 electro-Magnet current fault
129871 | 42391.9525 | 1 | LI19 | 41 | 4 | HARDWARE | HW Problem | 19-4 HV meter not working
129968 | 42397.1719 |  | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 frequently faulting on HVR
131477 | 42486.5826 | 1.5 | LI30 | 11 | 4 | HARDWARE | HW Problem | 30-1 T20 output resistor is showing signs of arcing/burning Grid trigger chassis also needs to be replaced.
128926 | 42333.5485 | 0.75 | LI22 | 31 | 4 | HARDWARE | HW Problem | 22-3 MKSU Trigger fault, won't reset.
128960 | 42335.9607 | 1 | LI20 | 61 | 4 | HARDWARE | HW Problem | LCLS gun (20-6) has zero amplitude
129012 | 42339.0486 | 2 | LI12 | 31 | 4 | HARDWARE | HW Problem | 12-3 MOD and won't reset
129260 | 42352.0064 | 1 | LI09 | 11 | 4 | HARDWARE | HW Problem | 9-1 MOD fault.
129300 | 42353.5307 | 2 | LI03 | 41 | 4 | HARDWARE | HW Problem | 3-4: "HVR". Turned on remotely to EOLC faults.
129448 | 42374.8283 | 0.5 | LI24 | 41 | 4 | HARDWARE | HW Problem | 24-4 intermittent DeQuing under current faults. Current fluctuates between 10 and 9.5 amps.
129654 | 42382.5228 | 0.3 | LI26 | 31 | 4 | HARDWARE | HW Problem | 26-3: Seeing some beam volts jitter.
129654 | 42382.5228 | 1 | LI26 | 31 | 4 | HARDWARE | HW Problem | 26-3: Seeing some beam volts jitter.
129808 | 42389.5507 | 4 | LI20 | 81 | 4 | HARDWARE | HW Problem | 20-8 Modulator: Oil leaking out of right front of modulator. Replace T20
130051 | 42401.4498 | 0.5 | LI29 | 61 | 4 | HARDWARE | HW Problem | 29-6 has HVR fault.
130051 | 42401.4498 | 1 | LI29 | 61 | 4 | HARDWARE | HW Problem | 29-6 has HVR fault.
130051 | 42401.4498 | 2 | LI29 | 61 | 4 | HARDWARE | HW Problem | 29-6 has HVR fault.
130051 | 42401.4498 | 2 | LI29 | 61 | 4 | HARDWARE | HW Problem | 29-6 has HVR fault.
130051 | 42401.4498 | 2 | LI29 | 61 | 4 | HARDWARE | HW Problem | 29-6 has HVR fault.
130051 | 42401.4498 | 0.25 | LI29 | 61 | 4 | HARDWARE | HW Problem | 29-6 has HVR fault.
130051 | 42401.4498 |  | LI29 | 61 | 4 | HARDWARE | HW Problem | 29-6 has HVR fault.
130051 | 42401.4498 | 2 | LI29 | 61 | 4 | HARDWARE | HW Problem | 29-6 has HVR fault.
131234 | 42464.9865 |  | LI27 | 11 | 4 | HARDWARE | HW Problem | Kly Mod 27-1: core bias current unstable. Fuse in AC controller opens repeatedly.
131234 | 42464.9865 | 1 | LI27 | 11 | 4 | HARDWARE | HW Problem | Kly Mod 27-1: core bias current unstable. Fuse in AC controller opens repeatedly.
131234 | 42464.9865 |  | LI27 | 11 | 4 | HARDWARE | HW Problem | Kly Mod 27-1: core bias current unstable. Fuse in AC controller opens repeatedly.
131234 | 42464.9865 | 1.5 | LI27 | 11 | 4 | HARDWARE | HW Problem | Kly Mod 27-1: core bias current unstable. Fuse in AC controller opens repeatedly.
131234 | 42464.9865 | 3 | LI27 | 11 | 4 | HARDWARE | HW Problem | Kly Mod 27-1: core bias current unstable. Fuse in AC controller opens repeatedly.
131234 | 42464.9865 | 2 | LI27 | 11 | 4 | HARDWARE | HW Problem | Kly Mod 27-1: core bias current unstable. Fuse in AC controller opens repeatedly.
131234 | 42464.9865 | 2 | LI27 | 11 | 4 | HARDWARE | HW Problem | Kly Mod 27-1: core bias current unstable. Fuse in AC controller opens repeatedly.
131546 | 42492.9523 | 0.5 | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2 has a lot of beam volt jitter
131649 | 42501.2371 | 1 | LI28 | 51 | 4 | HARDWARE | HW Problem | 28-5 has very high phase and amplitude jitter.
131664 | 42502.2806 |  | LI20 | 71 | 4 | HARDWARE | HW Problem | L0A is experiencing beam volts dropouts. See this elog entry for evidence: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=844430
128933 | 42333.8404 | 0 | LI27 | 21 | 4 | HARDWARE | HW Problem | Kly. LI 27 - 2 phase jitter jumped to 1.5 after being at 1.0 before the POM.
129219 | 42349.3311 | 0.1 | LI03 | 51 | 4 | HARDWARE | HW Problem | Klystron 3-5 MOD fault with no remote reset
129308 | 42353.6674 | 1 | LI11 | 61 | 4 | HARDWARE | HW Problem | Kly. LI 11 - 6 tripped on "MOD" fault and won't reset.
129462 | 42375.2916 |  | LI29 | 71 | 4 | HARDWARE | HW Problem | 29-7 FLK from many Core Bias faults
129728 | 42386.3344 | 3 | LI28 | 51 | 4 | HARDWARE | HW Problem | 28-5 HVOV again, random trips
129846 | 42391.3837 | 4 | LI15 | 81 | 4 | HARDWARE | HW Problem | 15-8 BVLTs low, can't make up amplitude
129846 | 42391.3837 | 0.75 | LI15 | 81 | 4 | HARDWARE | HW Problem | 15-8 BVLTs low, can't make up amplitude
131527 | 42492.1018 | 2 | LI26 | 61 | 4 | HARDWARE | HW Problem | Klystron 26-6 VVS faulted, MOD off, no remote reset. PEM reports that the VVS breaker is tripped on this modulator, need a high voltage electrician to address it.
131602 | 42496.4601 |  | LI25 | 51 | 4 | HARDWARE | HW Problem | 25-5 seeing VERY infrequent (once every few minutes) bouts of BV dropouts: One bad shot followed by five good shots. Bouts last about three seconds then clear up. For FTPs see MCC ELOG entry: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=843702
131605 | 42496.6154 |  | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3: External Interlock faults. Suspect PIOP.
129066 | 42341.1812 | 0.8 | LI02 | 61 | 4 | HARDWARE | HW Problem | 2-6 bad phase & amplitude jitter.
129241 | 42350.1646 | 1 | LI22 | 21 | 4 | HARDWARE | HW Problem | Klystron 22-2 tripped on HVR with no remote reset
129311 | 42354.2398 | 1 | LI13 | 31 | 4 | HARDWARE | HW Problem | 13-3 WGV fault.
129363 | 42372.2877 | 1 | LI10 | 61 | 4 | HARDWARE | HW Problem | 10-6 has a WGV fault, Gauge Fail won't clear
129480 | 42375.7235 | 0.25 | LI29 | 11 | 4 | HARDWARE | HW Problem | Klystron 29-1 MOD faulted and can't be reset because it is in LOCAL
129570 | 42378.619 |  | LI26 | 61 | 4 | HARDWARE | HW Problem | Kly Mod 26-6: during next LCLS PAMM, open and inspect/repair KMB ckt. Station intermittently trips breaker and trim supply has been replaced. Wiring in mag junction box looks normal.
129613 | 42381.1912 |  | LI15 | 61 | 4 | HARDWARE | HW Problem | 15-6 The Plus5 and the Ref bot drift out of tols
129655 | 42382.5264 | 0.5 | LI28 | 71 | 4 | HARDWARE | HW Problem | 28-7: Lots of beam volts jitter.
130530 | 42425.1845 | 0.5 | LI28 | 11 | 4 | HARDWARE | HW Problem | Kly. LI 28 - 1 tripped off multiple times in the last 30 min.s After taking it off the beam, it's not trip in the 30 min.s
130592 | 42428.9476 | 1.5 | LI25 | 51 | 4 | HARDWARE | HW Problem | 25-5 tripped, called PEM to investigate. no remote reset.
130949 | 42447.5964 | 1.5 | LI13 | 61 | 4 | HARDWARE | HW Problem | 13-6 Core bias fault
131383 | 42478.5731 | 0.3 | LI30 | 41 | 4 | HARDWARE | HW Problem | Klystron 30-4 has a KMC fault. May be klystron klaxon. Recommend tell-tale be installed.
129014 | 42339.0851 | 1 | LI09 | 11 | 4 | HARDWARE | HW Problem | 9-1 MOD fault
129014 | 42339.0851 | 1 | LI09 | 11 | 4 | HARDWARE | HW Problem | 9-1 MOD fault
129553 | 42377.5354 | 2 | LI27 | 11 | 4 | HARDWARE | HW Problem | Klys 27-1 PHAS readback no good. PIOP IPL no help.
129560 | 42377.8025 | 1 | LI30 | 41 | 4 | HARDWARE | HW Problem | Klystron 30-4 MOD faulted and won't reset remotely
129565 | 42378.2785 | 1 | LI20 | 81 | 4 | HARDWARE | HW Problem | Station 20-8 (L0-B) is showing large amplitude deviations.
129565 | 42378.2785 | 2 | LI20 | 81 | 4 | HARDWARE | HW Problem | Station 20-8 (L0-B) is showing large amplitude deviations.
129635 | 42381.8384 |  | LI24 | 11 | 4 | HARDWARE | HW Problem | Klys 24-1 seems to be tripping on reflected energy faults. The RE value is slightly higher than last run.
129635 | 42381.8384 |  | LI24 | 11 | 4 | HARDWARE | HW Problem | Klys 24-1 seems to be tripping on reflected energy faults. The RE value is slightly higher than last run.
129883 | 42392.8212 | 2 | LI09 | 11 | 4 | HARDWARE | HW Problem | Klystron LI09-11 is tripped off on Mod Heater Delay.
129926 | 42395.472 |  | LI28 | 71 | 4 | HARDWARE | HW Problem | KYS:LI28:7 tripped off. reset and turned on twice. BVT. ran up to about 250 and tripped off again.
130462 | 42421.853 | 1 | LI17 | 11 | 4 | HARDWARE | HW Problem | Klys 17-1 is MOD and won't reset remotely
129063 | 42341.0188 | 1 | LI06 | 21 | 4 | HARDWARE | HW Problem | 6-2 WGV flt caused by low vac. When the filament is juiced there is no meter movement reaction. Connections are good. Possible faulting Cold Cathode tube. Transfer cater to AMRF to investigate.
129221 | 42349.4063 | 4.5 | LI04 | 41 | 4 | HARDWARE | HW Problem | Kly Mod 4-4: appears CB 7 is bad--has troubles opening and closing. Replace CB 7. TS/R thyratron overcurrents.
129597 | 42380.4744 | 0.25 | LI09 | 11 | 4 | HARDWARE | HW Problem | KLYS 9-1 MOD fault will not reset.
129597 | 42380.4744 | 0.25 | LI09 | 11 | 4 | HARDWARE | HW Problem | KLYS 9-1 MOD fault will not reset.
130507 | 42423.8647 | 2 | LI05 | 71 | 4 | HARDWARE | HW Problem | Klystron 5-7 TRG fault
130507 | 42423.8647 | 3 | LI05 | 71 | 4 | HARDWARE | HW Problem | Klystron 5-7 TRG fault
130507 | 42423.8647 | 1 | LI05 | 71 | 4 | HARDWARE | HW Problem | Klystron 5-7 TRG fault
130507 | 42423.8647 | 1.6 | LI05 | 71 | 4 | HARDWARE | HW Problem | Klystron 5-7 TRG fault
130507 | 42423.8647 | 0.5 | LI05 | 71 | 4 | HARDWARE | HW Problem | Klystron 5-7 TRG fault
130507 | 42423.8647 | 1 | LI05 | 71 | 4 | HARDWARE | HW Problem | Klystron 5-7 TRG fault
130811 | 42440.2546 | 1 | LI21 | 11 | 4 | HARDWARE | HW Problem | 22-5 WGV
130890 | 42445.4672 |  | LI21 | 41 | 4 | HARDWARE | HW Problem | 21-4 Mod has 2 ea badly burnt PFN straps that need to be replaced (Lower Middle )
130922 | 42446.4117 | 1 | LI24 | 51 | 4 | HARDWARE | HW Problem | 24-5: Lower PFN strap needs replacing.
131103 | 42456.9067 | 1 | LI02 | 71 | 4 | HARDWARE | HW Problem | LI02 - 7 kly. tripped on WGV. Attempted a manual vac. gauge reset using the filament button but it failed to cleared the fault.
131232 | 42464.9262 | 1 | LI30 | 71 | 4 | HARDWARE | HW Problem | Klys 30-7 keeps tripping. Resets remotely, briefly showing WKL faults. Wondering if that's related to CATER #130604
131232 | 42464.9262 | 1 | LI30 | 71 | 4 | HARDWARE | HW Problem | Klys 30-7 keeps tripping. Resets remotely, briefly showing WKL faults. Wondering if that's related to CATER #130604
131433 | 42482.2898 |  | LI22 | 21 | 4 | HARDWARE | HW Problem | 22-2 has had multiple water faults.
131609 | 42497.4585 |  | LI28 | 11 | 4 | HARDWARE | HW Problem | 28-1 runs fine at standby but trips immediately to HVR when set to accelerate. Switched to local, seems to run fine on accelerate when in local.
128924 | 42333.4884 | 1 | LI21 | 71 | 4 | HARDWARE | HW Problem | Kly Mod 21-7: BV backswing very unstable. Thyratron near end of life. Install tail clipper stack to temporarily improve BV amplitude stability until station replaced with AIP mod next week.
129121 | 42344.5809 | 1.5 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 Mod fault lockout, won't reset remotely.
129208 | 42348.5873 | 0.5 | LI21 | 61 | 4 | HARDWARE | HW Problem | 21-6 constantly cycling; see a klytron temperature fault coming in intermittently.
129790 | 42388.9422 |  | LI20 | 81 | 4 | HARDWARE | HW Problem | L0B (20-8) large amplitude jitter.
129863 | 42391.6081 | 1 | LI04 | 11 | 4 | HARDWARE | HW Problem | 4-1 trips off immediately when reset.
129640 | 42382.3002 |  | LI21 |  | 4 | HARDWARE | HW Problem | 21-3 and 21-4 will not turn on after VVS cb swap
129712 | 42384.6781 | 0.5 | LI28 | 51 | 4 | HARDWARE | HW Problem | Klys 28-5 MOD fault, no remote reset.
129786 | 42388.7394 | 1 | LI28 | 61 | 4 | HARDWARE | HW Problem | 28-6 mod fault
130652 | 42431.5449 | 1 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4: Replace upper PFN strap from 7-to-8.
130665 | 42431.7204 | 1 | LI23 | 41 | 4 | HARDWARE | HW Problem | Klystron 23-4 HVR fault We can reset it, but it trips off shortly afterwards.
130728 | 42435.563 | 1 | LI24 | 61 | 4 | HARDWARE | HW Problem | KLYS 24-6 is tripping repeatedly
130728 | 42435.563 | 1.3 | LI24 | 61 | 4 | HARDWARE | HW Problem | KLYS 24-6 is tripping repeatedly
128952 | 42335.5281 | 0.5 | LI11 | 61 | 4 | HARDWARE | HW Problem | 11-6 MOD fault, won't reset.
128953 | 42335.5284 | 0.5 | LI13 | 11 | 4 | HARDWARE | HW Problem | 13-1 MOD fault, won't reset.
129186 | 42347.7781 | 0 | LI27 | 81 | 4 | HARDWARE | HW Problem | Klystron 27-8 mod faulting incessantly. We are able to reset it, but it mod faults again seconds afterwards. It does not appear to happen while the klystron is on standby.
129258 | 42351.9673 | 1 | LI09 | 81 | 4 | HARDWARE | HW Problem | After an elevated pressure incident in sector 9 (whole sector WGV), once recovered 9-8 will not reset FLK.
129563 | 42377.9706 |  | LI18 | 61 | 4 | HARDWARE | HW Problem | 18-6 bad PIOP
129966 | 42396.7801 | 1 | LI17 | 71 | 4 | HARDWARE | HW Problem | 17-7: K13 fault stepper sticky and only resets one count. It may be prudent to doctor all stations with the old mechanical K13's for the optimization of the last FACET run.
129969 | 42397.1782 | 1 | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 MOD faulting frequently.
128905 | 42332.862 | 1 | LI29 | 71 | 4 | HARDWARE | HW Problem | 29-7 KLV fault
130415 | 42418.703 | 2 | LI29 | 31 | 4 | HARDWARE | HW Problem | Klystron 29-3 MOD faulted and won't reset remotely. HV on, thyratron not firing.
130586 | 42428.0212 | 1 | LI03 | 71 | 4 | HARDWARE | HW Problem | Kly. LI03 - 7 phase jitter is getting large. 0.3 when I scanned it.
130891 | 42445.4687 | 1 | LI22 | 31 | 4 | HARDWARE | HW Problem | 22-3 Mod has a broken 15K bleeder resistor in the PFN cab
130968 | 42449.4149 | 0.2 | LI22 | 41 | 4 | HARDWARE | HW Problem | 22-4 FTPs show intermittent amplitude dropouts.
131017 | 42451.5524 | 1 | LI30 | 51 | 4 | HARDWARE | HW Problem | 30-5: Seeing some beam volts jitter.
131047 | 42452.6835 | 1.5 | LI20 | 41 | 4 | HARDWARE | HW Problem | FACET XTCAV Klystron 20-4B appears to have something wrong with the TWT.
129004 | 42338.5902 | 2 | LI27 | 81 | 4 | HARDWARE | HW Problem | 27-8 needs PFN tune
129496 | 42376.2097 | 1 | LI26 | 51 | 4 | HARDWARE | HW Problem | 26-5 CBL flt will not clear
129619 | 42381.3177 | 0.1 | LI15 | 31 | 4 | HARDWARE | HW Problem | 15-3 PAD needs replaced. out of tols diplay cleared with PAD subst. Box.
129952 | 42396.469 |  | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2 tripped off "KMC" no remote reset.
129974 | 42397.4289 | 1 | LI19 | 31 | 4 | HARDWARE | HW Problem | LI19-3 KLY water faults last night and this morning
130831 | 42441.2496 | 1 | LI05 | 71 | 4 | HARDWARE | HW Problem | 5-7 CBL FLT. At the station there is a PIOP to MKSU connection. I disconnected from LI05 CR04 slot 2 and connected to slot 3 or 5-8 modulator. The MKSU on 5-7 worked fine with a rate. Problem point to the PIOP in slot 2 or the Crate PIOP connection.
130835 | 42441.9892 |  | LI20 | 71 | 4 | HARDWARE | HW Problem | Bookkeeping Cater for time accounting: Station 20-7 (a.k.a. L0-A) restart caused a vacuum burst that caused valve 155 to go in (tripped gauge 365).
130897 | 42445.5951 | 0.3 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3 Mod Grid trigger assembly has a broken component and needs to be replaced.
130912 | 42446.0101 | 2 | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7 MKSU Swap. Time accounting cater.
130933 | 42447.3849 |  | LI15 | 71 | 4 | HARDWARE | HW Problem | Klystron 15-7 says "IPL required", probably needs a PIOP swap.
130963 | 42448.6616 | 1 | LI25 | 81 | 4 | HARDWARE | HW Problem | Kly Mod 25-8: frequently FLKs and trips off on ext. interlock.
130995 | 42450.7079 | 1 | LI15 | 71 | 4 | HARDWARE | HW Problem | Klystron 15-7 mod faulted and won't reset remotely.
131089 | 42455.5273 | 1.5 | LI01 | 51 | 4 | HARDWARE | HW Problem | 1-5 MOD fault. No remote reset.
128988 | 42338.1982 | 0.5 | LI14 | 81 | 4 | HARDWARE | HW Problem | 14-8 MOD fault and won't reset
128994 | 42338.321 | 4 | LI11 | 61 | 4 | HARDWARE | HW Problem | 11-6 immediately goes to MOD fault whenever reset.
129067 | 42341.1817 | 1 | LI03 | 71 | 4 | HARDWARE | HW Problem | 3-7 bad phase & amplitude jitter.
129199 | 42348.4878 | 1 | LI09 | 11 | 4 | HARDWARE | HW Problem | 9-1 mod fault, can't reset remotely.
129246 | 42350.5728 | 2 | LI22 | 21 | 4 | HARDWARE | HW Problem | Klys 22-2 went HVR in local.
129246 | 42350.5728 | 2 | LI22 | 21 | 4 | HARDWARE | HW Problem | Klys 22-2 went HVR in local.
129246 | 42350.5728 | 5 | LI22 | 21 | 4 | HARDWARE | HW Problem | Klys 22-2 went HVR in local.
129246 | 42350.5728 | 2 | LI22 | 21 | 4 | HARDWARE | HW Problem | Klys 22-2 went HVR in local.
129246 | 42350.5728 | 3 | LI22 | 21 | 4 | HARDWARE | HW Problem | Klys 22-2 went HVR in local.
129246 | 42350.5728 | 1 | LI22 | 21 | 4 | HARDWARE | HW Problem | Klys 22-2 went HVR in local.
129419 | 42374.2867 | 3 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 has a klystron delta temp and KMC faults showing on the MKSU also, SLED motor cable is unplugged from MKSU. 60-20 focus magnet supply appears ok, problem may be the MKSU
129652 | 42382.4981 | 7 | LI21 | 71 | 4 | HARDWARE | HW Problem | 21-7: PFN has unreasonable slope. Please don't close until fully evaluated.
129652 | 42382.4981 | 3 | LI21 | 71 | 4 | HARDWARE | HW Problem | 21-7: PFN has unreasonable slope. Please don't close until fully evaluated.
129766 | 42388.4671 | 3 | LI28 | 11 | 4 | HARDWARE | HW Problem | 28-1 MOD fault, doesn't reset remotely.
129766 | 42388.4671 | 3 | LI28 | 11 | 4 | HARDWARE | HW Problem | 28-1 MOD fault, doesn't reset remotely.
129766 | 42388.4671 | 1 | LI28 | 11 | 4 | HARDWARE | HW Problem | 28-1 MOD fault, doesn't reset remotely.
129858 | 42391.4919 | 1 | LI13 | 21 | 4 | HARDWARE | HW Problem | 13-2 MOD fault, doesn't reset remotely.
129936 | 42396.2459 | 1 | LI19 | 41 | 4 | HARDWARE | HW Problem | 19-4 Door No. 5 locking latch under the key lock missing.
129951 | 42396.4683 | 0.2 | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3 HVR no remote turn on as station is in LOCAL.
135746 | 42908.6829 | 1 | LI28 | 81 | 4 | HARDWARE | HW Problem | 28-8 is in mod fault lock out and cannot be reset remotely.
135124 | 42873.4314 | 0.5 | LI22 | 81 | 4 | HARDWARE | HW Problem | station 22-8 faults repeatedly--3 times in the last hour
135129 | 42873.5961 | 1 | LI28 | 11 | 4 | HARDWARE | HW Problem | 28-1 tripped on Core Bias
135249 | 42878.6642 | 1.3 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4 core bias current reads OK, core bias voltage abnormally high. Troubleshoot and repair
136538 | 42980.276 | 1.5 | LI26 | 71 | 4 | HARDWARE | HW Problem | 26-7: Fire wire fault.
136539 | 42980.277 | 1.5 | LI30 | 71 | 4 | HARDWARE | HW Problem | 30-7: Fire wire fault.
136540 | 42980.29 | 0.03 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4 FLK.
136543 | 42980.4558 | 1 | LI26 | 21 | 4 | HARDWARE | HW Problem | Kly Mod 26-2: station tripped off due to external faults.
136544 | 42980.4768 | 0 | LI26 | 21 | 4 | HARDWARE | HW Problem | 26-2 tripped HVOV reported.
136545 | 42980.4773 | 0 | LI29 | 81 | 4 | HARDWARE | HW Problem | 29-8 tripped ACOC reported
136547 | 42980.5254 | 0.5 | LI22 | 21 | 4 | HARDWARE | HW Problem | 22-2 WGV appears to be asleep needs filament juiced.
136549 | 42980.6179 |  | LI27 | 41 | 4 | HARDWARE | HW Problem | 27-4 is tripping on WA1 flowswitch faults.
136549 | 42980.6179 | 1 | LI27 | 41 | 4 | HARDWARE | HW Problem | 27-4 is tripping on WA1 flowswitch faults.
136550 | 42980.6322 | 0.1 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4 tripped on klystron vacuum. Readout history of klys vac looks suspiciously non-physical with huge diurnal temperature correlated swings.
136553 | 42980.6988 |  | LI27 | 31 | 4 | HARDWARE | HW Problem | 27-3 focus fault.
136556 | 42980.737 | 2 | LI28 | 61 | 4 | HARDWARE | HW Problem | 28-6 fire wire trip.
136557 | 42980.7779 | 2 | LI25 | 51 | 4 | HARDWARE | HW Problem | 25-5 MOD fault lockout. It was MOD VVS faulted and changed to FLK. Unclear what is actually wrong.
134398 | 42817.5762 | 1.2 | LI22 | 11 | 4 | HARDWARE | HW Problem | 22-1: problems found: 1. 480VAC cables not properly secured. 2. Primary snubber resistors: wires and lugs burnt. 3. Two filter cap 4.75uF missing green ground wires. 4. The big ground green terminal block that mounted to top left corner in the T-20 compartment" has cable loose.
134955 | 42867.2185 |  | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7 has low beam volts.
135570 | 42897.4317 | 0.5 | LI26 | 11 | 4 | HARDWARE | HW Problem | 26-1 Modulator problems found: 1. Primary snubber resistor R260, R261 burnt lugs. 2. HV filter 4.7uF, 30,000VDC cap missing ground wires. 4. T2-xformer 120VAC ground wire loose. 5. Core Bias Assembly: TB1 phoenix block has 120VAC BLK & WHT wires: strands protruding from the connection.
135570 | 42897.4317 | 0.5 | LI26 | 11 | 4 | HARDWARE | HW Problem | 26-1 Modulator problems found: 1. Primary snubber resistor R260, R261 burnt lugs. 2. HV filter 4.7uF, 30,000VDC cap missing ground wires. 4. T2-xformer 120VAC ground wire loose. 5. Core Bias Assembly: TB1 phoenix block has 120VAC BLK & WHT wires: strands protruding from the connection.
135570 | 42897.4317 | 1 | LI26 | 11 | 4 | HARDWARE | HW Problem | 26-1 Modulator problems found: 1. Primary snubber resistor R260, R261 burnt lugs. 2. HV filter 4.7uF, 30,000VDC cap missing ground wires. 4. T2-xformer 120VAC ground wire loose. 5. Core Bias Assembly: TB1 phoenix block has 120VAC BLK & WHT wires: strands protruding from the connection.
135570 | 42897.4317 | 1 | LI26 | 11 | 4 | HARDWARE | HW Problem | 26-1 Modulator problems found: 1. Primary snubber resistor R260, R261 burnt lugs. 2. HV filter 4.7uF, 30,000VDC cap missing ground wires. 4. T2-xformer 120VAC ground wire loose. 5. Core Bias Assembly: TB1 phoenix block has 120VAC BLK & WHT wires: strands protruding from the connection.
135570 | 42897.4317 | 1 | LI26 | 11 | 4 | HARDWARE | HW Problem | 26-1 Modulator problems found: 1. Primary snubber resistor R260, R261 burnt lugs. 2. HV filter 4.7uF, 30,000VDC cap missing ground wires. 4. T2-xformer 120VAC ground wire loose. 5. Core Bias Assembly: TB1 phoenix block has 120VAC BLK & WHT wires: strands protruding from the connection.
135571 | 42897.4431 | 1 | LI26 | 21 | 4 | HARDWARE | HW Problem | 26-2 Modulator problems found: 1.Primary snubber resistor R261 burnt wire, wire insulation cracked. 2. HV filter 4.7uF 30,000VDC cap missing ground wires. 4. Two De-spiking coils covers cracked (see attached picture).
135572 | 42897.4496 | 1 | LI26 | 31 | 4 | HARDWARE | HW Problem | 26-3 Modulator problems found: 1. Two primary snubber caps: jumper wires between them are burnt. 2. HV filter 4.7uF 30,000VDC cap missing ground wires. 3. Door#2: top 600VAC holes measured= 2 1/2 inches diameter holes should be covered.
135574 | 42897.4525 | 0.5 | LI26 | 41 | 4 | HARDWARE | HW Problem | 26-4 Modulator problems found: 1. Primary snubber resistor R261 lug burnt. 2. HV filter 4.7uF 30,000VDC cap missing ground wires. 4. Bad Grid Trigger chassis. Found screw that hold the resistor and coils together is broken. ( item #4 fixed by PN)
135574 | 42897.4525 | 1 | LI26 | 41 | 4 | HARDWARE | HW Problem | 26-4 Modulator problems found: 1. Primary snubber resistor R261 lug burnt. 2. HV filter 4.7uF 30,000VDC cap missing ground wires. 4. Bad Grid Trigger chassis. Found screw that hold the resistor and coils together is broken. ( item #4 fixed by PN)
135575 | 42897.599 | 1 | LI26 | 51 | 4 | HARDWARE | HW Problem | 26-5 Modulator problems found: 1. Primary snubber resistor R261 lug burnt. 2. HV filter 4.7uF 30,000VDC cap missing ground wires. 4. PFN strap bottom (RHS #4).
135771 | 42912.2418 | 2 | LI29 | 51 | 4 | HARDWARE | HW Problem | 29-5 modulator tripping repeatedly on End of Line Clipper fault.
135955 | 42924.6461 | 2 | LI26 | 61 | 4 | HARDWARE | HW Problem | Kly Mod 26-6: station tripped on variety of faults including ACOC and HVOC. Station in local mode.
135955 | 42924.6461 | 1 | LI26 | 61 | 4 | HARDWARE | HW Problem | Kly Mod 26-6: station tripped on variety of faults including ACOC and HVOC. Station in local mode.
135957 | 42925.2914 | 1 | LI22 | 41 | 4 | HARDWARE | HW Problem | Klystron 22-4 tripping periodically on ACOC faults.Trips several times a day but does reset remotely.
136561 | 42981.4696 | 0.5 | LI27 | 41 | 4 | HARDWARE | HW Problem | 27-4 is in FLK
136578 | 42983.5422 | 1 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3 MOD fault, no reset
135027 | 42868.727 |  | LI25 | 21 | 4 | HARDWARE | HW Problem | 25-2 Has a waveguide vacuum fault.
135028 | 42868.7292 |  | LI25 | 11 | 4 | HARDWARE | HW Problem | 25-1 has low beam voltage at 313KV expected is 350KV.
135029 | 42868.7302 |  | LI26 | 11 | 4 | HARDWARE | HW Problem | 26-1 has low beam voltage at 320KV expected is 350KV.
135031 | 42868.7334 | 0.5 | LI26 | 51 | 4 | HARDWARE | HW Problem | 26-5 has low beam voltage at 325KV expected is 345KV.
135045 | 42869.3111 | 2 | LI21 | 41 | 4 | HARDWARE | HW Problem | 21-4 WCS LCW leak from flow restrictor. Butler is on it.
135587 | 42898.5396 | 1 | LI26 | 81 | 4 | HARDWARE | HW Problem | 26-8 is off on KHD and won't clear
135588 | 42898.5402 | 0.5 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4 is KMC
135596 | 42898.8159 | 1 | LI20 | 81 | 4 | HARDWARE | HW Problem | Klystron 20-8 (L0B) has a large amount of beam volts jitter
136047 | 42932.291 |  | LI26 | 51 | 4 | HARDWARE | HW Problem | 26-5 window 2 temperature (KLYS:LI26:51:WNDW2_TEMP) is 149deg F.
136047 | 42932.291 |  | LI26 | 51 | 4 | HARDWARE | HW Problem | 26-5 window 2 temperature (KLYS:LI26:51:WNDW2_TEMP) is 149deg F.
136051 | 42932.5127 | 1.5 | LI29 | 61 | 4 | HARDWARE | HW Problem | 29-8 frequent trips "MOD EVOC"
136054 | 42932.6951 |  | LI21 | 41 | 4 | HARDWARE | HW Problem | 21-4 CBL fault does not reset with PIOP IPL or PDU reset.
136054 | 42932.6951 | 1.5 | LI21 | 41 | 4 | HARDWARE | HW Problem | 21-4 CBL fault does not reset with PIOP IPL or PDU reset.
136059 | 42933.0376 | 0.5 | LI21 | 51 | 4 | HARDWARE | HW Problem | 21-5 has MKT fault. PDU reset, PIOP reset, or IPL do not recover.
136579 | 42983.6405 | 1.5 | LI29 | 21 | 4 | HARDWARE | HW Problem | 29-2: The PIOP has been frozen since 9/5/17 @ 1517 and will not clear with an IPL.
136586 | 42983.8491 | 0.5 | LI25 | 81 | 4 | HARDWARE | HW Problem | Klystron 25-8 MOD FLK and won't reset remotely
136592 | 42984.4604 | 1.5 | LI24 | 81 | 4 | HARDWARE | HW Problem | 24-8 Klystron beam voltage reads 32KV low on PIOP data.
135171 | 42875.3209 | 2 | LI21 | 11 | 4 | HARDWARE | HW Problem | L1S window 2 temp is reading 180 degF should be 95 degF. Has been this way for two days now. Ops noticed this while troubleshooting observed HV problems with L1S modulator. The value seems high and should be addressed asap. Steele 5/21/17: This problem started on May first, before RF had been turned on. Must be a controls readback problem, over to CTL. This is a duplicate of CATER 134920.
135618 | 42899.5239 | 1 | LI24 | 41 | 4 | HARDWARE | HW Problem | 24-4 tripped off. Can reset but it trips immediately.
136034 | 42930.5645 | 0.5 | LI28 | 61 | 4 | HARDWARE | HW Problem | 28-6: Seeing some beam volts jitter.
136043 | 42931.653 | 1.5 | LI29 | 31 | 4 | HARDWARE | HW Problem | 29-3 MOD fault no reset, fault is ACOC.
136060 | 42933.3867 | 1.5 | LI25 | 81 | 4 | HARDWARE | HW Problem | When assigned to accelerate time, 25-8 trips on "MKT" fault every few minutes. This does not happen when assigned to standby time.
136270 | 42949.3592 | 0.5 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4 Station wave guide VAC-ION does not respond when set point buttons are pushed.
136596 | 42984.6239 |  | LI16 | 41 | 4 | HARDWARE | HW Problem | 16-4B S/N 396C and 16-4C S/N 250A have poor klystron vacuum.
136603 | 42985.4147 | 0.15 | LI24 | 61 | 4 | HARDWARE | HW Problem | Klystron LI24-61 is tripped off on a WGV fault and needs a local reset.
135178 | 42875.7134 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X (21-2) goes MKT (MKSU trigger enable fault) fairly frequently for several minutes at a time. This happens during the heat of the day. It seems to heat up to the point where it trips MK, then cools down, runs for 30 minutes, then trips again. It stops when it is sufficiently cool outside. Also a downtime accounting CATER.
135182 | 42876.2819 | 4 | LI20 | 81 | 4 | HARDWARE | HW Problem | 20-8 modulator: replace bad T20
135186 | 42876.4732 | 2 | LI26 | 81 | 4 | HARDWARE | HW Problem | 26-8: Klystron heater problems. Bad power supply and/or klystron heater problem.
135190 | 42876.6441 |  | LI20 | 81 | 4 | HARDWARE | HW Problem | 20-8: Raise BV's back to 350kV, check PIOP ratios.
135197 | 42877.1942 |  | LI20 | 81 | 4 | HARDWARE | HW Problem | L0B still having bouts of instability. Amplitude drifts away with no sign of BV variation . Maybe check the drive. The amplitude is very stable at 10Hz and maintained stability for two hours. When going back to 120Hz, The ampl immediately dropped. Zeroed offsets and ran pau re-synch script and the ampl is stable now (not as much as 10Hz). Stable enough for the water flow to handle the amount of change. After about 5 minutes, the amplitude excursions start to grow.
136063 | 42933.6372 | 1 | LI28 | 71 | 4 | HARDWARE | HW Problem | 28-7 tripping frequently ACOC (four times last hour)
136080 | 42934.9144 | 1 | LI20 | 71 | 4 | HARDWARE | HW Problem | Klystron LI20-71 (L0A) is tripped off on Water Summary Fault. Accelerator Water #2 fault indicated.
136081 | 42935.0194 | 0.5 | LI21 | 41 | 4 | HARDWARE | HW Problem | 21-4 has an MKSU Protect fault and does not reset remotely. PIOP data shows Electro-Magnet Tol fault. Beam current and Beam volt readouts are unrealistically high.
136243 | 42947.3215 | 1.5 | LI20 | 81 | 4 | HARDWARE | HW Problem | L0B has high BV jitter.
136243 | 42947.3215 |  | LI20 | 81 | 4 | HARDWARE | HW Problem | L0B has high BV jitter.
136246 | 42947.5925 | 1.2 | LI27 | 31 | 4 | HARDWARE | HW Problem | 27-3 Modulator: inspect T20 for leakage, clean leaked oil and check (and top off if needed) oil level.
136289 | 42951.8849 | 2 | LI28 | 81 | 4 | HARDWARE | HW Problem | 28-8 is mod faulted and will not reset remotely.
136292 | 42952.5597 | 1 | LI29 | 71 | 4 | HARDWARE | HW Problem | Klystron 29-7 appears to have tripped its fire wire.
136292 | 42952.5597 | 2 | LI29 | 71 | 4 | HARDWARE | HW Problem | Klystron 29-7 appears to have tripped its fire wire.
136294 | 42953.571 | 1 | LI21 | 51 | 4 | HARDWARE | HW Problem | Klystron 21-5 MKSU trigger enable fault Happened again, 8/11/2017 11:31
136294 | 42953.571 | 0.5 | LI21 | 51 | 4 | HARDWARE | HW Problem | Klystron 21-5 MKSU trigger enable fault Happened again, 8/11/2017 11:31
136296 | 42953.8694 | 1 | LI24 | 31 | 4 | HARDWARE | HW Problem | Klystron 24-3 MOD fault lockout
136306 | 42954.5693 |  | LI20 | 81 | 4 | HARDWARE | HW Problem | L0B MOD faulting repeatedly. Now locked out.
136615 | 42986.4642 |  | LI25 | 41 | 4 | HARDWARE | HW Problem | Klystron LI25-41 is in Mod Fault Lockout and will not reset remotely.
136616 | 42986.4769 | 1 | LI25 | 41 | 4 | HARDWARE | HW Problem | 25-4: Tripping on EOLC faults.
136616 | 42986.4769 | 50 | LI25 | 41 | 4 | HARDWARE | HW Problem | 25-4: Tripping on EOLC faults.
136619 | 42986.9835 |  | LI23 | 41 | 4 | HARDWARE | HW Problem | 23-4 shows BVJ but jitter FTPs do not show substantial RMS jitter.
136620 | 42986.9839 |  | LI21 | 51 | 4 | HARDWARE | HW Problem | 21-5 shows BVJ but jitter FTPs do not show substantial RMS jitter.
136622 | 42987.6365 | 2 | LI23 | 31 | 4 | HARDWARE | HW Problem | Klys 23-3 showing bad cable status. Tried IPL, failed hard, might be PIOP, unclear.
136622 | 42987.6365 | 1.5 | LI23 | 31 | 4 | HARDWARE | HW Problem | Klys 23-3 showing bad cable status. Tried IPL, failed hard, might be PIOP, unclear.
135210 | 42877.413 | 1 | LI27 | 71 | 4 | HARDWARE | HW Problem | 27-7 is FLK
135222 | 42878.3311 |  | LI25 | 11 | 4 | HARDWARE | HW Problem | Klystrons 25-1, 26-6, 27-7 MOD faulted and won't reset remotely
136253 | 42948.2116 | 1 | LI24 | 31 | 4 | HARDWARE | HW Problem | Klystron 24-3 faulting frequently. Three times in as many minutes. Edit 8/9/2017 08:30- this station still faulting frequently. Edit 8/23/2017 21:08 - This station is still faulting frequently, seems to fault more often when it is on accelerate than in standby. Edit 9/8/2017- this station still faulting frequently
136253 | 42948.2116 |  | LI24 | 31 | 4 | HARDWARE | HW Problem | Klystron 24-3 faulting frequently. Three times in as many minutes. Edit 8/9/2017 08:30- this station still faulting frequently. Edit 8/23/2017 21:08 - This station is still faulting frequently, seems to fault more often when it is on accelerate than in standby. Edit 9/8/2017- this station still faulting frequently
136253 | 42948.2116 | 0.5 | LI24 | 31 | 4 | HARDWARE | HW Problem | Klystron 24-3 faulting frequently. Three times in as many minutes. Edit 8/9/2017 08:30- this station still faulting frequently. Edit 8/23/2017 21:08 - This station is still faulting frequently, seems to fault more often when it is on accelerate than in standby. Edit 9/8/2017- this station still faulting frequently
136264 | 42948.5802 |  | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2 showing large amounts of phase jitter on the SsSB output.
136308 | 42954.7243 | 1 | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2 Faults on klystron focus current (KMC and EMC, regulator 1 and 2) when focus current is changed.
136308 | 42954.7243 | 0.5 | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2 Faults on klystron focus current (KMC and EMC, regulator 1 and 2) when focus current is changed.
136312 | 42954.7907 |  | LI28 | 21 | 4 | HARDWARE | HW Problem | Station 28-2 modulator is running seemingly normally by the CUD status PV: CUDKLYS:LI28:2:STATUS reports HVR
136351 | 42960.6496 | 1 | LI26 | 61 | 4 | HARDWARE | HW Problem | 26-6 tripped on VVS breaker. Requires electrician to reset with VVS ON. Please reset Monday morning, thanks. This is the second occurrence in three weeks. Please see: Cater 136216 7-27-2017
136388 | 42964.4273 | 1 | LI24 | 21 | 4 | HARDWARE | HW Problem | 24-2: PFN has unreasonable slope and ripple.
134325 | 42807.5573 |  | LI29 | 41 | 4 | HARDWARE | HW Problem | 29-4 klystron has leaky filament connector.
135230 | 42878.3969 |  | LI25 | 41 | 4 | HARDWARE | HW Problem | 25-4 Waveguide vacuum controller is faulted.
135232 | 42878.415 | 2 | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 L1X - Klystron focus current readbacks are all zero.
135288 | 42880.8424 | 0.5 | LI25 | 31 | 4 | HARDWARE | HW Problem | Klystron LI25-3 will not go back to SLED mode remotely after changing to non-SLED mode.
136147 | 42940.3065 |  | LI20 | 81 | 4 | HARDWARE | HW Problem | L0B tripped on over-voltage three times on 7/24 owl shift.
136147 | 42940.3065 |  | LI20 | 81 | 4 | HARDWARE | HW Problem | L0B tripped on over-voltage three times on 7/24 owl shift.
136265 | 42948.6261 |  | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2 Cannot saturate klystron. DAC out is maxed out at 32K.
136395 | 42965.1214 | 0 | LI23 | 81 | 4 | HARDWARE | HW Problem | 23-8 starting tripping off incessantly starting at ~1:30am on 8/18 with about 10 trips in about an hour.
134351 | 42813.2635 | 2 | LI21 | 21 | 4 | HARDWARE | HW Problem | LI21, Power supply GEN 12.5-60, voltage and current are unstable. It located at rack KF21-2C elevation 23 near Modulator 21-2.
134352 | 42813.2723 | 1 | LI20 | 71 | 4 | HARDWARE | HW Problem | 20-7: Core Bias Assembly - 120VAC harness "P2" the secure clip that plug into the Pheonix connector is broken.
134355 | 42814.2688 | 1 | LI21 | 11 | 4 | HARDWARE | HW Problem | 21-1: Core Bias Assembly: DC harness P1" secure clip is broken.
134357 | 42814.2768 | 1 | LI21 | 31 | 4 | HARDWARE | HW Problem | 21-3 Problems found: 1. Core Bias Assembly, 120VAC harness P2" secure clip broken.
135337 | 42883.5397 | 1 | LI29 | 51 | 4 | HARDWARE | HW Problem | 29-5 MOD fault
135339 | 42883.6684 | 2 | LI26 | 11 | 4 | HARDWARE | HW Problem | 26-1 trips every several minutes.
135345 | 42884.3447 | 1.5 | LI24 | 41 | 4 | HARDWARE | HW Problem | 21-4 Station needs PFN tuning for a flatter pulse top and a slope adjustment better than .5.
136321 | 42956.4285 | 0.5 | LI22 | 71 | 4 | HARDWARE | HW Problem | 22-7 showing large phase jitter, visible effect on beam.
136329 | 42956.8205 | 1.2 | LI29 | 71 | 4 | HARDWARE | HW Problem | 29-7 tripped fire wire and has "phase" and "VVS" faults showing on PLC. Locally no fire no smoke no smell of smoke.
136408 | 42968.6817 | 1 | LI25 | 21 | 4 | HARDWARE | HW Problem | Klys 25-2 cycling frequently. Also, it shows 255 errors for random faults on error display and the AMPL readback jumps to 68.73 frequently. Right now showing FOX error. Seems like problem with PAD, but unclear.
136416 | 42969.5591 |  | LI25 | 21 | 4 | HARDWARE | HW Problem | 25-2 PIOP froze after IPL and may needs swap.
135363 | 42885.426 | 2 | LI25 | 11 | 4 | HARDWARE | HW Problem | 25-1: PAD phase read back not working (shows nan). Replace PAD.
135365 | 42885.4518 | 0.5 | LI22 | 31 | 4 | HARDWARE | HW Problem | 22-3 is faulted on WGV
136077 | 42934.5462 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | Klystron 21-2 MOD faulted and won't reset remotely
136331 | 42957.3303 | 1 | LI22 | 71 | 4 | HARDWARE | HW Problem | 22-7: Tune PFN for a flatter signal top. Analysis slope is unreasonable (1).
136333 | 42957.5402 |  | LI21 | 31 | 4 | HARDWARE | HW Problem | 21-3:DQ > 400 amps on EPICS even after upgrades.
136334 | 42957.5422 |  | LI25 | 21 | 4 | HARDWARE | HW Problem | Klystron 25-2 tripping repeatedly on HVOV faults.
136338 | 42957.7619 |  | LI28 | 21 | 4 | HARDWARE | HW Problem | RFS 28-2 focus breaker is tripped.
136338 | 42957.7619 | 1 | LI28 | 21 | 4 | HARDWARE | HW Problem | RFS 28-2 focus breaker is tripped.
136338 | 42957.7619 | 1 | LI28 | 21 | 4 | HARDWARE | HW Problem | RFS 28-2 focus breaker is tripped.
136341 | 42958.4738 | 3 | LI23 | 81 | 4 | HARDWARE | HW Problem | 23-8: ACR log reports rash of HVOC faults.
136341 | 42958.4738 | 0 | LI23 | 81 | 4 | HARDWARE | HW Problem | 23-8: ACR log reports rash of HVOC faults.
136342 | 42958.5248 | 1 | LI21 | 51 | 4 | HARDWARE | HW Problem | 21-5: DQ current appears to be stuck at 20 amps.
134413 | 42820.5542 | 0.5 | LI29 | 11 | 4 | HARDWARE | HW Problem | 29-1 Modulator: Problems found: 1. Burnt secondary snubber resistor lug that cable connects to R1 of xformer. 2. Burnt core bias filter reactor, must be replaced before turn modulator back on. (Put blue block with tag) 3. Thyratron base bolts rusted. 4. The connector that plug to Klystron tank, the thread is damaged.
134414 | 42820.5605 | 0.5 | LI29 | 21 | 4 | HARDWARE | HW Problem | 29-2 Modulator: Problems found: 1. Burnt secondary snubber resistor lugs. 2. HV filter caps 4.75uF missing two ground wires. 3. P2 pheonix cable connector that connects to core bias assembly, the secure clip is broken. 4. PFN strap #4 (LHS) upper loose 5. De-Spiking coil resistors, there is one resistor has wrong resistance value=15000 ohm, should be 2000 ohm.
134415 | 42820.5669 | 0.5 | LI29 | 41 | 4 | HARDWARE | HW Problem | 29-3 Modulator problems found: 1. Core Bias Assembly: 120VAC P2" cable connector secure clip is broken.
134418 | 42820.5831 | 1 | LI29 | 61 | 4 | HARDWARE | HW Problem | 29-6 Modulator problems found: 1. Primary snubber resistor: All wires, lugs and bottom resistor R262" should be replaced.
134419 | 42820.5855 | 0.5 | LI29 | 71 | 4 | HARDWARE | HW Problem | 29-7 Modulator problems found: 1. Burnt primary wires and lugs, they need to be replaced. 2. Two filter cap cases are missing ground wires.
134420 | 42820.5868 | 0.5 | LI29 | 81 | 4 | HARDWARE | HW Problem | 29-8: Two HV filter cap cases are missing ground wires.
135380 | 42885.7005 | 1.5 | LI20 | 61 | 4 | HARDWARE | HW Problem | 20-6 is unstable. two trips within as many minutes. EVO faults before FLK.
135380 | 42885.7005 | 2 | LI20 | 61 | 4 | HARDWARE | HW Problem | 20-6 is unstable. two trips within as many minutes. EVO faults before FLK.
135888 | 42918.4945 |  | LI28 | 31 | 4 | HARDWARE | HW Problem | Klystron LI28-31 is tripped on Mod Fault and will not reset remotely.
136143 | 42939.3858 | 6 | LI20 | 61 | 4 | HARDWARE | HW Problem | Getting amplitude dropouts on klys 20-6 (gun). Not seen in modulator. Appears to be LLRF problem. Problem got very bad starting around 7/23 @ 09:00.
136344 | 42958.6048 | 1 | LI25 | 41 | 4 | HARDWARE | HW Problem | 25-4: Tune PFN for flatter signal top. Slope is presently -1.7.
136346 | 42958.8887 | 1.5 | LI21 | 51 | 4 | HARDWARE | HW Problem | 21-5 MKSU trigger enable fault. No triggers to modulator. See 135294.
136347 | 42958.9497 | 0.7 | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3 after bout of very frequent trips (not clear source of trips) it is no Fault Lockout and will not reset. See 136253.
136357 | 42961.6641 |  | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2 Phase feedback appears to be broken. tkurty 8/14/17 16:00: T4 phase feedback working correctly, T1 feedback was stuck between -180 and 180 deg. Changing PDES to 10 deg allowed T1 to leave the phase-flip region and stabilize to the same value as T4. Reset PDES to 0 and both timeslot feedbacks are working Broken state: [softegr@lcls-srv01 tkurty]$ caget KLYS:LI28:21:T{1,4}:CORRPHASE KLYS:LI28:21:T1:CORRPHASE 75.4542 KLYS:LI28:21:T4:CORRPHASE 92.8884 After resetting PDES: [softegr@lcls-srv01 tkurty]$ caget KLYS:LI28:21:T{1,4}:CORRPHASE KLYS:LI28:21:T1:CORRPHASE 90.9837 KLYS:LI28:21:T4:CORRPHASE 92.4988
136424 | 42970.5459 | 1 | LI20 | 51 | 4 | HARDWARE | HW Problem | 20-5 Modulator problems found: 1. T20 - phase B R2" base loose ground.
136424 | 42970.5459 | 1 | LI20 | 51 | 4 | HARDWARE | HW Problem | 20-5 Modulator problems found: 1. T20 - phase B R2" base loose ground.
136426 | 42970.6969 |  | LI30 | 11 | 4 | HARDWARE | HW Problem | Klystron 30-1 tripped its VVS breaker
136432 | 42971.4455 | 1 | LI20 | 61 | 4 | HARDWARE | HW Problem | 20-6 Modulator problems found: 1. T20 compartment: 120VAC T2" xformer - loose ground wire.
136432 | 42971.4455 | 1 | LI20 | 61 | 4 | HARDWARE | HW Problem | 20-6 Modulator problems found: 1. T20 compartment: 120VAC T2" xformer - loose ground wire.
136432 | 42971.4455 | 1 | LI20 | 61 | 4 | HARDWARE | HW Problem | 20-6 Modulator problems found: 1. T20 compartment: 120VAC T2" xformer - loose ground wire.
136439 | 42972.3156 |  | LI24 | 21 | 4 | HARDWARE | HW Problem | 24-2 is running without DQing regulation =PHJ, AMJ and BVJ.
134430 | 42822.4497 | 4.5 | LI20 | 71 | 4 | HARDWARE | HW Problem | 20-7: Replace old ground hook on door 4 with new style hook. Add hanger for hook safety.
134431 | 42822.4515 | 2 | LI23 | 51 | 4 | HARDWARE | HW Problem | 23-5: Replace old ground hook on door 4 with new style hook. Add hanger for hook safety.
134432 | 42822.453 | 2 | LI26 | 11 | 4 | HARDWARE | HW Problem | 26-1: Replace old ground hook on door 4 with new style hook. Add hanger for hook safety.
134433 | 42822.4548 | 2.5 | LI30 | 41 | 4 | HARDWARE | HW Problem | 30-4: Replace old ground hook on door 4 with new style hook. Add hanger for hook safety.
135385 | 42886.0145 | 0.1 | LI26 | 31 | 4 | HARDWARE | HW Problem | LI26-31 repeatedly flashes "Stat_Bad_Camac" and does not have a PAD Phase value.
135395 | 42886.5375 | 0.5 | LI27 | 31 | 4 | HARDWARE | HW Problem | Klystron 27-3 MOD fault and no reset
135403 | 42886.8037 |  | LI20 | 81 | 4 | HARDWARE | HW Problem | Klys 20-8 (L0B) has a tough time regulating at 1 Hz. See it in Beam Volts/Phas/AMPL. Not sure if this is timing or LLRF feedback problem. Unable to get FTP at 1 Hz to see if firing time is shifted. Might need a scope.
136358 | 42961.8082 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 L1X is showing frequent instances of running into saturation and causing the LLRF feedback to push to drive power to the software clipping limit (over saturation). See MCCELOG: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=904787 MOD 8/14/17 19:41 HVS Further investigation indicates that the saturation might be in the drive amplifier rather than the klystron itself. See MCCELOG: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=904788
136359 | 42961.9879 | 1 | LI20 | 81 | 4 | HARDWARE | HW Problem | Seeing increased jitter from 20-8 L0B modulator. Please check history of jitter PV SIOC:SYS0:ML00:AO216 and scope modulator.
136366 | 42962.3892 | 2 | LI26 | 61 | 4 | HARDWARE | HW Problem | 26-6 VVS CB/ VVS13 is running hot (See pic). Needs to be inspected.
136367 | 42962.4334 | 1 | LI22 | 11 | 4 | HARDWARE | HW Problem | Klystron 22-1 having numerous phase dropouts noticeable on FEL at low energy.
136442 | 42972.5198 | 1 | LI26 | 31 | 4 | HARDWARE | HW Problem | 26-3 FLK, has an EOLC fault.
136454 | 42973.6122 | 2 | LI20 | 81 | 4 | HARDWARE | HW Problem | Kly Mod 20-8 (L0B): station experiencing BV dropouts.
135425 | 42887.7492 |  | LI20 | 61 | 4 | HARDWARE | HW Problem | GUN PAC LLRF Feedback (AMPL) is railed I have had to increase the MAX setpoint in the LLRF feedback from 5100 to 5800 to maintain 6MeV off the gun. New setpoint settles at 5420 counts but is now decreasing as I write this entry. it is down to 5300 now.
135429 | 42888.0683 | 1 | LI20 | 81 | 4 | HARDWARE | HW Problem | L0B is MOD faulted.
135429 | 42888.0683 | 1 | LI20 | 81 | 4 | HARDWARE | HW Problem | L0B is MOD faulted.
136374 | 42963.2856 | 1 | LI30 | 41 | 4 | HARDWARE | HW Problem | 30-4 Has no core bias voltage.
136379 | 42963.4877 | 2 | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3 DQing current can't be adjusted below 25 amps
136466 | 42974.2725 | 1 | LI20 | 71 | 4 | HARDWARE | HW Problem | 20-7 Modulator problem found: 1. Primary snubber resistor R261: lug and wire should be replaced. 2. Pulse cable seems over heat, burnt the insulation.
136468 | 42974.3141 | 1 | LI26 | 71 | 4 | HARDWARE | HW Problem | 26-7 Modulator problems found: 1. Primary snubber resistors: R260 & R261 lugs need to be replaced. R262 lug & wire need to be replaced. 2. T20 compartment: filter caps missing ground wires. 3. Xformer T2" Loose screw, see up load picture."
136469 | 42974.3227 | 0.8 | LI26 | 81 | 4 | HARDWARE | HW Problem | 26-8 Modulator problems found: 1. Primary snubber resistors: R261 and R262 burnt lugs
134621 | 42848.3031 | 2 | LI24 | 81 | 4 | HARDWARE | HW Problem | 24-8 Modulator problems found: 1. Wires touching the primary resistors body. 3. Oil leaks around the T-20 base and cables. 4. HV filter caps missing two ground wires. 5. Under Thyratron: two 115VAC covers are still unsafe. They need to be remake. 6. Door #7: the third 20 ohm carborundum res from left need to be replaced.
134622 | 42848.4619 | 1.5 | LI25 | 11 | 4 | HARDWARE | HW Problem | 25-1 Modulator problems found: 1. Replace lug and wire for primary snubber res R261. 2. HV filter caps missing two ground wires. 4. Grid trigger spring coil is rusted.
134623 | 42848.4715 | 1 | LI25 | 21 | 4 | HARDWARE | HW Problem | 25-2 Modulator problems found: 1. Primary snubber resistors: R260, R262 only lugs need to be replaced. R261 lug & wire need to be replaced. 3. HV filter caps missing two ground wires. 4. Terminal block TB13 firewire lug is corrosion. 5. T2 120VAC xformer on floor loose ground wire. 6. Door #5: #2AWG ground wire loose (top right)
134624 | 42848.5268 | 1 | LI25 | 31 | 4 | HARDWARE | HW Problem | 25-3 Modulator problems found: 2. Primary snubber resistors: a)wire touching R260 body, reroute it away ; b)replace lugs for R261, R262 3. Door #4 top loose ground wires #2AWG. 4. HV filter caps missing two ground wires.
135446 | 42888.4617 | 1.5 | LI23 | 31 | 4 | HARDWARE | HW Problem | Klys 23-3 seeing large (10 deg or more) phase jumps frequently.
136475 | 42974.6729 | 1 | LI27 | 31 | 4 | HARDWARE | HW Problem | 27-3 KMC faulted.
136476 | 42974.6835 | 2 | LI21 | 41 | 4 | HARDWARE | HW Problem | 21-4 MOD faulted with no remote reset. Appears to be a tripped fire wire.
136486 | 42975.5612 |  | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3: Tripping on SCR klixon interlock fault.
135523 | 42893.444 | 1 | LI25 | 41 | 4 | HARDWARE | HW Problem | 25-4: KMC fault
135530 | 42893.6268 | 0.5 | LI21 | 41 | 4 | HARDWARE | HW Problem | Klys 21-4 showing BVLT jitter.
136498 | 42976.4242 |  | LI20 | 71 | 4 | HARDWARE | HW Problem | 20-7 (aka L0A): Water leaking from PAD hose connection.
136500 | 42976.5709 | 2 | LI27 | 11 | 4 | HARDWARE | HW Problem | 27-1: Seeing some beam volts jitter.
134613 | 42845.5087 | 1 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 Modulator problems found: 2. Primary resisistors: R261, R262 bunrt lugs & wires. 3. HV filter caps missing two ground wires to EG. 4. Cable K" from the thyratron connect to the keep alive is touching the metal sharp edge.
134514 | 42834.6002 | 0.7 | LI23 | 51 | 4 | HARDWARE | HW Problem | 23-5 Modulator problems found: 1. Primary snubber resistors: lugs and wires need to be replaced. 2. HV filter caps missing two ground wires. 3. 120Vac T2" xfomer: loose ground.
134899 | 42865.6904 | 2 | LI20 | 61 | 4 | HARDWARE | HW Problem | LI20 klys focus magnet power supplies reading 0 V for 20-6 and 20-7. PS shows ON (normal color) Resis = -1 ohms (purple) voltage - 0V (purple) Current = 14.8 amps (red)
136220 | 42944.368 | 1 | LI24 | 11 | 4 | HARDWARE | HW Problem | 24-1 tripped and will not reset. PEM to reset please.
136227 | 42944.6009 | 0.2 | LI23 | 21 | 4 | HARDWARE | HW Problem | 23-2: Phase jitter above spec limit. Amplitude jitter very low.
136508 | 42977.5727 | 2 | LI26 | 71 | 4 | HARDWARE | HW Problem | 26-7: Seeing lots of beam volts jitter.
134394 | 42817.5618 | 1 | LI21 | 51 | 4 | HARDWARE | HW Problem | 21-5: Primary snubber resistors: R261, R262 wires and lugs need to be replaced.
134395 | 42817.5646 | 0.5 | LI21 | 61 | 4 | HARDWARE | HW Problem | 21-6: Primary snubber resistors: R261 wire and lug burnt; R262 loose screw
134397 | 42817.5714 | 1.2 | LI21 | 81 | 4 | HARDWARE | HW Problem | 21-8: Problems found: 1. 480VAC cables are touching the metal sharp edges of xfomer. 2. The snubber resistors: R261 and R262 have wires and lugs burnt.
134397 | 42817.5714 | 2 | LI21 | 81 | 4 | HARDWARE | HW Problem | 21-8: Problems found: 1. 480VAC cables are touching the metal sharp edges of xfomer. 2. The snubber resistors: R261 and R262 have wires and lugs burnt.
135553 | 42895.3836 | 1 | LI26 | 51 | 4 | HARDWARE | HW Problem | Klystron LI26-51 has amplitude and beam volt jitter that exceeds tolerance levels.
135554 | 42895.4239 | 1.5 | LI29 | 11 | 4 | HARDWARE | HW Problem | 29-1: Remove thyratron for use at 20-6. Install new thyratron in this station.
135559 | 42895.5499 |  | LI25 | 51 | 4 | HARDWARE | HW Problem | 25-5C Repair klystron ion pump electrical leakage
135681 | 42904.6343 | 1 | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3 has mod faulted and cannot be reset remotely
135084 | 42871.944 | 1 | LI20 | 61 | 4 | HARDWARE | HW Problem | Frequent Gun dropouts. Klystron 20-6 is having periodic bouts of extreme amplitude dropouts. See E-Log entry: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=890229
136513 | 42978.2466 | 1 | LI27 | 41 | 4 | HARDWARE | HW Problem | 27-4 is MOD faulted and won't reset
136518 | 42978.678 | 1 | LI27 | 31 | 4 | HARDWARE | HW Problem | Klystron 27-3 KMC faulted.
136523 | 42979.4244 | 1 | LI30 | 61 | 4 | HARDWARE | HW Problem | 30-6: Seeing lots of beam volts jitter.
136524 | 42979.4451 |  | LI23 | 31 | 4 | HARDWARE | HW Problem | station 23-3 has bad Cable status Verified it wasn't the PAD or MKSU PIOP swap doesn't seem to have helped - tried two spares just in case, left second spare in place. - M.Harms 5-Sep-17
136524 | 42979.4451 |  | LI23 | 31 | 4 | HARDWARE | HW Problem | station 23-3 has bad Cable status Verified it wasn't the PAD or MKSU PIOP swap doesn't seem to have helped - tried two spares just in case, left second spare in place. - M.Harms 5-Sep-17
136525 | 42979.4722 | 0.1 | LI30 | 61 | 4 | HARDWARE | HW Problem | station 30-6 MOD faults reset, but only lasts a couple of minutes before faulting again
136532 | 42979.6729 | 0.5 | LI24 | 61 | 4 | HARDWARE | HW Problem | station 24-6 Klystron Waveguide fault
134741 | 42858.6097 | 1 | LI30 | 51 | 4 | HARDWARE | HW Problem | 30-5: Keep alive power supply has no output.
134748 | 42859.2748 |  | LI20 | 51 | 4 | HARDWARE | HW Problem | 20-5 PIOP appears in-op. repair or replace
134749 | 42859.2765 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 PIOP appears in-op. repair or replace
134750 | 42859.2778 |  | LI23 | 41 | 4 | HARDWARE | HW Problem | 23-4 PIOP appears in-op. repair or replace
135128 | 42873.5873 | 2 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 tripping on thyratron trigger faults
135135 | 42873.8594 | 2.5 | LI23 | 11 | 4 | HARDWARE | HW Problem | 23-1 getting intermittent accelerate triggers. Seems to mostly not trigger correctly but sometimes does.
135142 | 42874.2712 | 1 | LI26 | 41 | 4 | HARDWARE | HW Problem | 26-4 trips on EOLCs
135142 | 42874.2712 | 1 | LI26 | 41 | 4 | HARDWARE | HW Problem | 26-4 trips on EOLCs
134497 | 42831.5672 | 1.2 | LI22 | 31 | 4 | HARDWARE | HW Problem | 22-3 Modulator: Problems found: 1. Primary snubber resistors: R261 and R262 burnt lugs & wires. 2. Fire wire lugs that connected to TB13, the lug is corrosion.
134498 | 42831.5737 | 1.2 | LI22 | 41 | 4 | HARDWARE | HW Problem | 22-4 Modulator: problems found: 1. Primary snubber resistors: R260, R261, and R262 burnt lugs/wires 2. Fire wires TB12: lug is corrosion (see picture).
134500 | 42831.5974 | 1.2 | LI22 | 61 | 4 | HARDWARE | HW Problem | 22-6 Modulator, Problems found: 1. Primary snubber resistors: R261 & R262 burnt lugs/wires.
134500 | 42831.5974 | 1.2 | LI22 | 61 | 4 | HARDWARE | HW Problem | 22-6 Modulator, Problems found: 1. Primary snubber resistors: R261 & R262 burnt lugs/wires.
134501 | 42831.5996 | 1.2 | LI22 | 71 | 4 | HARDWARE | HW Problem | 22-7 Modulator: problems found: 1. Primary snubber resistors: R260, R261 and R262 lugs/wires burnt. 2. Despiking Coil insulation burnt.
134510 | 42834.5794 | 1 | LI23 | 11 | 4 | HARDWARE | HW Problem | 23-1 Modulator problems found: 1. T-20 : oil leaks 2. Primary snubber resistors: R260, R261, R262 lugs and wires burnt 3. HV filter caps missing two ground wires. 4. Core Bias Assembly: the two wires cross the AC/DC current xductor are over heat and melted the plastic of the AC/DC current xdu5ctor. 5. De-Spring Coil cover is cracked.
134511 | 42834.5887 | 0.75 | LI23 | 21 | 4 | HARDWARE | HW Problem | 23-2 Modulator problems found: 1. Primary snubber resistors: R261, R262 lugs and wires burnt. 2. HV filter caps are missing two ground wires.
134778 | 42862.5913 | 1.5 | LI25 | 51 | 4 | HARDWARE | HW Problem | 25-5 Modulator problems found: 1. T-20 leaks oil on cables and the T-20 base. 2. Primary snubber resistors: R261, R262 lugs and wires burnt. 3. HV filter caps missing ground wires. 4. Door #2: top 600VAC holes = 2 1/2 inches should be covered. 5. PFN strap #7 top (LHS) loose, strap needs to be replaced.
134780 | 42862.5993 | 2 | LI25 | 61 | 4 | HARDWARE | HW Problem | 25-6 Modulator problems found: 1. Primary snubber resistor R261: lug &wire burnt. 2. HV filter caps missing ground wires. 3. Door#2: top 600VAC holes = 2 1/2 inches should be covered.
134781 | 42862.6032 | 2 | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 Modulator problems found: 1. T2-xformer 120VAC ground wire loose. 2. T-20 base ground wire loose. 3. Primary snubber resistor R260 burnt lug. 4. HV filter cap missing ground wires. 5. Door#2: top 600VAC holes = 2 1/2 inches should be covered. 6. TB1 phoenix connector 120VAC black wires should be re-terminated and re-insert properly. I think it is a Core bias output wire, not 120 VAC -p.stiles
134550 | 42838.5892 | 1 | LI24 | 11 | 4 | HARDWARE | HW Problem | 24-1 Modulator problems found: 1. Primary snubber resistors: R261, R262 Lugs/wires need to be replaced.
134551 | 42838.5924 | 2 | LI24 | 21 | 4 | HARDWARE | HW Problem | 24-2 Modulator problems found: 1. Primary snubber resistors: R261, R262 lugs / wires need to be replaced. 2. Door #5: Top right corner - ground wire loose. 3. Despiking coil cover is cracked.
134553 | 42838.5961 | 1 | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3 Modulator problems found: 1. Primary snubber resistors: R261 lug /wire needs to be replaced. 2. HV filter caps missing ground wires. 3. T2- 120VAC xformer base : loose ground wire. 4. EOLC stack missing mounting screw. 5. Cable is touching the despiking coil body.
134553 | 42838.5961 | 1 | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3 Modulator problems found: 1. Primary snubber resistors: R261 lug /wire needs to be replaced. 2. HV filter caps missing ground wires. 3. T2- 120VAC xformer base : loose ground wire. 4. EOLC stack missing mounting screw. 5. Cable is touching the despiking coil body.
134557 | 42839.4406 | 0.5 | LI23 | 21 | 4 | HARDWARE | HW Problem | 23-2 Has a klystron vacuum fault.
134770 | 42860.4607 | 1 | LI25 | 11 | 4 | HARDWARE | HW Problem | 25-1 KMC fault. Trim p/s breaker trips.
135159 | 42874.6007 | 1 | LI25 | 41 | 4 | HARDWARE | HW Problem | 25-4 Replace HV regulator board.
135163 | 42874.7551 |  | LI24 | 11 | 4 | HARDWARE | HW Problem | 24-1 tripping regularly; have had to reset several times in last two hours.
135163 | 42874.7551 | 1.5 | LI24 | 11 | 4 | HARDWARE | HW Problem | 24-1 tripping regularly; have had to reset several times in last two hours.
134512 | 42834.5921 | 0.5 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3 Moduator problems found: 1. Primary snubber resistor R260: burnt lug 2. HV filter caps missing two ground wires.
134513 | 42834.5951 | 0.5 | LI23 | 41 | 4 | HARDWARE | HW Problem | 23-4 Modulator problems found: 1. T-20 ground loose screw. 2. Primary snubber resistor R260 burnt lug; R261, R262 burnt wires. 3. HV filter caps missing ground wires 4. Cable C" below Thyratron, the insulation is cracked."
134516 | 42834.604 | 0.5 | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7 Modulator problem found: 1. Primary snubber resistors R261, R262: lugs/wires burnt 2. HV filter caps missing two ground wires.
134517 | 42834.6061 | 0.8 | LI23 | 81 | 4 | HARDWARE | HW Problem | 23-8 Modulator problems found: 1. Primary snubber resistor R260: lug burnt 2. HV filter caps missing two ground wires. 3. Door #5: the ground wire that located on top right corner loose.
135289 | 42880.8432 |  | LI27 | 41 | 4 | HARDWARE | HW Problem | Klystron LI27-4 will not go back to SLED mode remotely after changing to non-SLED mode.
135297 | 42881.2713 | 3 | LI22 | 61 | 4 | HARDWARE | HW Problem | Klys 22-6 experiencing frequent KLV faults.
135297 | 42881.2713 | 1.2 | LI22 | 61 | 4 | HARDWARE | HW Problem | Klys 22-6 experiencing frequent KLV faults.
135297 | 42881.2713 |  | LI22 | 61 | 4 | HARDWARE | HW Problem | Klys 22-6 experiencing frequent KLV faults.
135242 | 42878.5845 | 1.5 | LI25 | 11 | 4 | HARDWARE | HW Problem | 25-1 MOD faulted
135331 | 42882.5983 | 0.5 | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3 was tripping on Reflected Energy
135332 | 42882.6208 | 1 | LI25 | 71 | 4 | HARDWARE | HW Problem | 25-7 Reflected energy fault
135450 | 42888.5924 | 2 | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X occasionally shows MKT which trips the beam. If we are lucky, the fault goes away in 5 minutes. We are not sure what is causing this fault. 6/5/2017 edit: although L1X stopped going MKT after swapping the MKSU and SCR, it is now doing so again (has gone MKT three times in last two hours, for roughly ~8minutes each time).
135496 | 42892.0446 | 1 | LI23 | 61 | 4 | HARDWARE | HW Problem | Klystron LI23-61 has high beam volts jitter.
135499 | 42892.2962 | 1 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3 HVR reset .
135510 | 42892.5973 | 2.5 | LI26 | 61 | 4 | HARDWARE | HW Problem | 26-6 is KHD
135300 | 42881.3495 | 0.7 | LI28 | 71 | 4 | HARDWARE | HW Problem | staion 28-7 Control Power Fault
134826 | 42864.5718 |  | LI30 | 21 | 4 | HARDWARE | HW Problem | 30-2: Touch panel display blanks out and goes back to "HOME" while doing the set-up. It happens quite often during the settings.
135686 | 42904.7374 | 1 | LI21 | 51 | 4 | HARDWARE | HW Problem | 21-5 MOD fault, no reset: NOTE: problem is Fire Wire short - Modulator was VERY hot, I staged a high voltage rectifier at the modulator but was sent to SPEAR and didn't make it back.
135712 | 42906.7422 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 (L1X) modulator HV has a strong 16.6 Hz timing jitter, see http://physics-elog.slac.stanford.edu/lclselog/show.jsp?dir=/2017/25/20.06&pos=2017-06-20T17:35:31
135713 | 42906.7959 | 8 | LI20 | 81 | 4 | HARDWARE | HW Problem | 20-8 L0B MOD fault. Attempts to restart greeted immediately with ACOC fault.
135722 | 42907.5159 |  | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2 KMC fault.
135314 | 42881.6303 | 4 | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3 trips immediately after turn on with HV overcurrent fault showing.
135320 | 42881.952 | 1 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4 has MASSIVE beam volts jitter.
135830 | 42915.442 | 1.3 | LI29 | 61 | 4 | HARDWARE | HW Problem | 29-6 has intermittent BV dropouts.
135845 | 42916.295 |  | LI20 | 51 | 4 | HARDWARE | HW Problem | LI20 TCAV0 trip on WG1 last night 2140 hrs.
134854 | 42865.2582 |  | LI22 | 21 | 4 | HARDWARE | HW Problem | 22-2 klystron focus supply does not respond to on/off commands, shows purple status on EPICS.
134873 | 42865.4063 |  | LI21 | 81 | 4 | HARDWARE | HW Problem | 21-8 has a LVPS fault that won't clear
134874 | 42865.4095 |  | LI21 | 31 | 4 | HARDWARE | HW Problem | 21-3 has a T20 fan fault that won't clear
135731 | 42908.2445 | 0.3 | LI20 | 61 | 4 | HARDWARE | HW Problem | 20-6 (GUN): showing occasional bouts of beam volts jitter Scope hooked up. Thyratron replace with good broken-in one from 29-1 Running OK, not great. ACR happy.
135920 | 42922.1306 | 2.5 | LI23 | 61 | 4 | HARDWARE | HW Problem | 23-6 ACOCs on turn on.
135921 | 42922.2888 | 1 | LI29 | 61 | 4 | HARDWARE | HW Problem | 29-6: Shows AMJ. causing 30% jitter on FEL intensity. DEACT'd, pulse stability improved immediately.
135924 | 42922.4105 | 2 | LI30 | 41 | 4 | HARDWARE | HW Problem | 30-4: Thyratron becoming unstable, nearing EOL.
135926 | 42922.4135 | 2 | LI26 | 11 | 4 | HARDWARE | HW Problem | 26-1: Thyratron becoming unstable, maybe near EOL.
135927 | 42922.4152 | 2 | LI27 | 81 | 4 | HARDWARE | HW Problem | 27-8: Thyratron becoming unstable, maybe near EOL.
135943 | 42923.5632 | 1 | LI26 | 61 | 4 | HARDWARE | HW Problem | 26-6 modulator T20 has oil around the base. Inspect bushings (input and output) for sighs of new leakage. Inspect oil level.
135948 | 42923.6601 |  | LI25 | 11 | 4 | HARDWARE | HW Problem | 25-1: Leaking oil at bottom of pulse cable and oil dropping on the floor from the klystron.
135954 | 42924.5915 | 0.5 | LI21 | 51 | 4 | HARDWARE | HW Problem | 21-5 tripled off on MKT fault could not reset remotely or at the station.
135965 | 42926.2675 | 2.5 | LI23 | 71 | 4 | HARDWARE | HW Problem | Klystron 23-7 has a lot of phase jitter and spikes on the beam volts
135982 | 42927.0386 | 1.5 | LI26 | 61 | 4 | HARDWARE | HW Problem | Klys 26-6 has VVS breaker fault. Be good if it can get fixed before 09:00 on 7/11.
136029 | 42930.4264 | 0.5 | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2: Thyratron Compartment Fan Fault
136106 | 42936.3422 | 2 | LI21 | 81 | 4 | HARDWARE | HW Problem | 21-8 amplitude is dropping out very frequently.
136106 | 42936.3422 | 1.3 | LI21 | 81 | 4 | HARDWARE | HW Problem | 21-8 amplitude is dropping out very frequently.
136139 | 42938.4848 | 0.5 | LI25 | 21 | 4 | HARDWARE | HW Problem | 25-2: Station keeps dropping out on external interlock. follow up: water trap shows many ACC1 trips, meter checking -24v suggests very fast trip
136139 | 42938.4848 | 1 | LI25 | 21 | 4 | HARDWARE | HW Problem | 25-2: Station keeps dropping out on external interlock. follow up: water trap shows many ACC1 trips, meter checking -24v suggests very fast trip
136139 | 42938.4848 |  | LI25 | 21 | 4 | HARDWARE | HW Problem | 25-2: Station keeps dropping out on external interlock. follow up: water trap shows many ACC1 trips, meter checking -24v suggests very fast trip
136140 | 42938.5049 | 1 | LI29 | 51 | 4 | HARDWARE | HW Problem | 29-5 has frequent Amplitude drop outs. Ranging does not cure. Res = 6.5 volts. Edit 7/22/2017- swing shift, 29-5 is now tripped and trips off immediately on overcurrent if we try to bring back on.
136140 | 42938.5049 | 3 | LI29 | 51 | 4 | HARDWARE | HW Problem | 29-5 has frequent Amplitude drop outs. Ranging does not cure. Res = 6.5 volts. Edit 7/22/2017- swing shift, 29-5 is now tripped and trips off immediately on overcurrent if we try to bring back on.
136140 | 42938.5049 | 1 | LI29 | 51 | 4 | HARDWARE | HW Problem | 29-5 has frequent Amplitude drop outs. Ranging does not cure. Res = 6.5 volts. Edit 7/22/2017- swing shift, 29-5 is now tripped and trips off immediately on overcurrent if we try to bring back on.
136001 | 42928.0527 | 1 | LI23 | 11 | 4 | HARDWARE | HW Problem | Klystron 23-1 having bouts of serious amplitude problems. Beam volts and phase appear fine but amplitude jumps up problematically.
136009 | 42929.1129 | 0.5 | LI24 | 61 | 4 | HARDWARE | HW Problem | Klystron 24-6 tripping frequently on temperature faults. KLT faults. Temperature readback seems to drift low. See ELOG entry: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=900623
136019 | 42929.8995 | 0.5 | LI20 | 61 | 4 | HARDWARE | HW Problem | station 20-6 (GUN) is tripped and showing VVS fault
135882 | 42917.2898 | 1 | LI30 | 41 | 4 | HARDWARE | HW Problem | 30-4: DQ'g current dropped to <1A.
135883 | 42917.5963 | 1 | LI21 | 11 | 4 | HARDWARE | HW Problem | LCLS INJ: 21-1 and 21-2 did not reset remotely after PPS fault.
135983 | 42927.2657 | 4 | LI28 | 41 | 4 | HARDWARE | HW Problem | Klys 28-4 tripped on modulator control power fault.
135992 | 42927.5898 | 0.5 | LI21 | 51 | 4 | HARDWARE | HW Problem | 21-5 is in MKSU Trigger Enable Fault and will not reset remotely.
134930 | 42866.6243 | 3 | LI20 | 71 | 4 | HARDWARE | HW Problem | Kly Mod 20-7, L0A: begin modulator startup. Process new E2V CX1937AX thyratron for HV ops.
134930 | 42866.6243 | 2.5 | LI20 | 71 | 4 | HARDWARE | HW Problem | Kly Mod 20-7, L0A: begin modulator startup. Process new E2V CX1937AX thyratron for HV ops.
134930 | 42866.6243 | 2.5 | LI20 | 71 | 4 | HARDWARE | HW Problem | Kly Mod 20-7, L0A: begin modulator startup. Process new E2V CX1937AX thyratron for HV ops.
136199 | 42942.6732 | 0.5 | LI25 | 71 | 4 | HARDWARE | HW Problem | 25-7 MOD fault, won't reset remotely.
136202 | 42943.2183 | 1 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4 has high BV jitter.
136203 | 42943.2187 | 1 | LI29 | 71 | 4 | HARDWARE | HW Problem | 29-7 has high amplitude jitter.
136208 | 42943.4083 | 0.5 | LI22 | 21 | 4 | HARDWARE | HW Problem | 22-2 Has an intermittent T-20 fan fault.
136209 | 42943.4131 | 1 | LI22 | 21 | 4 | HARDWARE | HW Problem | 22-2 Outboard filter cap looks like is bulging. May need replace.
136210 | 42943.4227 | 1 | LI25 | 21 | 4 | HARDWARE | HW Problem | Klystron 25-2 MOD faulted and won't reset remotely
136211 | 42943.4548 |  | LI25 | 21 | 4 | HARDWARE | HW Problem | 25-2 Trips on WG1.
136214 | 42943.599 | 1 | LI29 | 71 | 4 | HARDWARE | HW Problem | 29-7 Has a PFN fan fault.
136216 | 42943.7444 | 0.5 | LI26 | 61 | 4 | HARDWARE | HW Problem | Klystron 26-6 has a MOD VVS fault and cannot be reset remotely. Also, check the J-Box on top of the modulator.
136022 | 42929.9003 | 0.25 | LI20 | 61 | 4 | HARDWARE | HW Problem | LCLS gun klystron (20-6) has a VVS voltage fault after site-wide power dip.
136025 | 42929.9655 | 2 | LI20 | 61 | 4 | HARDWARE | HW Problem | 20-6: Gun beam voltage is extremely unstable after recovery from power dip. Edit 07:20 7/14/2017- The beam volts still become unstable; it seems to come in spurts. Station still needs attention.
136026 | 42929.9686 | 1 | LI20 | 81 | 4 | HARDWARE | HW Problem | L0B MOD faulted, won't reset.
136035 | 42930.5995 | 1.5 | LI23 | 21 | 4 | HARDWARE | HW Problem | 23-2: Thyratron nearing EOL.
136036 | 42930.6162 | 1 | LI29 | 41 | 4 | HARDWARE | HW Problem | 29-4: Keep alive power supply went from 280mA to 430mA after recent POMM work.
136038 | 42930.6798 |  | LI23 | 11 | 4 | HARDWARE | HW Problem | 23-1: Tune PFN for less slope and ripple.
136101 | 42935.7078 |  | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2 tripped on KMC after about 5 minutes of running.
136082 | 42935.0207 |  | LI25 | 21 | 4 | HARDWARE | HW Problem | Klystron Li25-21 is Mod faulted and will not reset remotely.
136091 | 42935.4787 | 0.5 | LI25 | 41 | 4 | HARDWARE | HW Problem | 25-4 station appears to be running fine but the MOD PLC EPICs PVs are not connecting.
136093 | 42935.5111 |  | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 frequent MOD PLC external interlock faults and PIOP MKSU trigger enable faults
134949 | 42867.1885 | 1.5 | LI22 | 51 | 4 | HARDWARE | HW Problem | Kly Mod 22-5 stuck in klystron heater delay for the past 18 hours.
134954 | 42867.2183 |  | LI22 | 41 | 4 | HARDWARE | HW Problem | 22-4 has low beam volts.
134956 | 42867.2188 |  | LI25 | 11 | 4 | HARDWARE | HW Problem | 25-1 has low beam volts.
134957 | 42867.219 |  | LI26 | 11 | 4 | HARDWARE | HW Problem | 26-1 has low beam volts.
136122 | 42937.263 | 1.5 | LI25 | 51 | 4 | HARDWARE | HW Problem | Many trips from 25-5 this shift. Nothing visible on the error counter.
134966 | 42867.2375 |  | LI21 | 11 | 4 | HARDWARE | HW Problem | 21-1 (L1S) only reaches 14 MV before hitting LLRF amplitude feedback limits (needs to get up to about 115 MV). Waveforms from the PAD look strange.
134967 | 42867.2851 | 0.5 | LI30 | 81 | 4 | HARDWARE | HW Problem | RF Station 30-8 not getting triggers--shows CAMAC error on CUD
134986 | 42867.511 | 2 | LI22 | 61 | 4 | HARDWARE | HW Problem | 22-6 has a tripped CB7 and Safety Relay & VVS faults indicated in plc fault log
135627 | 42900.2381 |  | LI22 | 11 | 4 | HARDWARE | HW Problem | 22-1 modulator panel shows a dequing under current fault.
135628 | 42900.2401 | 1 | LI26 | 51 | 4 | HARDWARE | HW Problem | 26-5 modulator panel shows a dequing under current fault.
135636 | 42900.4356 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | intermittent Amplitude spikes on L1X observed. Directly correlated with FEL intensity jitter. Not sure if HV or LLRF is the problem. see elog entry: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=896147 shows HV spikes correlate with bunch length spikes. tripping MPS pics.
135646 | 42900.7213 | 1.5 | LI24 | 61 | 4 | HARDWARE | HW Problem | 24-6 MKT fault.
136138 | 42938.3639 | 1.5 | LI30 | 11 | 4 | HARDWARE | HW Problem | 30-1: After HV regulator swap, DQ'g at 33A, amplitude and phase jitter very high.
136138 | 42938.3639 | 1 | LI30 | 11 | 4 | HARDWARE | HW Problem | 30-1: After HV regulator swap, DQ'g at 33A, amplitude and phase jitter very high.
136159 | 42941.5524 | 2 | LI28 | 81 | 4 | HARDWARE | HW Problem | 28-8 T20 input snubber assembly has several burnt wires that need to be replaced.
136160 | 42941.5573 | 2.2 | LI27 | 11 | 4 | HARDWARE | HW Problem | 27-1 Modulator problems found: 1. Primary snubber resistors: R260, R261 wires need to be replaced. 2. HV filter 4.7uF cap missing ground wires. 3. Core Bias TB1: pheonix connector (P1) Blk wire, the strands protruding from the connection. 4. Under thyratron at right side , the 115VAC cover came off and moved near the thyratron blower.
136161 | 42941.5581 | 1.5 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 Modulator EOLC stack needs to be replaced
136164 | 42941.5633 | 0.5 | LI23 | 31 | 4 | HARDWARE | HW Problem | 23-3: Door #3 ground hook wire too long.
136165 | 42941.5646 | 25 | LI27 | 31 | 4 | HARDWARE | HW Problem | 27-3 Modulator problems found: 1. T20: oil leak around the base and T-20 terminal (R3). 2. Primary snubber resistor R262 wire need to be replaced. 3. HV filter cap 4.7uF missing ground wires. 4. Core Bias: Pheonix connector (P2) orange clip that secured the connector is broken. 5. Thyratron compartment light out. 6. Keep alive wires hanging out. they need to be secured.
136167 | 42941.5651 | 2 | LI23 | 51 | 4 | HARDWARE | HW Problem | Grounding nut needed on main filter cap on the following stations
136167 | 42941.5651 | 3 | LI23 | 51 | 4 | HARDWARE | HW Problem | Grounding nut needed on main filter cap on the following stations
136167 | 42941.5651 | 3 | LI23 | 51 | 4 | HARDWARE | HW Problem | Grounding nut needed on main filter cap on the following stations
136167 | 42941.5651 | 2 | LI23 | 51 | 4 | HARDWARE | HW Problem | Grounding nut needed on main filter cap on the following stations
136167 | 42941.5651 | 2 | LI23 | 51 | 4 | HARDWARE | HW Problem | Grounding nut needed on main filter cap on the following stations
136168 | 42941.566 | 0.5 | LI27 | 41 | 4 | HARDWARE | HW Problem | 27-4 Modulator problem found: 1. Grounding nut needed on main filter cap.
136171 | 42941.5829 | 1.2 | LI27 | 71 | 4 | HARDWARE | HW Problem | 27-7 Modulator problems found: 1. Primary snubber caps: the jumpers on the caps should be replaced. 2. Core Bias: 120VAC Pheonix connector (P2) the orange clip is broken. 3. Under Thyratron compartment: the big white cable (H1) insulation cut open by the sharp metal edge.
136172 | 42941.5839 | 0.5 | LI27 | 81 | 4 | HARDWARE | HW Problem | 27-8 Modulator problem found: HV filter cap 4.7uF missing ground wires.
136172 | 42941.5839 | 1 | LI27 | 81 | 4 | HARDWARE | HW Problem | 27-8 Modulator problem found: HV filter cap 4.7uF missing ground wires.
136103 | 42935.8263 | 0.5 | LI25 | 21 | 4 | HARDWARE | HW Problem | Klystron LI25-21 is tripped on Mod Fault and will not reset remotely.
136128 | 42937.3665 | 0.5 | LI25 | 21 | 4 | HARDWARE | HW Problem | Klystron 25-2 MOD faulted and won't reset remotely
134987 | 42867.5125 | 1 | LI22 | 41 | 4 | HARDWARE | HW Problem | 22-4 is MOD faulted - indicates ACOC
134988 | 42867.5125 |  | LI22 | 41 | 4 | HARDWARE | HW Problem | 22-4 is MOD faulted - indicates ACOC
134989 | 42867.5139 | 1 | LI23 | 71 | 4 | HARDWARE | HW Problem | 23-7 is faulting on klystron O/C & O/V faults
134990 | 42867.5208 | 2 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4 tripped LVPS, F12 on relay board showing open,
135006 | 42868.1605 | 0.5 | LI20 | 51 | 4 | HARDWARE | HW Problem | station 20-5 PIOP looks bad
135651 | 42901.2013 |  | LI22 | 51 | 4 | HARDWARE | HW Problem | station 22-5 faulted HVR multiple times. We see 19 AMPL and 2 REFL faults on the error display since 08:00 yesterday.
135663 | 42901.952 | 2 | LI28 | 61 | 4 | HARDWARE | HW Problem | Klystron 28-6 MOD faulted and, when it is reset remotely, immediately trips off again.
135767 | 42910.5423 | 0.1 | LI25 | 21 | 4 | HARDWARE | HW Problem | 25-2 WGV faulted.
135767 | 42910.5423 | 2 | LI25 | 21 | 4 | HARDWARE | HW Problem | 25-2 WGV faulted.
135769 | 42910.9994 |  | LI20 | 81 | 4 | HARDWARE | HW Problem | 20-8 L0B trips on restart HVOV
135774 | 42912.3906 | 0.5 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 is getting OV faults. CB 7 tripped.
135781 | 42912.8061 | 1.3 | LI28 | 41 | 4 | HARDWARE | HW Problem | Klys 28-4 showing Beam Volts jitter.
136178 | 42941.7421 | 2 | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2 has a KMC fault that will not clear even though focus current is at 15A.
136192 | 42942.4755 | 1.5 | LI20 | 71 | 4 | HARDWARE | HW Problem | Klystron 20-7 (L0A) does not have PPS permit.
136192 | 42942.4755 | 0.2 | LI20 | 71 | 4 | HARDWARE | HW Problem | Klystron 20-7 (L0A) does not have PPS permit.
135020 | 42868.5661 | 0.5 | LI27 | 51 | 4 | HARDWARE | HW Problem | 25-7 Klystron magnet power supply has no output.
135022 | 42868.5965 | 3 | LI24 | 61 | 4 | HARDWARE | HW Problem | Station 24-6 not staying on. Frequently HVR faults.
135022 | 42868.5965 | 1 | LI24 | 61 | 4 | HARDWARE | HW Problem | Station 24-6 not staying on. Frequently HVR faults.
135023 | 42868.6494 | 2 | LI24 | 11 | 4 | HARDWARE | HW Problem | Klys 24-1 is MOD; no remote reset
135674 | 42902.6686 | 3.5 | LI20 | 81 | 4 | HARDWARE | HW Problem | Klystron 20-8 (L0B) MOD faulted and won't reset remotely
135675 | 42903.1114 | 2.2 | LI30 | 61 | 4 | HARDWARE | HW Problem | Klys 30-6 keeps going HVR
135677 | 42904.4689 | 1 | LI22 | 41 | 4 | HARDWARE | HW Problem | Klys 22-4 is tripping frequently on over voltage.
135684 | 42904.6766 | 0 | LI20 | 81 | 4 | HARDWARE | HW Problem | L0B's modulator interlocks will sometimes only reset via the PPS reset button, even when the interlock is not PPS related. See attachment: 28-2,20-5,6,7,8,21-1,21-2,b921.pdf and HV_boards_upgrade.xlsx for list of modulators upgraded.
135684 | 42904.6766 |  | LI20 | 81 | 4 | HARDWARE | HW Problem | L0B's modulator interlocks will sometimes only reset via the PPS reset button, even when the interlock is not PPS related. See attachment: 28-2,20-5,6,7,8,21-1,21-2,b921.pdf and HV_boards_upgrade.xlsx for list of modulators upgraded.
135689 | 42905.3675 |  | LI21 | 21 | 4 | HARDWARE | HW Problem | L1X amplitude jitter high; varying significantly shot to shot. Don't see correlation to beam volts, do see correlation to drive.
135787 | 42913.4043 | 1.3 | LI27 | 81 | 4 | HARDWARE | HW Problem | 27-8 has high BV jitter.
135788 | 42913.4055 | 1.3 | LI26 | 11 | 4 | HARDWARE | HW Problem | 26-1 has high BV jitter.
135789 | 42913.407 | 0.1 | LI26 | 81 | 4 | HARDWARE | HW Problem | 26-8 fast time plots aren't working. Tried plotting amplitude & BV jitter and got unrealistic-looking plots that didn't change from one plot request to the next. I don't think fast time plotting issues reflect a modulator problem. p.stiles
135789 | 42913.407 | 1 | LI26 | 81 | 4 | HARDWARE | HW Problem | 26-8 fast time plots aren't working. Tried plotting amplitude & BV jitter and got unrealistic-looking plots that didn't change from one plot request to the next. I don't think fast time plotting issues reflect a modulator problem. p.stiles
135790 | 42913.417 | 1.5 | LI30 | 31 | 4 | HARDWARE | HW Problem | 30-3 Has intermittent BV dropouts.
135244 | 42878.5984 | 2 | LI27 | 71 | 4 | HARDWARE | HW Problem | 27-7 drops to HVR / loses HV CMD signal to regulator board with no faults showing.
135381 | 42885.8519 | 1.5 | LI21 | 11 | 4 | HARDWARE | HW Problem | L1S beam volts jitter causing 20% jitter on FEL. Correlated well with phase jitter
135398 | 42886.5579 | 2 | LI27 | 41 | 4 | HARDWARE | HW Problem | Station trips once on HVOC and it goes into Mod Fault lock-out
134358 | 42814.2781 | 1.2 | LI21 | 41 | 4 | HARDWARE | HW Problem | 21-4: Core Bias Assembly, 120Vac "P2" secure clip broken.
135826 | 42915.4376 | 1 | LI27 | 61 | 4 | HARDWARE | HW Problem | 27-6 has intermittent large BV dropouts.
135827 | 42915.4392 | 1.3 | LI27 | 11 | 4 | HARDWARE | HW Problem | 27-1 has high BV jitter.
135828 | 42915.4402 | 1.3 | LI26 | 11 | 4 | HARDWARE | HW Problem | 26-1 has very high phase jitter.
135829 | 42915.4412 | 1.3 | LI29 | 51 | 4 | HARDWARE | HW Problem | 29-5 has high phase jitter.
134574 | 42842.5815 |  | LI25 | 41 | 4 | HARDWARE | HW Problem | 25-4 Modulator's PLC will not remember settings. Suspect PLC processor module. Replace if needed.
134575 | 42842.5872 | 1 | LI25 | 81 | 4 | HARDWARE | HW Problem | 25-8 Modulator Fire Wire Fault. Replace faulty fire wire.
134578 | 42842.6044 | 1 | LI30 | 61 | 4 | HARDWARE | HW Problem | 30-6 modulator: Does not show an External Interlock. It should because MKSU is showing a KMC fault.
135438 | 42888.3459 | 3 | LI20 | 81 | 4 | HARDWARE | HW Problem | Klystron 20-8 MOD faulted and won't reset remotely.
134690 | 42857.3447 | 1 | LI29 | 71 | 4 | HARDWARE | HW Problem | 29-7 mod. has a PFN rooftop fan fault.
135832 | 42915.444 | 1.3 | LI22 | 81 | 4 | HARDWARE | HW Problem | 22-8 has intermittent single shot phase jumps.
135834 | 42915.546 | 1.3 | LI28 | 31 | 4 | HARDWARE | HW Problem | Klystron 28-3 MOD faulted with no remote reset.
135834 | 42915.546 | 1.3 | LI28 | 31 | 4 | HARDWARE | HW Problem | Klystron 28-3 MOD faulted with no remote reset.
135843 | 42916.1193 | 0.1 | LI26 | 31 | 4 | HARDWARE | HW Problem | 26-3 has a Klystron Temperature fault.
135465 | 42889.9157 | 2 | LI26 | 41 | 4 | HARDWARE | HW Problem | 26-4 tripping off frequently.
135465 | 42889.9157 | 0.5 | LI26 | 41 | 4 | HARDWARE | HW Problem | 26-4 tripping off frequently.
135470 | 42890.2514 | 2 | LI26 | 41 | 4 | HARDWARE | HW Problem | Klystron LI26-41 is tripping off on EVOC faults.
135470 | 42890.2514 |  | LI26 | 41 | 4 | HARDWARE | HW Problem | Klystron LI26-41 is tripping off on EVOC faults.
135473 | 42890.5959 | 0.5 | LI28 | 71 | 4 | HARDWARE | HW Problem | station 28-7 Mod faulted-won't stay on after reset.
134742 | 42858.6166 | 1.5 | LI30 | 81 | 4 | HARDWARE | HW Problem | 30-8: DQ'g divider has slow oil leak.
135211 | 42877.4162 | 1 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4 is faulted on KMC
136194 | 42942.5822 | 1.5 | LI30 | 81 | 4 | HARDWARE | HW Problem | 30-8 Modulator shows HV but no current. Triggers are present.
135474 | 42890.6 | 1 | LI26 | 61 | 4 | HARDWARE | HW Problem | 26-6 Core Bias supply unstable, 15.0 amps +/- 0.1 amp at 5.65 volts
135481 | 42891.2527 | 1 | LI22 | 31 | 4 | HARDWARE | HW Problem | 22-3 is MOD fault and won't reset
135482 | 42891.2558 | 1 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4 is MOD fault and won't reset
135488 | 42891.4621 | 2 | LI21 | 21 | 4 | HARDWARE | HW Problem | 21-2 no focusing current read back on EPICS. SAM Module MODU:LI21:208-LI21 Cr 2 M08 - S21-F2A (21-2) shows zeros on EPICS. OK LED not flashing on the module. Suspect SAM module.
135491 | 42891.5845 | 1 | LI26 | 31 | 4 | HARDWARE | HW Problem | Klystron 26-3 is fault lockout. Cannot reset remotely.
135497 | 42892.0499 | 1 | LI23 | 41 | 4 | HARDWARE | HW Problem | Klystron LI23-41 has high amplitude jitter.
135047 | 42869.5956 |  | LI23 | 71 | 4 | HARDWARE | HW Problem | Klystron 23-7 tripping off. Overvoltage errors, trips immediately when mod is on.
135077 | 42871.6262 | 3 | LI20 | 81 | 4 | HARDWARE | HW Problem | Kly Mod 20-8 (L0B) has excessive phase and amplitude jitter.
135086 | 42872.3741 |  | LI26 | 61 | 4 | HARDWARE | HW Problem | No 110vac control power to modulator 26-6 although feeder breaker at panel 2P26G1B - #2 is on, not tripped. Cycling breaker #2 has no effect. Other circuits in panel are ok.
135917 | 42921.9586 | 0.5 | LI28 | 41 | 4 | HARDWARE | HW Problem | 28-4: Shows AMJ. causing 30% jitter on FEL intensity. DEACT'd, pulse stability improved immediately.
134689 | 42857.3425 | 1 | LI29 | 51 | 4 | HARDWARE | HW Problem | 29-5 Mod. has a Thy Fan Flt.
135511 | 42892.9219 | 0 | LI21 | 51 | 4 | HARDWARE | HW Problem | 21-5 having multiple Mod faults (water faults)
135640 | 42900.6086 |  | LI20 | 81 | 4 | HARDWARE | HW Problem | L0B tripped, required PEM to lower res voltage to make stable. (-.5H)
135096 | 42872.4702 | 1 | LI29 | 71 | 4 | HARDWARE | HW Problem | 29-7 tripping on thyratron filament O/C
135104 | 42872.5929 |  | LI27 | 51 | 4 | HARDWARE | HW Problem | 27-5 drops to HVR / loses HV CMD signal to regulator board with no faults showing.
135106 | 42872.6043 | 2 | LI30 | 61 | 4 | HARDWARE | HW Problem | 30-6 KMC
135108 | 42872.6196 |  | LI25 | 41 | 4 | HARDWARE | HW Problem | TBR 25-4 Suspected klystron magnet short.
135933 | 42922.7457 | 1.5 | LI26 | 61 | 4 | HARDWARE | HW Problem | 26-6 tripped on MOD fault; can't reset.
135933 | 42922.7457 | 2 | LI26 | 61 | 4 | HARDWARE | HW Problem | 26-6 tripped on MOD fault; can't reset.
134705 | 42857.4297 | 1 | LI25 | 41 | 4 | HARDWARE | HW Problem | 25-4 modulator PLC has zeros for all input values
134708 | 42857.4545 | 1 | LI20 | 71 | 4 | HARDWARE | HW Problem | 20-7 Trim P/S Genesys needs replace. No output.
134709 | 42857.4562 | 1 | LI22 | 41 | 4 | HARDWARE | HW Problem | 22-4 Trim P/S Display is flashing. Touch panel shows no output from unit.
134710 | 42857.4573 | 1 | LI22 | 71 | 4 | HARDWARE | HW Problem | 22-7 Trim P/S Display is flashing. Touch panel shows no output from unit.
136899 | 43006.914 | 0.5 | LI23 | 11 | 4 | HARDWARE | HW Problem | 23-1 poor focus regulation see ACR ELOG entry: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=911103
136684 | 42991.5834 |  | LI20 | 81 | 4 | HARDWARE | HW Problem | RFS 20-8 L0B after 9/13/17 POMM 80kV thyratron replacement, we are seeing frequent instances of no HV for one pulse then slow (several pulses) recovery to normal. No interlock faults seen.
136685 | 42991.5847 | 1 | LI25 | 71 | 4 | HARDWARE | HW Problem | 25-7: Thyratron nearing EOL.
136692 | 42991.9797 |  | LI20 | 81 | 4 | HARDWARE | HW Problem | The symptom observed is a non-latching MPS fault. This fault was also observed when TD-11 is in. It is suspected that these faults are due to a combination of L0B, and a yet to be identified L2 station. Another issue is that the MPS fault does not persist for very long, and is not logged in the MPS fault log in the MPS GUI.
136766 | 42998.5178 |  | LI30 | 31 | 4 | HARDWARE | HW Problem | 30-3 WGV fault.
136804 | 43000.3403 | 1.5 | LI23 | 81 | 4 | HARDWARE | HW Problem | 23-8 has been in KHD for over an hour.
136679 | 42991.4855 |  | LI23 | 51 | 4 | HARDWARE | HW Problem | 23-5 FLK attempts to reset it trips again on External Interlock lockout. I saw some PIOP messages during turn on attempt indicating klys over voltage.
136700 | 42992.4446 |  | LI20 | 81 | 4 | HARDWARE | HW Problem | L0B showing frequent BVLT dropouts
136716 | 42993.4118 |  | LI20 | 81 | 4 | HARDWARE | HW Problem | LOB (20-8) repeatedly tipping off on ACOC faults. Plan per PD and Axel- swap thyratron during PAMM next week, run with lower amplitude until then.
136729 | 42995.0826 |  | LI22 | 51 | 4 | HARDWARE | HW Problem | 22-5 is experiencing frequent BV dropouts.
136729 | 42995.0826 | 0.5 | LI22 | 51 | 4 | HARDWARE | HW Problem | 22-5 is experiencing frequent BV dropouts.
136813 | 43001.3332 | 3 | LI24 | 21 | 4 | HARDWARE | HW Problem | 24-2: MOD faulted and will not reset remotely. Station getting no triggers. That problem fixed, still has instabilities.
136815 | 43001.5496 | 2 | LI20 | 81 | 4 | HARDWARE | HW Problem | Kly Mod 20-8 (L0B): station again experiencing BV drop outs.
136818 | 43002.5321 |  | LI23 | 81 | 4 | HARDWARE | HW Problem | station 27-8 had at least 8 AMPL faults in the last 15 minutes.
136820 | 43002.5872 | 4 | LI30 | 61 | 4 | HARDWARE | HW Problem | station 30-6 MOD fault--won't reset remotely
136821 | 43002.5877 | 2 | LI27 | 81 | 4 | HARDWARE | HW Problem | 27-8 ACOC interlocks.
136821 | 43002.5877 | 0.8 | LI27 | 81 | 4 | HARDWARE | HW Problem | 27-8 ACOC interlocks.
136825 | 43003.2798 | 1 | LI21 | 81 | 4 | HARDWARE | HW Problem | 21-8 MOD fault, no reset
136798 | 42999.7505 | 3 | LI26 | 81 | 4 | HARDWARE | HW Problem | Klystron LI26-81 goes Mod Fault immediately after resetting.
136838 | 43003.6947 | 0.5 | LI27 | 61 | 4 | HARDWARE | HW Problem | 27-6 reporting rapid trips on WA#1 flowswitch
136831 | 43003.5383 | 2 | LI24 | 21 | 4 | HARDWARE | HW Problem | 24-2 at 300kv, DQing current = 265 amps indicated on PLC
136839 | 43003.749 | 0.5 | LI27 | 61 | 4 | HARDWARE | HW Problem | station 27-6 shows WA1 faults. Error display shows 1640 water faults since 08:00 this morning.
136842 | 43004.1097 | 0.5 | LI28 | 31 | 4 | HARDWARE | HW Problem | Klys 28-3 is MOD; no remote reset. Ops tried resetting locally, but we know less than PEM about how to do that, so please draw no conclusions from us :)
136843 | 43004.178 | 0.5 | LI22 | 81 | 4 | HARDWARE | HW Problem | Klystron 22-8 experienced quite a few trips in a very short time span and then wouldn't reset
136843 | 43004.178 | 0.5 | LI22 | 81 | 4 | HARDWARE | HW Problem | Klystron 22-8 experienced quite a few trips in a very short time span and then wouldn't reset
136844 | 43004.2666 | 0.5 | LI24 | 21 | 4 | HARDWARE | HW Problem | 24-2 had the MKSU changed and also the Beam voltage cable changed. Need to do PIOP ratios and change cable if necessary.
136859 | 43005.0764 | 2 | LI28 | 31 | 4 | HARDWARE | HW Problem | Klys 28-3 is MOD with no remote reset
136860 | 43005.1185 | 1 | LI22 | 81 | 4 | HARDWARE | HW Problem | Klystron LI22-81 is Mod Faulted and won't reset remotely. Klystron had been dropping out sporadically before faulting off. Possible water problem; thought water summary fault flashed before tripping. Update: water kludge added and indicated an ACC1 fault - need flow / flow switch checked
136862 | 43005.3808 | 0 | LI27 | 81 | 4 | HARDWARE | HW Problem | Modulator 27-8 MOD faults frequently. We are able to reset it, but it trips off again less than a minute later.
136878 | 43006.2904 | 0 | LI30 | 61 | 4 | HARDWARE | HW Problem | Klystron LI30-61 has tripped off 6 times over the last shift and is now not able to reset remotely. Errors panel does not show reason for faulting.
136741 | 42997.3311 | 0.5 | LI22 | 51 | 4 | HARDWARE | HW Problem | 22-5 PHJ due to near zero crossing on 360 scan
136751 | 42998.3195 |  | LI26 | 21 | 4 | HARDWARE | HW Problem | 26-2 CAMAC blower chassis fan is noisy.
136754 | 42998.4505 | 0.5 | LI24 | 11 | 4 | HARDWARE | HW Problem | 24-1: SCR snubbers have open resistors on phase A and phase C.
136755 | 42998.4524 | 0.5 | LI24 | 21 | 4 | HARDWARE | HW Problem | 24-2: SCR snubbers have open resistor on phase A and blown resistor on phase C.
136756 | 42998.4541 | 0.8 | LI24 | 31 | 4 | HARDWARE | HW Problem | 24-3: SCR snubbers have open resistors on phase B and phase C, and phase A board burnt.
136730 | 42995.8592 | 0.5 | LI28 | 21 | 4 | HARDWARE | HW Problem | 28-2 has beam voltage jitter.
136759 | 42998.4589 | 1 | LI24 | 81 | 4 | HARDWARE | HW Problem | 24-8: SCR snubbers have open resistors on phase B and phase C.
136767 | 42998.5723 | 3.5 | LI22 | 11 | 4 | HARDWARE | HW Problem | 22-1 Modultor problems found: 1. Primary snubber resistors: R260, R262 lugs need to be replaced. A 2. Current transducer case of Core Bias Assembly is burnt. 3. AC wires and thyratron cables should be separated. Wires insulation burnt. 4. Cover for 120VAC should be secured with tie-wrap.
136768 | 42998.5871 | 2.5 | LI22 | 21 | 4 | HARDWARE | HW Problem | 22-2 Modulator problems found: 1. Ground hook holder on the PFN door need to be replaced. 2. EOLC stack is missing the stand-off (see attached picture). 3. All ground hooks in this modulator, screw that hold cable and ground hook together need a lock washer to keep it tighten.
136787 | 42999.3632 | 1 | LI30 | 81 | 4 | HARDWARE | HW Problem | 30-8: MOD faults, ACOC's, no reset.
136788 | 42999.3858 | 1 | LI23 | 81 | 4 | HARDWARE | HW Problem | 23-8: Modulator still going into KHD. Klystron voltage dropping out, current staying steady.
136624 | 42988.5589 | 2 | LI20 | 81 | 4 | HARDWARE | HW Problem | Klys 20-8 (L0B) showing beam volts dropouts. See 136243.
136662 | 42991.122 |  | LI21 | 11 | 4 | HARDWARE | HW Problem | Klystron 21-1 FLK and won't reset remotely
136667 | 42991.2309 | 1 | LI20 | 71 | 4 | HARDWARE | HW Problem | Mod 20-7 KLT ( klystron delta temperature) fault
136668 | 42991.2322 |  | LI22 | 81 | 4 | HARDWARE | HW Problem | 22-8 waveguide vacuum fault
136669 | 42991.2342 | 1 | LI23 | 31 | 4 | HARDWARE | HW Problem | Modulator 23-3 has ACOC faults even when modulator high voltage is lowered to 1 kV or modulator triggers are removed.
136671 | 42991.2536 | 1 | LI22 | 81 | 4 | HARDWARE | HW Problem | 22-8 tripped on WGV, wont reset.
136672 | 42991.2627 |  | LI24 | 31 | 4 | HARDWARE | HW Problem | Klystron 24-3 MOD faulting frequently
136911 | 43007.7514 | 4 | LI28 | 31 | 4 | HARDWARE | HW Problem | Modulator 28-3: from ACR it appears CB7 has again tripped.
136913 | 43008.4548 | 4 | LI25 | 51 | 4 | HARDWARE | HW Problem | Klystron LI25-51 is tripped on "Mod fault" and will not reset remotely.
136640 | 42990.6155 | 2 | LI20 | 81 | 4 | HARDWARE | HW Problem | 20-8 L0B several (3) HVOV trips in last half hour.
136701 | 42992.4852 | 3 | LI28 | 31 | 4 | HARDWARE | HW Problem | 28-3 MOD fault with no remote reset.
136710 | 42992.9849 | 3 | LI23 | 81 | 4 | HARDWARE | HW Problem | Klys 23-8 is on mystery KHD
136727 | 42994.3576 | 1 | LI30 | 71 | 4 | HARDWARE | HW Problem | 30-7: MOD fault, tripping off randomly.
136727 | 42994.3576 | 0.5 | LI30 | 71 | 4 | HARDWARE | HW Problem | 30-7: MOD fault, tripping off randomly.
136801 | 43000.2763 | 0.5 | LI26 | 31 | 4 | HARDWARE | HW Problem | 26-3 has a KLT fault.
136802 | 43000.2774 | 0.2 | LI25 | 71 | 4 | HARDWARE | HW Problem | 25-7 is FLK.
136803 | 43000.3202 | 1 | LI21 | 81 | 4 | HARDWARE | HW Problem | Klystron 21-8 was experiencing frequent amplitude dropouts
136808 | 43000.5005 | 2 | LI20 | 81 | 4 | HARDWARE | HW Problem | Klystron 20-8 frequent beam volt dropouts. See 136814.