## Sheet1
PROB_ID | CREATED_DATE | SOLVE_HOURS | MICRO | UNIT | STATUS_CHK | DESCRIPTION | PROB_TYPE_CHK
133417 | 2016-10-18 | 0.75 | LI23 | 41 | 4 | Klys 23-4 has AMJ and PHJ problems that are pretty visible on the beam | HARDWARE
133500 | 2016-10-28 | 1.4 | LI24 | 51 | 4 | 24-5 has bad phase jitter. | HARDWARE
133548 | 2016-11-02 | 0.75 | LI22 | 71 | 4 | station 22-7 has growing BeamVolts jitter | HARDWARE
133653 | 2016-11-10 | 2.5 | LI24 | 81 | 4 | 24-8 install and runup new deuterium thyratron from spares inventory, update depot | HARDWARE
133800 | 2016-11-28 | 1 | LI24 | 21 | 4 | 24-2 is MOD faulted. Touch panel has a "Host Communication err" popup and is unresponsive. | HARDWARE
133804 | 2016-11-28 |  | LI20 | 81 | 4 | 20-8 (L0B) Core bias power supply experienced current noise after Nov 13th. Open the cabinet and inspect possible noise source (arcing somewhere) including pulse cable. | HARDWARE
133806 | 2016-11-28 | 2.5 | LI20 | 71 | 4 | Klystron 20-7 trips incessantly on MOD fault lockout. We can reset it temporarily, but it trips off again a minute later. | HARDWARE
133850 | 2016-11-30 | 1 | LI22 | 61 | 4 | 22-6 Showing frequent bouts of intense phase jitter. Virtually unusable. | HARDWARE
133204 | 2016-10-01 | 1 | LI23 | 21 | 4 | klys 23-2 tripped, no reset. | HARDWARE
133419 | 2016-10-19 |  | LI22 | 71 | 4 | 22-7 exhibiting single shot phase excursions. | HARDWARE
133513 | 2016-10-28 | 1 | LI24 | 51 | 4 | 24-5 ACOC faults on turn on. | HARDWARE
133513 | 2016-10-28 | 0.2 | LI24 | 51 | 4 | 24-5 ACOC faults on turn on. | HARDWARE
133672 | 2016-11-11 |  | LI29 | 71 | 4 | Klys 29-7 suddenly showing frequent large beam volts dropout. Wouldn't be surprised if it goes away on its own at some point as it started very suddenly. | HARDWARE
133719 | 2016-11-17 | 1 | LI27 | 41 | 4 | Klystron 27-4 and 29-1 MOD faulted and won't reset remotely | HARDWARE
133175 | 2016-09-29 |  | LI22 | 71 | 4 | 22-7 BV FTPs not working (show 0 BV, station is at nominal voltage). | HARDWARE
133131 | 2016-09-26 | 1 | LI28 | 21 | 4 | Klys 28-2 FLK, no reset. | HARDWARE
133173 | 2016-09-29 | 1.2 | LI23 | 41 | 4 | 23-4 has high beam volts jitter. | HARDWARE
133179 | 2016-09-29 |  | LI24 | 31 | 4 | Klys 24-3 has been tripping frequently on this shift | HARDWARE
133185 | 2016-09-30 |  | LI22 | 71 | 4 | 22-7 FTPs for BV and BI show enormous (false) shot to shot jitter. | HARDWARE
133211 | 2016-10-02 | 0.5 | LI24 | 61 | 4 | 24-6 tripping to HVR | HARDWARE
133094 | 2016-09-23 |  | LI26 | 21 | 4 | 26-2 showing beam volt jitter and amplitude jitter on FTPs, affects beam quality when it's on the beam. | HARDWARE
133205 | 2016-10-02 | 0.5 | LI23 | 21 | 4 | Kly. LI 23 - 2 tripped on 1 hour Klystron heater delay | HARDWARE
133778 | 2016-11-23 |  | LI22 | 71 | 4 | Duplicate of 132111 - KLYS:LI22:71:WNDW1_TEMP Thermal couple looks to be broken. Read back is -10982.3 | HARDWARE
133781 | 2016-11-23 | 0.5 | LI28 | 31 | 4 | 28-3 is tripping on klystron water faults. | HARDWARE
133789 | 2016-11-24 | 0.75 | LI25 | 31 | 4 | Kly. LI 25 - 3 trips off when I tried to turn it on. | HARDWARE
133794 | 2016-11-26 | 1 | LI23 | 31 | 4 | 23-3 has tripped off at least three times on owl shift; EDM Error panel shows 6 'CURR' faults. | HARDWARE
133794 | 2016-11-26 | 3 | LI23 | 31 | 4 | 23-3 has tripped off at least three times on owl shift; EDM Error panel shows 6 'CURR' faults. | HARDWARE
133220 | 2016-10-04 | 1 | LI24 | 31 | 4 | 24-3 HVOC and ACOC. tripped off countless times during swing shift and continues to trip off on HVR. Resets each time but then trips again later. | HARDWARE
133266 | 2016-10-07 | 0.1 | LI26 | 31 | 4 | LI 26 Kly. - 3 tripped on WGV. informed PEM | HARDWARE
133348 | 2016-10-13 | 1.5 | LI23 | 11 | 4 | 23-1 HVOC's possibly caused by PFN capacitor going bad | HARDWARE
133076 | 2016-09-20 | 2 | LI21 | 11 | 4 | 21-1 occasional faults for SCR pan fan | HARDWARE
133206 | 2016-10-02 | 2 | LI21 | 11 | 4 | Klys. LI 21 - 1 tripping off mult. times. PIOP message show Focus P.S. faults. PEM suspects MSKU unit maybe at fault. | HARDWARE
133943 | 2016-12-08 | 1.2 | LI27 | 81 | 4 | station 27-8 has High Beam Volts Jitter --see mccElog entry today at 10:00 for pictures | HARDWARE
133366 | 2016-10-14 | 6 | LI21 | 21 | 4 | Time Accounting: Had to replace cable for L1X jitter. | HARDWARE
133394 | 2016-10-17 | 4 | LI24 | 31 | 4 | Klys 24-3 cycling freqently, error count worthless. | HARDWARE
133407 | 2016-10-18 | 1.5 | LI21 | 41 | 4 | 21-4 is having lots of BVLT drop outs. | HARDWARE
133422 | 2016-10-19 | 0.5 | LI25 | 11 | 4 | 25-1 is MOD faulted, trips off immediately when it is turned on. | HARDWARE
133429 | 2016-10-19 | 1.5 | LI22 | 71 | 4 | 22-7 Beam Current and Beam voltage fast time plots are in-operative, suspect bad PIOP. This problem began after PIOP swap in August | HARDWARE
134154 | 2017-02-02 |  | LI26 | 51 | 4 | 26-5: Dead man switch failure at door 2. | HARDWARE
134164 | 2017-02-05 | 0.5 | LI21 | 21 | 4 | 21-2: Remove tail clipper assembly. | HARDWARE
134165 | 2017-02-05 | 0.8 | LI21 | 31 | 4 | 21-3: Remove tail clipper assembly. | HARDWARE
133200 | 2016-10-01 | 2 | LI29 | 11 | 4 | Kly LI 29 - 1 tripped off. Now on 1 hour heater delay. | HARDWARE
133230 | 2016-10-05 |  | LI23 | 31 | 4 | Klys 23-3 lost AMPL/PHAS readback. | HARDWARE
133249 | 2016-10-06 | 1 | LI24 | 31 | 4 | 24-3 focus monitor current varying +/- 1 amp, noisy on history plot | HARDWARE
133292 | 2016-10-09 |  | LI23 | 51 | 4 | Quad LI24 713 (QM21) won't trim after power dip. | HARDWARE
133309 | 2016-10-11 |  | LI21 | 41 | 4 | 21-4 amplitude dropped out momentarily about half a dozen times on 10/11 owl shift. Does not register on the sector error counts and it appears that the beam volts don't drop during these temporary excursions. | HARDWARE
133318 | 2016-10-11 | 2 | LI26 | 61 | 4 | Kly. LI 26 - 6 tripped the VVS breaker. | HARDWARE
133439 | 2016-10-20 | 1 | LI26 | 61 | 4 | 26-6 MOD/HVR faults, no remote reset. | HARDWARE
133494 | 2016-10-27 |  | LI25 |  | 4 | Sector 25, Need to run POLYs on these Stations: 25-5,6,8 before running a Poly on the SBST. Put into CKP when complete. | HARDWARE
133499 | 2016-10-28 | 1.2 | LI29 | 51 | 4 | 29-5 has bad phase and amplitude jitter. | HARDWARE
133120 | 2016-09-26 |  | LI28 | 31 | 4 | 28-3 seeing WKL FS trips when LI29 KCS LCW is taxed by additional heat load (as in when SBST 28 trips off). | HARDWARE
133387 | 2016-10-17 | 1 | LI27 | 41 | 4 | 27-4 has had excessive HVOC trips | HARDWARE
134287 | 2017-03-02 | 1 | LI28 | 31 | 4 | 28-3 has water leaking out the magnet weep hole | HARDWARE
133337 | 2016-10-12 | 1 | LI22 | 61 | 4 | Kly LI 22 - 6 tripped on "MOD" fault and won't reset. | HARDWARE
133339 | 2016-10-12 | 1 | LI20 | 81 | 4 | 20-8 has Keep Alive undercurrent fault. Can't bring it on remotely. | HARDWARE
133526 | 2016-10-30 | 3 | LI23 | 31 | 4 | Klys 23-3 trips frequently. | HARDWARE
133420 | 2016-10-19 | 0.75 | LI22 | 51 | 4 | 22-5 has relatively high beam volts jitter, which is in turn causing phase jitter. | HARDWARE
133896 | 2016-12-04 | 1 | LI23 | 41 | 4 | 23-4 showing lots of phase jitter causing FEL spikes down | HARDWARE
133897 | 2016-12-04 | 1 | LI23 | 51 | 4 | 23-5 showing lots of phase jitter causing FEL spikes down | HARDWARE
133899 | 2016-12-05 | 0 | LI24 | 31 | 4 | 24-3 has occasional BV dropouts, bad amplitude jitter. | HARDWARE
133900 | 2016-12-05 | 1.2 | LI23 | 31 | 4 | 23-3 has high amplitude jitter. | HARDWARE
133601 | 2016-11-07 | 1 | LI21 | 81 | 4 | station 21-8 high Amplitude Jitter | HARDWARE
133601 | 2016-11-07 |  | LI21 | 81 | 4 | station 21-8 high Amplitude Jitter | HARDWARE
133601 | 2016-11-07 | 1 | LI21 | 81 | 4 | station 21-8 high Amplitude Jitter | HARDWARE
133649 | 2016-11-09 |  | LI23 | 71 | 4 | KLYS:LI23:71:Phase Jitter is 1.5 %. Was 1.0% before the PAMM. | HARDWARE
133656 | 2016-11-10 |  | LI26 | 31 | 4 | 26-3 WGV fault. No observable vacuum issues, probably just needs gauge juiced. | HARDWARE
133670 | 2016-11-11 |  | LI24 | 31 | 4 | 24--3 shows no amplitude. It is not on the beam, but it is behaving exactly as 24-1 did earlier (see CATER 133667). Initially there was zero amplitude shown, tried to IPL PIOP, after the station has zero amplitude but this does not propagate an error up to the klystron CUD, and there is no phase readback. | HARDWARE
133053 | 2016-09-18 | 1 | LI29 | 51 | 4 | 29-5 has high phase jitter. | HARDWARE
133260 | 2016-10-07 | 1 | LI26 | 11 | 4 | Klystron 26-1 HVR and won't reset remotely. It comes on briefly, but trips off again afterwards. | HARDWARE
133388 | 2016-10-17 | 1 | LI24 | 51 | 4 | 24-5 The PFN need tuning, slope =-0.6 | HARDWARE
133996 | 2016-12-13 | 1 | LI23 | 41 | 4 | 23-4 has large Beam Volts jitter | HARDWARE
133674 | 2016-11-12 | 1.4 | LI29 | 41 | 4 | 29-1 WGV. Locally vacuum gauge controller is in "Protect" mode reporting to control system a non-plausible 2e-4 torr. | HARDWARE
133716 | 2016-11-16 |  | LI20 | 81 | 4 | Klystron Li20-81 (L0-B) shows intermittent Mod Fault but klystron is still functioning. Bad status. | HARDWARE
133725 | 2016-11-17 |  | LI25 | 81 | 4 | 25-8 CBL fault; resetting/IPLing PIOP doesn't help. | HARDWARE
133728 | 2016-11-18 |  | LI27 | 31 | 4 | Kly. LI 27 - 3 tripped "MOD' fault. Tried to reset 3 times but faulted each time. | HARDWARE
133939 | 2016-12-07 |  | LI30 | 71 | 4 | Kly. LI 30 - 7 BVLT. came back 10 volts lower than when the unit went down on Mon. | HARDWARE
133770 | 2016-11-23 | 1 | LI23 | 81 | 4 | Klys 23-8 is cycling/tripping frequently. Error count only shows AMPL errors. | HARDWARE
133922 | 2016-12-05 | 0.5 | LI24 | 21 | 4 | 24-2 having amplitude and phase dropouts, enough to trip off the beam often. Put to standby. | HARDWARE
133928 | 2016-12-06 | 0.2 | LI22 | 51 | 4 | Klys 22-5 is WGV | HARDWARE
133928 | 2016-12-06 | 0.2 | LI22 | 51 | 4 | Klys 22-5 is WGV | HARDWARE
133984 | 2016-12-11 |  | LI30 | 71 | 4 | 30-7 BVLT. came back at 337. Was at 347 before last down time. | HARDWARE
134234 | 2017-02-21 | 2 | LI22 | 81 | 4 | 22-8: MKSU problem causing PPS red lights to flash in sector 22. | HARDWARE
134271 | 2017-02-26 | 0.5 | LI27 | 41 | 4 | 27-4: Anode reactor broken on PFN. | HARDWARE
134272 | 2017-02-26 | 1.2 | LI21 | 41 | 4 | 21-4: TB7 in thyratron compartment broken. | HARDWARE
134138 | 2017-01-28 | 2 | LI30 | 11 | 4 | Kly Mod 30-1: PFN cabinet ground hook broken. PFN capacitor #4 of upper string (cap is upper right hand in upper string or string #1) leaking from bushing cap. Replace capacitor and ground hook. | HARDWARE
132941 | 2016-09-06 | 3 | LI20 | 71 | 4 | Kly Mod 20-7(L0A): replace L4888 thyratron and thyratron stand with E2V CX1937AX deuterium thyratron and modified stand with trigger box and new blower. | HARDWARE
132941 | 2016-09-06 | 3 | LI20 | 71 | 4 | Kly Mod 20-7(L0A): replace L4888 thyratron and thyratron stand with E2V CX1937AX deuterium thyratron and modified stand with trigger box and new blower. | HARDWARE
132941 | 2016-09-06 | 7 | LI20 | 71 | 4 | Kly Mod 20-7(L0A): replace L4888 thyratron and thyratron stand with E2V CX1937AX deuterium thyratron and modified stand with trigger box and new blower. | HARDWARE
132941 | 2016-09-06 | 7.5 | LI20 | 71 | 4 | Kly Mod 20-7(L0A): replace L4888 thyratron and thyratron stand with E2V CX1937AX deuterium thyratron and modified stand with trigger box and new blower. | HARDWARE
133273 | 2016-10-07 | 1 | LI28 | 21 | 4 | 28-2 counting up WA#2 trips. | HARDWARE
133359 | 2016-10-14 | 1.2 | LI23 | 21 | 4 | 23-2 is experiencing beam volts dropouts. | HARDWARE
133360 | 2016-10-14 | 1.2 | LI23 | 41 | 4 | 23-4 is experiencing beam volts dropouts. | HARDWARE
133524 | 2016-10-30 | 0.8 | LI25 | 31 | 4 | 25-3 HVR with no remote reset. Comes back on then immediately trips again. | HARDWARE
133536 | 2016-10-31 | 0.8 | LI22 | 81 | 4 | KLYS:LI22:81 Beam voltage jitter is raising again up to 0.35 | HARDWARE
132064 | 2016-06-13 | 3.5 | LI28 | 21 | 4 | Modulator 28-2; PLC and Epics code modification request. This is in support of the LLRF test. We are now in the install and test phase of the LLRF mods. After the mod all commands and status signals will pass between the Modulator PLC and EPICs via the Ethernet port. The cable between the Modulator and MKSU, currently carrying those signals will be removed for the test. This mod is needed to make room for the PLC expansion needed for the LLRF test. Charge code: 17145-A8001 | HARDWARE
132064 | 2016-06-13 | 6.5 | LI28 | 21 | 4 | Modulator 28-2; PLC and Epics code modification request. This is in support of the LLRF test. We are now in the install and test phase of the LLRF mods. After the mod all commands and status signals will pass between the Modulator PLC and EPICs via the Ethernet port. The cable between the Modulator and MKSU, currently carrying those signals will be removed for the test. This mod is needed to make room for the PLC expansion needed for the LLRF test. Charge code: 17145-A8001 | HARDWARE
133643 | 2016-11-09 | 1.2 | LI29 | 81 | 4 | Mod. 29-7 Touch panel showing SCR fan fault. On the PLC relay PCB, the rearmost SCR fan indicator LED is blinking. Troubleshoot and repair as necessary. | HARDWARE
133684 | 2016-11-14 | 0.5 | LI29 | 11 | 4 | Klystron LI29-11 trips off on "Mod Fault" immediately after resetting. | HARDWARE
133589 | 2016-11-05 | 0.75 | LI24 | 31 | 4 | 24-3 trips immediately upon turn on. | HARDWARE
133591 | 2016-11-05 | 1 | LI21 | 11 | 4 | 21-1 L1S seeing spikes (increase) pulses in MOD waveform. These trip the AMO beam. See lclselog time stamp of this CATER as well as MCCELOG time of this CATER. | HARDWARE
133598 | 2016-11-07 |  | LI24 | 31 | 4 | 24-3 thyratron appears to have occasional miss-fires | HARDWARE
133598 | 2016-11-07 | 1 | LI24 | 31 | 4 | 24-3 thyratron appears to have occasional miss-fires | HARDWARE
133229 | 2016-10-05 | 0.5 | LI26 | 31 | 4 | 26-3 WGV fault | HARDWARE
133616 | 2016-11-08 | 2 | LI29 | 11 | 4 | 29-1 is MOD faulting frequently. (Not easy to tell given the LI29 CAMAC IOC problem but every time I visit the modulator panel I find it faulted.) | HARDWARE
133619 | 2016-11-08 | 1.3 | LI28 | 31 | 4 | 28 - 3 measure klystron heater AC waveform, voltage and current, with AC power analyzer | HARDWARE
133622 | 2016-11-08 |  | LI29 | 81 | 4 | 29 - 8 (or nearby) suspect failing PFN cap due to signature odor | HARDWARE
133625 | 2016-11-08 |  | LI21 | 11 | 4 | Downtime accounting: L1S (Klys 21-1) BVLT jitter made beam unusable. Had to swap out thyratron from klys 24-8. | HARDWARE
133667 | 2016-11-11 |  | LI24 | 11 | 4 | 24-1 has zero amplitude readback. Station appears to be contributing energy to the beam. Recommend IPL, IOC. OPs tried IPL, this did not recover the amplitude but did result in the amplitude error not propagating to the klystron CUD or other displays. The phase readback also became nan after IPL. Station seems to be still triggering and contributing expected amount of energy. The same symptoms have since arisen for 24-3 as well (CATER 133669) | HARDWARE
133290 | 2016-10-09 | 1 | LI24 | 61 | 4 | 24-6 tripping to HVR again | HARDWARE
133291 | 2016-10-09 |  | LI23 | 51 | 4 | Klystron 23-5 KMC after power dip. No remote reset. | HARDWARE
133291 | 2016-10-09 | 1 | LI23 | 51 | 4 | Klystron 23-5 KMC after power dip. No remote reset. | HARDWARE
133395 | 2016-10-17 | 1.5 | LI24 | 11 | 4 | Mod 24-1 Thyratron needs replacement: PFN and Beam volts (and other things) unstable. Ranged thy. res. from 6.35 to 7.3 Volts. Some improvement but still not good. | HARDWARE
133418 | 2016-10-18 | 0.75 | LI23 | 21 | 4 | Klys 23-2 has BVJ and PHJ problems. Think it is getting unwholesome ideas from 23-4 | HARDWARE
133489 | 2016-10-27 |  | LI24 | 31 | 4 | Klystron LI24-31 amplitude has suddenly dropped from 69.2 to 49.5. Zimmer edit 10/27 12:40: This may have been a rate dependent issue. It seems to somewhat correlate with us having the full machine at 10Hz rate for a while (instead of 120Hz). | HARDWARE
133527 | 2016-10-30 | 1.5 | LI27 | 11 | 4 | KLYS:LI27:1 is showing 1.2 % phase jitter. | HARDWARE
133879 | 2016-12-02 | 1.2 | LI30 | 41 | 4 | station 30-4 resets but won't turn on Edit 12/3/2016 20:49- this appears to be the case again, station immediately trips off when we try to bring it back on. | HARDWARE
133879 | 2016-12-02 | 1.2 | LI30 | 41 | 4 | station 30-4 resets but won't turn on Edit 12/3/2016 20:49- this appears to be the case again, station immediately trips off when we try to bring it back on. | HARDWARE
133883 | 2016-12-02 | 1 | LI20 | 81 | 4 | L0B, klys (20-8) suddenly dropped into heater delay. | HARDWARE
133883 | 2016-12-02 | 1 | LI20 | 81 | 4 | L0B, klys (20-8) suddenly dropped into heater delay. | HARDWARE
133884 | 2016-12-02 | 2 | LI21 | 41 | 4 | Klys 21-4 AMPL dropped to 3% suddenly. AMRF recommends changing PIOP. | HARDWARE
133887 | 2016-12-02 | 1.5 | LI25 | 31 | 4 | Station 25-3 mod faulted; won't reset. Edit 12/3/2016 20:27- this station is off again, and now immediately trips off when we try to reset it. | HARDWARE
133887 | 2016-12-02 | 1 | LI25 | 31 | 4 | Station 25-3 mod faulted; won't reset. Edit 12/3/2016 20:27- this station is off again, and now immediately trips off when we try to reset it. | HARDWARE
133887 | 2016-12-02 | 1.5 | LI25 | 31 | 4 | Station 25-3 mod faulted; won't reset. Edit 12/3/2016 20:27- this station is off again, and now immediately trips off when we try to reset it. | HARDWARE
133888 | 2016-12-02 | 2 | LI26 | 81 | 4 | 26-8 started tripping off constantly, took off the beam. See many amplitude drops and one reflected power error on EDM error display. | HARDWARE
133892 | 2016-12-03 | 2 | LI24 | 81 | 4 | 24-8 Thyratron needs replace due to instability | HARDWARE
133892 | 2016-12-03 | 1.5 | LI24 | 81 | 4 | 24-8 Thyratron needs replace due to instability | HARDWARE
133892 | 2016-12-03 | 2.3 | LI24 | 81 | 4 | 24-8 Thyratron needs replace due to instability | HARDWARE
133893 | 2016-12-03 |  | LI23 | 61 | 4 | 23-6 shows beam volts dropouts on FTP; observed correlation between phase jitter and LCLS xray jitter. | HARDWARE
133112 | 2016-09-25 | 12 | LI29 | 11 | 4 | station 29-1 reports Klys Heater Delay, | HARDWARE
133191 | 2016-09-30 |  | LI26 | 21 | 4 | 26-2 MOD with no remote reset. Went into KHD. | HARDWARE
133210 | 2016-10-02 | 1 | LI21 | 31 | 4 | Klystron LI21-31 is tripped off on Mod fault and will not reset remotely. | HARDWARE
133731 | 2016-11-18 |  | LI21 | 31 | 4 | 21-3 shows no amplitude and phase reads NaN. Doesn't seem to be producing any actual amplitude as tested with beam. Resetting/IPLing PIOP doesn't help. PAD has been replaced two times and the problem is still there. Recommend to replace PIOP. Refer to the follwing E-logs https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=873779 https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=873777 | HARDWARE
133741 | 2016-11-20 | 0.75 | LI21 | 71 | 4 | 21-7 having BVLT drop outs that are impeding lasing. | HARDWARE
133951 | 2016-12-08 | 2 | LI20 | 61 | 4 | INJ valves IN. Checked gauges and pressure looked normal so we attempted to pull the valves but they just fell back IN. Experts arrived and also noticed no obvious vac activity that would cause the valves to trip IN so after a few more attempts of pulling valves we decided to turn the GUN HV OFF then pull valves. With RF off, the valves stay open. We then lowered the HV for 20-6 to 15KV (from 17.45KV) and then lowered to AMPL setpoint in the LLRF VME display to 1MV. Next we energized the gun and ramped up the HV and AMPL setpoint slowly while watching vac activity in the INJ. We then increased the ampl until the setpoint reaches its max allowed then increased the HV until we got to 17.45KV. (-2H) temps stabilized, beam was delivered. Seems to be caused by GUN RF but no smoking gun found. | HARDWARE
133953 | 2016-12-08 | 1.2 | LI23 | 61 | 4 | 23-6 has phase, amplitude and beam voltage jitter + 0.2 to 2.5 jitter. (see pic.s in MCC log) | HARDWARE
133213 | 2016-10-03 | 1 | LI24 | 31 | 4 | 24-3 tripped on "MOD" fault. Won't reset. | HARDWARE
133227 | 2016-10-04 | 2 | LI25 | 31 | 4 | Klystron LI25-31 cycles and trips off repeatedly. At this time it shows 10 voltage faults and 3 reflected energy faults. Resets okay when it trips, but frequency of trips makes it nearly unusable for production running. | HARDWARE
133776 | 2016-11-23 |  | LI21 | 81 | 4 | Kly. LI 21 - 8 Phase jitter is jumping 3.0 deg.s (see pic.s in MCC log.) | HARDWARE
133790 | 2016-11-24 | 1.5 | LI22 | 51 | 4 | Kly. LI 22 - 5 tripped Mod - Fault Lockout. Won't reset. | HARDWARE
133867 | 2016-12-01 | 0.5 | LI30 | 41 | 4 | Klystron 30-4 has an Electro-Magnet Current fault | HARDWARE
133898 | 2016-12-04 | 1.2 | LI30 | 71 | 4 | 30-7 tripped, can't reset remotely. | HARDWARE
133923 | 2016-12-06 | 1 | LI27 | 61 | 4 | 27-6 is WGV faulted. | HARDWARE
134004 | 2016-12-13 |  | LI21 | 41 | 4 | Kly. LI 21 - 4 Amplit. drop from 50.1 to 44.1, causing the phase to change to - 39 deg.s from 0.0 .The FEL dropped to 20 from 41. Trimmed phase back up to 0.0 but Amplit. still low at 43.6. FEL back to 4.0 | HARDWARE
134004 | 2016-12-13 | 1.2 | LI21 | 41 | 4 | Kly. LI 21 - 4 Amplit. drop from 50.1 to 44.1, causing the phase to change to - 39 deg.s from 0.0 .The FEL dropped to 20 from 41. Trimmed phase back up to 0.0 but Amplit. still low at 43.6. FEL back to 4.0 | HARDWARE
133336 | 2016-10-12 | 1 | LI26 | 11 | 4 | Kly LI 26 - 1 tripped on "MOD" fault and won't reset. | HARDWARE
133553 | 2016-11-02 |  | LI23 | 11 | 4 | 23-1 seems to have lost accelerate triggers suddenly. Was on the beam and stopped contributing energy altogether, and showed no accelerate triggers. Swapped with 23-3 for now. 11/17/2016- happened again on swing shift tonight for a few hours 6/3/2017- happened again on swing shift at 20:28 | HARDWARE
133747 | 2016-11-21 |  | LI21 | 61 | 4 | 21-6 and 21-7 aren't getting any triggers at standby time. This isn't a major issue, because the stations still get triggers when they are on the beam, and we almost always use them. A fix for this can probably wait until a POMM or PAMM. | HARDWARE
134037 | 2016-12-16 | 1 | LI30 | 71 | 4 | Kly. LI 30 - 7 BVLT. failed to come back were it was on the previous on state. (see plot. in MCC log) | HARDWARE
134039 | 2016-12-17 | 1.3 | LI26 | 11 | 4 | 26-1 Trips on WGV when the DRIVE is at 50% | HARDWARE
134041 | 2016-12-18 | 2 | LI30 | 21 | 4 | 30-2 is wave guide vacuum fault. Op juiced filament to no avail. | HARDWARE
133174 | 2016-09-29 | 1 | LI28 | 61 | 4 | 28-6 has high beam volts jitter. | HARDWARE
133208 | 2016-10-02 | 0.5 | LI26 | 11 | 4 | 26-1 is MOD fault and no reset | HARDWARE
133298 | 2016-10-10 | 1 | LI29 | 11 | 4 | 29-1 showing DQing under current and high jitter | HARDWARE
133745 | 2016-11-21 | 2 | LI27 | 41 | 4 | 27-4 is FLK | HARDWARE
133642 | 2016-11-09 |  | LI29 | 61 | 4 | Following VVS 15 runup after the 11/9 PAMM, klystrons 29-6, 29-7, 29-8, 30-4, and 30-7 are all off and refusing to turn on. | HARDWARE
133650 | 2016-11-10 |  | LI23 | 41 | 4 | KLYS:LI23:41:Phase Jitter is 0.75 %. Was 0.75% before the PAMM. | HARDWARE
134009 | 2016-12-14 | 1 | LI24 | 31 | 4 | station 24-3 has high Beam Volts jitter. Has risen since 10:00 today. | HARDWARE
134012 | 2016-12-14 | 0.5 | LI28 | 41 | 4 | 28-4 shows Beam Volt jitter, steps of 1kV. Please have a look, res voltage may need an adjustment. | HARDWARE
134016 | 2016-12-14 | 1 | LI26 | 61 | 4 | Kly. LI 26 - 6 tripped it's VVS breaker. Need PEM / electricians to reset. | HARDWARE
134024 | 2016-12-15 | 2 | LI30 | 61 | 4 | 30-6 went into heater delay spontaneously. | HARDWARE
133459 | 2016-10-23 |  | LI23 | 11 | 4 | 23-1 lost accelerate triggers suddenly while on the beam. Klys status showed No Accelerate Triggers" even though the station was active on beam.
133460 | 2016-10-24 | 3 | LI28 | 11 | 4 | 28-1 has been in KHD for three hours. | HARDWARE
133068 | 2016-09-19 |  | LI24 | 31 | 4 | 24-3 PIOP bad CBL status | HARDWARE
134198 | 2017-02-09 | 1.5 | LI26 | 41 | 4 | 1 spare klystron in sector 26 has a dead KVGC | HARDWARE
134055 | 2016-12-19 | 0 | LI25 | 41 | 4 | LI25/26 focus supply bulk tripped off due to a ground current fault. PEM technician determined problem was with individual 25-4 klystron system. (DT SS 12/16/16 -0.3 hr MFX/CXI/XPP, 1.0 hr MEC) | HARDWARE
133314 | 2016-10-11 | 1.5 | LI24 | 11 | 4 | 24-1 Klystron focus readback (FMON) is noisey, jumps between 13 and 15 Amps. Expected is 14.5 Amps. This started when ferrites were added to the focus readback cable. | HARDWARE
133160 | 2016-09-28 | 1 | LI28 | 51 | 4 | 28-5: DQing cur is unstable. Sometimes, it dropped to 0A and then shot back up to 18A in few minutes. | HARDWARE
133184 | 2016-09-30 |  | LI26 | 11 | 4 | 26-1 no klystron filament heater voltage or current | HARDWARE
133723 | 2016-11-17 | 1.3 | LI23 | 41 | 4 | Kly. LI 23 - 4 showing Beam volt jitter. KLYS:LI23:41:BVLTHSTBPLOT shows 2 - 3 volt drops which might be causing this. KLYS:LI23:41:BVJTHSTBPLOT doesn't show any jitter. | HARDWARE
131795 | 2016-05-26 |  | LI28 | 11 | 4 | 28-1 Water leaks from the klystron hose. | HARDWARE
133310 | 2016-10-11 | 1.5 | LI24 | 11 | 4 | 24-1 had many RE faults on 10/11 owl shift. Ops lowered the drive request from 100% --> 89% as a temporary solution. | HARDWARE
133788 | 2016-11-24 | 2.3 | LI20 | 71 | 4 | Klys 20-7 (L0-A) is showing AMPL dropouts. I think this might be on beam volts as well, but I couldn't get those to plot for some reason. | HARDWARE
133798 | 2016-11-27 | 1.5 | LI28 | 31 | 4 | 28-3 tripped; can't reset remotely. | HARDWARE
133237 | 2016-10-05 | 1 | LI24 | 31 | 4 | station 24-3 trips repeatedly | HARDWARE
132000 | 2016-06-08 | 3 | LI24 | 81 | 4 | 24-8 (aka TCAV3): small T20 oil leak reported. | HARDWARE
133726 | 2016-11-18 | 1 | LI25 | 81 | 4 | 25-8 has faulted "CBL" status which does not remedy with PIOP IPL. | HARDWARE
134072 | 2017-01-04 | 1 | LI19 | 11 | 4 | 19-1 Remove battery ion pump supply and restore to AC when power is restored to sector 19 racks. | HARDWARE
133732 | 2016-11-18 |  | LI23 | 61 | 4 | 23-6 not contributing energy to beam, similar to ongoing problem with 23-1. Doesn't seem to be firing on accelerate. We are now down two stations in sector 23. | HARDWARE
133347 | 2016-10-13 | 1 | LI29 | 61 | 4 | KLYS:LI29:61:Beam voltage jitter is spiking from 0.1 to 0.5. Some drop outs on the Beam voltage read back. | HARDWARE
134080 | 2017-01-05 |  | LI24 | 41 | 4 | Klystron 24-4 shows signs of water puddling. Inspect klystron water system and weep hole. | HARDWARE
133299 | 2016-10-10 | 1 | LI30 | 41 | 4 | Modulator 30-4: core bias voltage transducer readback is about half of what it should be. | HARDWARE
133341 | 2016-10-12 | 2 | LI20 | 81 | 4 | L0B has a core bias fault. | HARDWARE
133771 | 2016-11-23 | 1 | LI23 | 61 | 4 | Klys 23-6 is cycling frequently. Tend to see MOD HVOC fault. | HARDWARE
133771 | 2016-11-23 | 1 | LI23 | 61 | 4 | Klys 23-6 is cycling frequently. Tend to see MOD HVOC fault. | HARDWARE
134079 | 2017-01-05 | 1 | LI22 | 41 | 4 | Klystron 22-4 shows signs of water puddling. Inspect klystron water system and weep hole. | HARDWARE
123580 | 2015-03-17 | 1.5 | LI25 |  | 4 | Sector 25 Modulators: HVPS Regulator and Enerpro boards need modification (1) | HARDWARE
125686 | 2015-06-12 |  | LI30 | 71 | 4 | The WVGC measures LO E-10 on the PIOP it reads 3.6 E-9. While the station was off, there was not a change in vacuum. Set point 1 would not be able to effectively pass NIRP. | HARDWARE
125307 | 2015-05-24 |  | LI30 | 11 | 4 | 30-1: Vacuum glitches causing station to drop out. | HARDWARE
126005 | 2015-07-01 |  | LI21 | 21 | 4 | Suspect standby PAC or modulator timing is not correct. Noticed this by observing the 6x6 commands for L1S phase and amplitude when changing rates. Temperature change is small, <0.5 deg F, and 6x6 is able to compensate. Need invasive time (POMM) to investigate (need beam rate to 10 hz, and disable LLRF feedback). | HARDWARE
127370 | 2015-09-29 |  | LI21 | 11 | 4 | L1S phase and amplitude change in response to LCLS rate changes. Presumably indicates that standby timing is not set up properly. See this elog entry: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=803194 | HARDWARE
128904 | 2015-11-24 |  | LI20 | 61 | 4 | 20-6 Bad high power attenuator. Drive does not go to zero. | HARDWARE
129820 | 2016-01-21 |  | LI30 | 41 | 4 | Sector 30 - The Klystron Cooling System water flows appears to be the lowest flow at 30-4. That station has been tripping on KMC, which could be improved by better flow. | HARDWARE
131523 | 2016-04-30 |  | LI21 | 81 | 4 | Klys 21-8 and 22-8 getting BVLT dropouts on the same pulse. Would CATER each individual klystron, but the problem is common to both of them as their BVLT dropouts are always on the same pulse. Makes it seem less like a thyratron problem, but unclear. This is too weird, not sure what could cause that. MOD 5/1/16 HVS this is a LLRF timing problem with LI22 FIDO. LI22 FIDO RF levels need to be checked. Timing of PDUs which use this FIDO are jumping ~500nS w.r.t. the rest of the timing system, causing these apparent" dropouts.
135746 | 2017-06-22 | 1 | LI28 | 81 | 4 | 28-8 is in mod fault lock out and cannot be reset remotely. | HARDWARE
135124 | 2017-05-18 | 0.5 | LI22 | 81 | 4 | station 22-8 faults repeatedly--3 times in the last hour | HARDWARE
135129 | 2017-05-18 | 1 | LI28 | 11 | 4 | 28-1 tripped on Core Bias | HARDWARE
135249 | 2017-05-23 | 1.3 | LI28 | 41 | 4 | 28-4 core bias current reads OK, core bias voltage abnormally high. Troubleshoot and repair | HARDWARE
136538 | 2017-09-02 | 1.5 | LI26 | 71 | 4 | 26-7: Fire wire fault. | HARDWARE
136539 | 2017-09-02 | 1.5 | LI30 | 71 | 4 | 30-7: Fire wire fault. | HARDWARE
136540 | 2017-09-02 | 0.03 | LI28 | 41 | 4 | 28-4 FLK. | HARDWARE
136543 | 2017-09-02 | 1 | LI26 | 21 | 4 | Kly Mod 26-2: station tripped off due to external faults. | HARDWARE
136544 | 2017-09-02 | 0 | LI26 | 21 | 4 | 26-2 tripped HVOV reported. | HARDWARE
136545 | 2017-09-02 | 0 | LI29 | 81 | 4 | 29-8 tripped ACOC reported | HARDWARE
136547 | 2017-09-02 | 0.5 | LI22 | 21 | 4 | 22-2 WGV appears to be asleep needs filament juiced. | HARDWARE
136549 | 2017-09-02 |  | LI27 | 41 | 4 | 27-4 is tripping on WA1 flowswitch faults. | HARDWARE
136549 | 2017-09-02 | 1 | LI27 | 41 | 4 | 27-4 is tripping on WA1 flowswitch faults. | HARDWARE
136550 | 2017-09-02 | 0.1 | LI28 | 41 | 4 | 28-4 tripped on klystron vacuum. Readout history of klys vac looks suspiciously non-physical with huge diurnal temperature correlated swings. | HARDWARE
136553 | 2017-09-02 |  | LI27 | 31 | 4 | 27-3 focus fault. | HARDWARE
136556 | 2017-09-02 | 2 | LI28 | 61 | 4 | 28-6 fire wire trip. | HARDWARE
136557 | 2017-09-02 | 2 | LI25 | 51 | 4 | 25-5 MOD fault lockout. It was MOD VVS faulted and changed to FLK. Unclear what is actually wrong. | HARDWARE
134398 | 2017-03-23 | 1.2 | LI22 | 11 | 4 | 22-1: problems found: 1. 480VAC cables not properly secured. 2. Primary snubber resistors: wires and lugs burnt. 3. Two filter cap 4.75uF missing green ground wires. 4. The big ground green terminal block that mounted to top left corner in the T-20 compartment" has cable loose.
134955 | 2017-05-12 |  | LI23 | 71 | 4 | 23-7 has low beam volts. | HARDWARE
135570 | 2017-06-11 | 0.5 | LI26 | 11 | 4 | 26-1 Modulator problems found: 1. Primary snubber resistor R260, R261 burnt lugs. 2. HV filter 4.7uF, 30,000VDC cap missing ground wires. 4. T2-xformer 120VAC ground wire loose. 5. Core Bias Assembly: TB1 phoenix block has 120VAC BLK & WHT wires: strands protruding from the connection. | HARDWARE
135570 | 2017-06-11 | 0.5 | LI26 | 11 | 4 | 26-1 Modulator problems found: 1. Primary snubber resistor R260, R261 burnt lugs. 2. HV filter 4.7uF, 30,000VDC cap missing ground wires. 4. T2-xformer 120VAC ground wire loose. 5. Core Bias Assembly: TB1 phoenix block has 120VAC BLK & WHT wires: strands protruding from the connection. | HARDWARE
135570 | 2017-06-11 | 1 | LI26 | 11 | 4 | 26-1 Modulator problems found: 1. Primary snubber resistor R260, R261 burnt lugs. 2. HV filter 4.7uF, 30,000VDC cap missing ground wires. 4. T2-xformer 120VAC ground wire loose. 5. Core Bias Assembly: TB1 phoenix block has 120VAC BLK & WHT wires: strands protruding from the connection. | HARDWARE
135570 | 2017-06-11 | 1 | LI26 | 11 | 4 | 26-1 Modulator problems found: 1. Primary snubber resistor R260, R261 burnt lugs. 2. HV filter 4.7uF, 30,000VDC cap missing ground wires. 4. T2-xformer 120VAC ground wire loose. 5. Core Bias Assembly: TB1 phoenix block has 120VAC BLK & WHT wires: strands protruding from the connection. | HARDWARE
135570 | 2017-06-11 | 1 | LI26 | 11 | 4 | 26-1 Modulator problems found: 1. Primary snubber resistor R260, R261 burnt lugs. 2. HV filter 4.7uF, 30,000VDC cap missing ground wires. 4. T2-xformer 120VAC ground wire loose. 5. Core Bias Assembly: TB1 phoenix block has 120VAC BLK & WHT wires: strands protruding from the connection. | HARDWARE
135571 | 2017-06-11 | 1 | LI26 | 21 | 4 | 26-2 Modulator problems found: 1.Primary snubber resistor R261 burnt wire, wire insulation cracked. 2. HV filter 4.7uF 30,000VDC cap missing ground wires. 4. Two De-spiking coils covers cracked (see attached picture). | HARDWARE
135572 | 2017-06-11 | 1 | LI26 | 31 | 4 | 26-3 Modulator problems found: 1. Two primary snubber caps: jumper wires between them are burnt. 2. HV filter 4.7uF 30,000VDC cap missing ground wires. 3. Door#2: top 600VAC holes measured= 2 1/2 inches diameter holes should be covered. | HARDWARE
135574 | 2017-06-11 | 0.5 | LI26 | 41 | 4 | 26-4 Modulator problems found: 1. Primary snubber resistor R261 lug burnt. 2. HV filter 4.7uF 30,000VDC cap missing ground wires. 4. Bad Grid Trigger chassis. Found screw that hold the resistor and coils together is broken. ( item #4 fixed by PN) | HARDWARE
135574 | 2017-06-11 | 1 | LI26 | 41 | 4 | 26-4 Modulator problems found: 1. Primary snubber resistor R261 lug burnt. 2. HV filter 4.7uF 30,000VDC cap missing ground wires. 4. Bad Grid Trigger chassis. Found screw that hold the resistor and coils together is broken. ( item #4 fixed by PN) | HARDWARE
135575 | 2017-06-11 | 1 | LI26 | 51 | 4 | 26-5 Modulator problems found: 1. Primary snubber resistor R261 lug burnt. 2. HV filter 4.7uF 30,000VDC cap missing ground wires. 4. PFN strap bottom (RHS #4). | HARDWARE
135771 | 2017-06-26 | 2 | LI29 | 51 | 4 | 29-5 modulator tripping repeatedly on End of Line Clipper fault. | HARDWARE
135955 | 2017-07-08 | 2 | LI26 | 61 | 4 | Kly Mod 26-6: station tripped on variety of faults including ACOC and HVOC. Station in local mode. | HARDWARE
135955 | 2017-07-08 | 1 | LI26 | 61 | 4 | Kly Mod 26-6: station tripped on variety of faults including ACOC and HVOC. Station in local mode. | HARDWARE
135957 | 2017-07-09 | 1 | LI22 | 41 | 4 | Klystron 22-4 tripping periodically on ACOC faults.Trips several times a day but does reset remotely. | HARDWARE
136561 | 2017-09-03 | 0.5 | LI27 | 41 | 4 | 27-4 is in FLK | HARDWARE
136578 | 2017-09-05 | 1 | LI23 | 31 | 4 | 23-3 MOD fault, no reset | HARDWARE
135027 | 2017-05-13 |  | LI25 | 21 | 4 | 25-2 Has a waveguide vacuum fault. | HARDWARE
135028 | 2017-05-13 |  | LI25 | 11 | 4 | 25-1 has low beam voltage at 313KV expected is 350KV. | HARDWARE
135029 | 2017-05-13 |  | LI26 | 11 | 4 | 26-1 has low beam voltage at 320KV expected is 350KV. | HARDWARE
135031 | 2017-05-13 | 0.5 | LI26 | 51 | 4 | 26-5 has low beam voltage at 325KV expected is 345KV. | HARDWARE
135045 | 2017-05-14 | 2 | LI21 | 41 | 4 | 21-4 WCS LCW leak from flow restrictor. Butler is on it. | HARDWARE
135587 | 2017-06-12 | 1 | LI26 | 81 | 4 | 26-8 is off on KHD and won't clear | HARDWARE
135588 | 2017-06-12 | 0.5 | LI28 | 41 | 4 | 28-4 is KMC | HARDWARE
135596 | 2017-06-12 | 1 | LI20 | 81 | 4 | Klystron 20-8 (L0B) has a large amount of beam volts jitter | HARDWARE
136047 | 2017-07-16 |  | LI26 | 51 | 4 | 26-5 window 2 temperature (KLYS:LI26:51:WNDW2_TEMP) is 149deg F. | HARDWARE
136047 | 2017-07-16 |  | LI26 | 51 | 4 | 26-5 window 2 temperature (KLYS:LI26:51:WNDW2_TEMP) is 149deg F. | HARDWARE
136051 | 2017-07-16 | 1.5 | LI29 | 61 | 4 | 29-8 frequent trips "MOD EVOC" | HARDWARE
136054 | 2017-07-16 |  | LI21 | 41 | 4 | 21-4 CBL fault does not reset with PIOP IPL or PDU reset. | HARDWARE
136054 | 2017-07-16 | 1.5 | LI21 | 41 | 4 | 21-4 CBL fault does not reset with PIOP IPL or PDU reset. | HARDWARE
136059 | 2017-07-17 | 0.5 | LI21 | 51 | 4 | 21-5 has MKT fault. PDU reset, PIOP reset, or IPL do not recover. | HARDWARE
136579 | 2017-09-05 | 1.5 | LI29 | 21 | 4 | 29-2: The PIOP has been frozen since 9/5/17 @ 1517 and will not clear with an IPL. | HARDWARE
136586 | 2017-09-05 | 0.5 | LI25 | 81 | 4 | Klystron 25-8 MOD FLK and won't reset remotely | HARDWARE
136592 | 2017-09-06 | 1.5 | LI24 | 81 | 4 | 24-8 Klystron beam voltage reads 32KV low on PIOP data. | HARDWARE
135171 | 2017-05-20 | 2 | LI21 | 11 | 4 | L1S window 2 temp is reading 180 degF should be 95 degF. Has been this way for two days now. Ops noticed this while troubleshooting observed HV problems with L1S modulator. The value seems high and should be addressed asap. Steele 5/21/17: This problem started on May first, before RF had been turned on. Must be a controls readback problem, over to CTL. This is a duplicate of CATER 134920. | HARDWARE
135618 | 2017-06-13 | 1 | LI24 | 41 | 4 | 24-4 tripped off. Can reset but it trips immediately. | HARDWARE
136034 | 2017-07-14 | 0.5 | LI28 | 61 | 4 | 28-6: Seeing some beam volts jitter. | HARDWARE
136043 | 2017-07-15 | 1.5 | LI29 | 31 | 4 | 29-3 MOD fault no reset, fault is ACOC. | HARDWARE
136060 | 2017-07-17 | 1.5 | LI25 | 81 | 4 | When assigned to accelerate time, 25-8 trips on "MKT" fault every few minutes. This does not happen when assigned to standby time. | HARDWARE
136270 | 2017-08-02 | 0.5 | LI28 | 41 | 4 | 28-4 Station wave guide VAC-ION does not respond when set point buttons are pushed. | HARDWARE
135741 | 2017-06-22 | 6 | LI20 | 81 | 4 | Building 15, South Bay Load Dock: remove PCB mineral oil (< 50 ppm) from T20 transformer, S/N 818 to haz waste barrel for disposal. T20 removed from klystron modulator 20-8 (L0B) yesterday. | HARDWARE
135741 | 2017-06-22 | 5 | LI20 | 81 | 4 | Building 15, South Bay Load Dock: remove PCB mineral oil (< 50 ppm) from T20 transformer, S/N 818 to haz waste barrel for disposal. T20 removed from klystron modulator 20-8 (L0B) yesterday. | HARDWARE
135741 | 2017-06-22 | 3 | LI20 | 81 | 4 | Building 15, South Bay Load Dock: remove PCB mineral oil (< 50 ppm) from T20 transformer, S/N 818 to haz waste barrel for disposal. T20 removed from klystron modulator 20-8 (L0B) yesterday. | HARDWARE
136596 | 2017-09-06 |  | LI16 | 41 | 4 | 16-4B S/N 396C and 16-4C S/N 250A have poor klystron vacuum. | HARDWARE
136603 | 2017-09-07 | 0.15 | LI24 | 61 | 4 | Klystron LI24-61 is tripped off on a WGV fault and needs a local reset. | HARDWARE
135178 | 2017-05-20 |  | LI21 | 21 | 4 | L1X (21-2) goes MKT (MKSU trigger enable fault) fairly frequently for several minutes at a time. This happens during the heat of the day. It seems to heat up to the point where it trips MK, then cools down, runs for 30 minutes, then trips again. It stops when it is sufficiently cool outside. Also a downtime accounting CATER. | HARDWARE
135182 | 2017-05-21 | 4 | LI20 | 81 | 4 | 20-8 modulator: replace bad T20 | HARDWARE
135186 | 2017-05-21 | 2 | LI26 | 81 | 4 | 26-8: Klystron heater problems. Bad power supply and/or klystron heater problem. | HARDWARE
135190 | 2017-05-21 |  | LI20 | 81 | 4 | 20-8: Raise BV's back to 350kV, check PIOP ratios. | HARDWARE
135197 | 2017-05-22 |  | LI20 | 81 | 4 | L0B still having bouts of instability. Amplitude drifts away with no sign of BV variation . Maybe check the drive. The amplitude is very stable at 10Hz and maintained stability for two hours. When going back to 120Hz, The ampl immediately dropped. Zeroed offsets and ran pau re-synch script and the ampl is stable now (not as much as 10Hz). Stable enough for the water flow to handle the amount of change. After about 5 minutes, the amplitude excursions start to grow. | HARDWARE
136063 | 2017-07-17 | 1 | LI28 | 71 | 4 | 28-7 tripping frequently ACOC (four times last hour) | HARDWARE
136080 | 2017-07-18 | 1 | LI20 | 71 | 4 | Klystron LI20-71 (L0A) is tripped off on Water Summary Fault. Accelerator Water #2 fault indicated. | HARDWARE
136081 | 2017-07-19 | 0.5 | LI21 | 41 | 4 | 21-4 has an MKSU Protect fault and does not reset remotely. PIOP data shows Electro-Magnet Tol fault. Beam current and Beam volt readouts are unrealistically high. | HARDWARE
136243 | 2017-07-31 | 1.5 | LI20 | 81 | 4 | L0B has high BV jitter. | HARDWARE
136243 | 2017-07-31 |  | LI20 | 81 | 4 | L0B has high BV jitter. | HARDWARE
136246 | 2017-07-31 | 1.2 | LI27 | 31 | 4 | 27-3 Modulator: inspect T20 for leakage, clean leaked oil and check (and top off if needed) oil level. | HARDWARE
136289 | 2017-08-04 | 2 | LI28 | 81 | 4 | 28-8 is mod faulted and will not reset remotely. | HARDWARE
136292 | 2017-08-05 | 1 | LI29 | 71 | 4 | Klystron 29-7 appears to have tripped its fire wire. | HARDWARE
136292 | 2017-08-05 | 2 | LI29 | 71 | 4 | Klystron 29-7 appears to have tripped its fire wire. | HARDWARE
136294 | 2017-08-06 | 1 | LI21 | 51 | 4 | Klystron 21-5 MKSU trigger enable fault Happened again, 8/11/2017 11:31 | HARDWARE
136294 | 2017-08-06 | 0.5 | LI21 | 51 | 4 | Klystron 21-5 MKSU trigger enable fault Happened again, 8/11/2017 11:31 | HARDWARE
136296 | 2017-08-06 | 1 | LI24 | 31 | 4 | Klystron 24-3 MOD fault lockout | HARDWARE
136306 | 2017-08-07 |  | LI20 | 81 | 4 | L0B MOD faulting repeatedly. Now locked out. | HARDWARE
136615 | 2017-09-08 |  | LI25 | 41 | 4 | Klystron LI25-41 is in Mod Fault Lockout and will not reset remotely. | HARDWARE
136616 | 2017-09-08 | 1 | LI25 | 41 | 4 | 25-4: Tripping on EOLC faults. | HARDWARE
136616 | 2017-09-08 | 50 | LI25 | 41 | 4 | 25-4: Tripping on EOLC faults. | HARDWARE
136619 | 2017-09-08 |  | LI23 | 41 | 4 | 23-4 shows BVJ but jitter FTPs do not show substantial RMS jitter. | HARDWARE
136620 | 2017-09-08 |  | LI21 | 51 | 4 | 21-5 shows BVJ but jitter FTPs do not show substantial RMS jitter. | HARDWARE
136622 | 2017-09-09 | 2 | LI23 | 31 | 4 | Klys 23-3 showing bad cable status. Tried IPL, failed hard, might be PIOP, unclear. | HARDWARE
136622 | 2017-09-09 | 1.5 | LI23 | 31 | 4 | Klys 23-3 showing bad cable status. Tried IPL, failed hard, might be PIOP, unclear. | HARDWARE
135210 | 2017-05-22 | 1 | LI27 | 71 | 4 | 27-7 is FLK | HARDWARE
135222 | 2017-05-23 |  | LI25 | 11 | 4 | Klystrons 25-1, 26-6, 27-7 MOD faulted and won't reset remotely | HARDWARE
136253 | 2017-08-01 | 1 | LI24 | 31 | 4 | Klystron 24-3 faulting frequently. Three times in as many minutes. Edit 8/9/2017 08:30- this station still faulting frequently. Edit 8/23/2017 21:08 - This station is still faulting frequently, seems to fault more often when it is on accelerate than in standby. Edit 9/8/2017- this station still faulting frequently | HARDWARE
136253 | 2017-08-01 |  | LI24 | 31 | 4 | Klystron 24-3 faulting frequently. Three times in as many minutes. Edit 8/9/2017 08:30- this station still faulting frequently. Edit 8/23/2017 21:08 - This station is still faulting frequently, seems to fault more often when it is on accelerate than in standby. Edit 9/8/2017- this station still faulting frequently | HARDWARE
136253 | 2017-08-01 | 0.5 | LI24 | 31 | 4 | Klystron 24-3 faulting frequently. Three times in as many minutes. Edit 8/9/2017 08:30- this station still faulting frequently. Edit 8/23/2017 21:08 - This station is still faulting frequently, seems to fault more often when it is on accelerate than in standby. Edit 9/8/2017- this station still faulting frequently | HARDWARE
136264 | 2017-08-01 |  | LI28 | 21 | 4 | 28-2 showing large amounts of phase jitter on the SsSB output. | HARDWARE
136308 | 2017-08-07 | 1 | LI28 | 21 | 4 | 28-2 Faults on klystron focus current (KMC and EMC, regulator 1 and 2) when focus current is changed. | HARDWARE
136308 | 2017-08-07 | 0.5 | LI28 | 21 | 4 | 28-2 Faults on klystron focus current (KMC and EMC, regulator 1 and 2) when focus current is changed. | HARDWARE
136312 | 2017-08-07 |  | LI28 | 21 | 4 | Station 28-2 modulator is running seemingly normally by the CUD status PV: CUDKLYS:LI28:2:STATUS reports HVR | HARDWARE
136351 | 2017-08-13 | 1 | LI26 | 61 | 4 | 26-6 tripped on VVS breaker. Requires electrician to reset with VVS ON. Please reset Monday morning, thanks. This is the second occurrence in three weeks. Please see: Cater 136216 7-27-2017 | HARDWARE
136388 | 2017-08-17 | 1 | LI24 | 21 | 4 | 24-2: PFN has unreasonable slope and ripple. | HARDWARE
134325 | 2017-03-13 |  | LI29 | 41 | 4 | 29-4 klystron has leaky filament connector. | HARDWARE
135230 | 2017-05-23 |  | LI25 | 41 | 4 | 25-4 Waveguide vacuum controller is faulted. | HARDWARE
135232 | 2017-05-23 | 2 | LI21 | 21 | 4 | 21-2 L1X - Klystron focus current readbacks are all zero. | HARDWARE
135288 | 2017-05-25 | 0.5 | LI25 | 31 | 4 | Klystron LI25-3 will not go back to SLED mode remotely after changing to non-SLED mode. | HARDWARE
136147 | 2017-07-24 |  | LI20 | 81 | 4 | L0B tripped on over-voltage three times on 7/24 owl shift. | HARDWARE
136147 | 2017-07-24 |  | LI20 | 81 | 4 | L0B tripped on over-voltage three times on 7/24 owl shift. | HARDWARE
136265 | 2017-08-01 |  | LI28 | 21 | 4 | 28-2 Cannot saturate klystron. DAC out is maxed out at 32K. | HARDWARE
136395 | 2017-08-18 | 0 | LI23 | 81 | 4 | 23-8 starting tripping off incessantly starting at ~1:30am on 8/18 with about 10 trips in about an hour. | HARDWARE
134351 | 2017-03-19 | 2 | LI21 | 21 | 4 | LI21, Power supply GEN 12.5-60, voltage and current are unstable. It located at rack KF21-2C elevation 23 near Modulator 21-2. | HARDWARE
134352 | 2017-03-19 | 1 | LI20 | 71 | 4 | 20-7: Core Bias Assembly - 120VAC harness "P2" the secure clip that plug into the Pheonix connector is broken. | HARDWARE
134355 | 2017-03-20 | 1 | LI21 | 11 | 4 | 21-1: Core Bias Assembly: DC harness P1" secure clip is broken.
134357 | 2017-03-20 | 1 | LI21 | 31 | 4 | 21-3 Problems found: 1. Core Bias Assembly, 120VAC harness P2" secure clip broken.
135337 | 2017-05-28 | 1 | LI29 | 51 | 4 | 29-5 MOD fault | HARDWARE
135339 | 2017-05-28 | 2 | LI26 | 11 | 4 | 26-1 trips every several minutes. | HARDWARE
135345 | 2017-05-29 | 1.5 | LI24 | 41 | 4 | 21-4 Station needs PFN tuning for a flatter pulse top and a slope adjustment better than .5. | HARDWARE
136283 | 2017-08-03 | 4 | LI28 | 41 | 4 | Kly Mod 28-4: install 5K : 1 Ross voltage divider in cabinet 3 of modulator. Run compensated cable to HVPS bulkhead feedthru connector, # FT-14. Tie off BNC jack next to isolated feedthru for installation with gen 2 regulator board. | HARDWARE
136321 | 2017-08-09 | 0.5 | LI22 | 71 | 4 | 22-7 showing large phase jitter, visible effect on beam. | HARDWARE
136323 | 2017-08-09 | 0 | LI21 | 51 | 4 | 21-5 Focus supply trips on current fault. Focus supply (Genesis) needs to be replaced. | HARDWARE
136329 | 2017-08-09 | 1.2 | LI29 | 71 | 4 | 29-7 tripped fire wire and has "phase" and "VVS" faults showing on PLC. Locally no fire no smoke no smell of smoke. | HARDWARE
136408 | 2017-08-21 | 1 | LI25 | 21 | 4 | Klys 25-2 cycling frequently. Also, it shows 255 errors for random faults on error display and the AMPL readback jumps to 68.73 frequently. Right now showing FOX error. Seems like problem with PAD, but unclear. | HARDWARE
136416 | 2017-08-22 |  | LI25 | 21 | 4 | 25-2 PIOP froze after IPL and may needs swap. | HARDWARE
135363 | 2017-05-30 | 2 | LI25 | 11 | 4 | 25-1: PAD phase read back not working (shows nan). Replace PAD. | HARDWARE
135365 | 2017-05-30 | 0.5 | LI22 | 31 | 4 | 22-3 is faulted on WGV | HARDWARE
136077 | 2017-07-18 |  | LI21 | 21 | 4 | Klystron 21-2 MOD faulted and won't reset remotely | HARDWARE
136331 | 2017-08-10 | 1 | LI22 | 71 | 4 | 22-7: Tune PFN for a flatter signal top. Analysis slope is unreasonable (1). | HARDWARE
136333 | 2017-08-10 |  | LI21 | 31 | 4 | 21-3:DQ > 400 amps on EPICS even after upgrades. | HARDWARE
136334 | 2017-08-10 |  | LI25 | 21 | 4 | Klystron 25-2 tripping repeatedly on HVOV faults. | HARDWARE
136338 | 2017-08-10 |  | LI28 | 21 | 4 | RFS 28-2 focus breaker is tripped. | HARDWARE
136338 | 2017-08-10 | 1 | LI28 | 21 | 4 | RFS 28-2 focus breaker is tripped. | HARDWARE
136338 | 2017-08-10 | 1 | LI28 | 21 | 4 | RFS 28-2 focus breaker is tripped. | HARDWARE
136341 | 2017-08-11 | 3 | LI23 | 81 | 4 | 23-8: ACR log reports rash of HVOC faults. | HARDWARE
136341 | 2017-08-11 | 0 | LI23 | 81 | 4 | 23-8: ACR log reports rash of HVOC faults. | HARDWARE
136342 | 2017-08-11 | 1 | LI21 | 51 | 4 | 21-5: DQ current appears to be stuck at 20 amps. | HARDWARE
134413 | 2017-03-26 | 0.5 | LI29 | 11 | 4 | 29-1 Modulator: Problems found: 1. Burnt secondary snubber resistor lug that cable connects to R1 of xformer. 2. Burnt core bias filter reactor, must be replaced before turn modulator back on. (Put blue block with tag) 3. Thyratron base bolts rusted. 4. The connector that plug to Klystron tank, the thread is damaged. | HARDWARE
134414 | 2017-03-26 | 0.5 | LI29 | 21 | 4 | 29-2 Modulator: Problems found: 1. Burnt secondary snubber resistor lugs. 2. HV filter caps 4.75uF missing two ground wires. 3. P2 pheonix cable connector that connects to core bias assembly, the secure clip is broken. 4. PFN strap #4 (LHS) upper loose 5. De-Spiking coil resistors, there is one resistor has wrong resistance value=15000 ohm, should be 2000 ohm. | HARDWARE
134415 | 2017-03-26 | 0.5 | LI29 | 41 | 4 | 29-3 Modulator problems found: 1. Core Bias Assembly: 120VAC P2" cable connector secure clip is broken.
134418 | 2017-03-26 | 1 | LI29 | 61 | 4 | 29-6 Modulator problems found: 1. Primary snubber resistor: All wires, lugs and bottom resistor R262" should be replaced.
134419 | 2017-03-26 | 0.5 | LI29 | 71 | 4 | 29-7 Modulator problems found: 1. Burnt primary wires and lugs, they need to be replaced. 2. Two filter cap cases are missing ground wires. | HARDWARE
134420 | 2017-03-26 | 0.5 | LI29 | 81 | 4 | 29-8: Two HV filter cap cases are missing ground wires. | HARDWARE
135380 | 2017-05-30 | 1.5 | LI20 | 61 | 4 | 20-6 is unstable. two trips within as many minutes. EVO faults before FLK. | HARDWARE
135380 | 2017-05-30 | 2 | LI20 | 61 | 4 | 20-6 is unstable. two trips within as many minutes. EVO faults before FLK. | HARDWARE
135888 | 2017-07-02 |  | LI28 | 31 | 4 | Klystron LI28-31 is tripped on Mod Fault and will not reset remotely. | HARDWARE
136143 | 2017-07-23 | 6 | LI20 | 61 | 4 | Getting amplitude dropouts on klys 20-6 (gun). Not seen in modulator. Appears to be LLRF problem. Problem got very bad starting around 7/23 @ 09:00. | HARDWARE
136344 | 2017-08-11 | 1 | LI25 | 41 | 4 | 25-4: Tune PFN for flatter signal top. Slope is presently -1.7. | HARDWARE
136346 | 2017-08-11 | 1.5 | LI21 | 51 | 4 | 21-5 MKSU trigger enable fault. No triggers to modulator. See 135294. | HARDWARE
136347 | 2017-08-11 | 0.7 | LI24 | 31 | 4 | 24-3 after bout of very frequent trips (not clear source of trips) it is no Fault Lockout and will not reset. See 136253. | HARDWARE
136357 | 2017-08-14 |  | LI28 | 21 | 4 | 28-2 Phase feedback appears to be broken. tkurty 8/14/17 16:00: T4 phase feedback working correctly, T1 feedback was stuck between -180 and 180 deg. Changing PDES to 10 deg allowed T1 to leave the phase-flip region and stabilize to the same value as T4. Reset PDES to 0 and both timeslot feedbacks are working Broken state: [softegr@lcls-srv01 tkurty]$ caget KLYS:LI28:21:T{1,4}:CORRPHASE KLYS:LI28:21:T1:CORRPHASE 75.4542 KLYS:LI28:21:T4:CORRPHASE 92.8884 After resetting PDES: [softegr@lcls-srv01 tkurty]$ caget KLYS:LI28:21:T{1,4}:CORRPHASE KLYS:LI28:21:T1:CORRPHASE 90.9837 KLYS:LI28:21:T4:CORRPHASE 92.4988 | HARDWARE
136424 | 2017-08-23 | 1 | LI20 | 51 | 4 | 20-5 Modulator problems found: 1. T20 - phase B R2" base loose ground.
136424 | 2017-08-23 | 1 | LI20 | 51 | 4 | 20-5 Modulator problems found: 1. T20 - phase B R2" base loose ground.
136426 | 2017-08-23 |  | LI30 | 11 | 4 | Klystron 30-1 tripped its VVS breaker | HARDWARE
136432 | 2017-08-24 | 1 | LI20 | 61 | 4 | 20-6 Modulator problems found: 1. T20 compartment: 120VAC T2" xformer - loose ground wire.
136432 | 2017-08-24 | 1 | LI20 | 61 | 4 | 20-6 Modulator problems found: 1. T20 compartment: 120VAC T2" xformer - loose ground wire.
136432 | 2017-08-24 | 1 | LI20 | 61 | 4 | 20-6 Modulator problems found: 1. T20 compartment: 120VAC T2" xformer - loose ground wire.
136439 | 2017-08-25 |  | LI24 | 21 | 4 | 24-2 is running without DQing regulation =PHJ, AMJ and BVJ. | HARDWARE
134430 | 2017-03-28 | 4.5 | LI20 | 71 | 4 | 20-7: Replace old ground hook on door 4 with new style hook. Add hanger for hook safety. | HARDWARE
134431 | 2017-03-28 | 2 | LI23 | 51 | 4 | 23-5: Replace old ground hook on door 4 with new style hook. Add hanger for hook safety. | HARDWARE
134432 | 2017-03-28 | 2 | LI26 | 11 | 4 | 26-1: Replace old ground hook on door 4 with new style hook. Add hanger for hook safety. | HARDWARE
134433 | 2017-03-28 | 2.5 | LI30 | 41 | 4 | 30-4: Replace old ground hook on door 4 with new style hook. Add hanger for hook safety. | HARDWARE
135385 | 2017-05-31 | 0.1 | LI26 | 31 | 4 | LI26-31 repeatedly flashes "Stat_Bad_Camac" and does not have a PAD Phase value. | HARDWARE
135395 | 2017-05-31 | 0.5 | LI27 | 31 | 4 | Klystron 27-3 MOD fault and no reset | HARDWARE
135403 | 2017-05-31 |  | LI20 | 81 | 4 | Klys 20-8 (L0B) has a tough time regulating at 1 Hz. See it in Beam Volts/Phas/AMPL. Not sure if this is timing or LLRF feedback problem. Unable to get FTP at 1 Hz to see if firing time is shifted. Might need a scope. | HARDWARE
136358 | 2017-08-14 |  | LI21 | 21 | 4 | 21-2 L1X is showing frequent instances of running into saturation and causing the LLRF feedback to push to drive power to the software clipping limit (over saturation). See MCCELOG: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=904787 MOD 8/14/17 19:41 HVS Further investigation indicates that the saturation might be in the drive amplifier rather than the klystron itself. See MCCELOG: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=904788 | HARDWARE
136359 | 2017-08-14 | 1 | LI20 | 81 | 4 | Seeing increased jitter from 20-8 L0B modulator. Please check history of jitter PV SIOC:SYS0:ML00:AO216 and scope modulator. | HARDWARE
136366 | 2017-08-15 | 2 | LI26 | 61 | 4 | 26-6 VVS CB/ VVS13 is running hot (See pic). Needs to be inspected. | HARDWARE
136367 | 2017-08-15 | 1 | LI22 | 11 | 4 | Klystron 22-1 having numerous phase dropouts noticeable on FEL at low energy. | HARDWARE
136442 | 2017-08-25 | 1 | LI26 | 31 | 4 | 26-3 FLK, has an EOLC fault. | HARDWARE
136454 | 2017-08-26 | 2 | LI20 | 81 | 4 | Kly Mod 20-8 (L0B): station experiencing BV dropouts. | HARDWARE
135425 | 2017-06-01 |  | LI20 | 61 | 4 | GUN PAC LLRF Feedback (AMPL) is railed I have had to increase the MAX setpoint in the LLRF feedback from 5100 to 5800 to maintain 6MeV off the gun. New setpoint settles at 5420 counts but is now decreasing as I write this entry. it is down to 5300 now. | HARDWARE
135429 | 2017-06-02 | 1 | LI20 | 81 | 4 | L0B is MOD faulted. | HARDWARE
135429 | 2017-06-02 | 1 | LI20 | 81 | 4 | L0B is MOD faulted. | HARDWARE
136374 | 2017-08-16 | 1 | LI30 | 41 | 4 | 30-4 Has no core bias voltage. | HARDWARE
136379 | 2017-08-16 | 2 | LI24 | 31 | 4 | 24-3 DQing current can't be adjusted below 25 amps | HARDWARE
136466 | 2017-08-27 | 1 | LI20 | 71 | 4 | 20-7 Modulator problem found: 1. Primary snubber resistor R261: lug and wire should be replaced. 2. Pulse cable seems over heat, burnt the insulation. | HARDWARE
136468 | 2017-08-27 | 1 | LI26 | 71 | 4 | 26-7 Modulator problems found: 1. Primary snubber resistors: R260 & R261 lugs need to be replaced. R262 lug & wire need to be replaced. 2. T20 compartment: filter caps missing ground wires. 3. Xformer T2" Loose screw, see up load picture." | HARDWARE
136469 | 2017-08-27 | 0.8 | LI26 | 81 | 4 | 26-8 Modulator problems found: 1. Primary snubber resistors: R261 and R262 burnt lugs | HARDWARE
134621 | 2017-04-23 | 2 | LI24 | 81 | 4 | 24-8 Modulator problems found: 1. Wires touching the primary resistors body. 3. Oil leaks around the T-20 base and cables. 4. HV filter caps missing two ground wires. 5. Under Thyratron: two 115VAC covers are still unsafe. They need to be remake. 6. Door #7: the third 20 ohm carborundum res from left need to be replaced. | HARDWARE
134622 | 2017-04-23 | 1.5 | LI25 | 11 | 4 | 25-1 Modulator problems found: 1. Replace lug and wire for primary snubber res R261. 2. HV filter caps missing two ground wires. 4. Grid trigger spring coil is rusted. | HARDWARE
134623 | 2017-04-23 | 1 | LI25 | 21 | 4 | 25-2 Modulator problems found: 1. Primary snubber resistors: R260, R262 only lugs need to be replaced. R261 lug & wire need to be replaced. 3. HV filter caps missing two ground wires. 4. Terminal block TB13 firewire lug is corrosion. 5. T2 120VAC xformer on floor loose ground wire. 6. Door #5: #2AWG ground wire loose (top right) | HARDWARE
134624 | 2017-04-23 | 1 | LI25 | 31 | 4 | 25-3 Modulator problems found: 2. Primary snubber resistors: a)wire touching R260 body, reroute it away ; b)replace lugs for R261, R262 3. Door #4 top loose ground wires #2AWG. 4. HV filter caps missing two ground wires. | HARDWARE
135446 | 2017-06-02 | 1.5 | LI23 | 31 | 4 | Klys 23-3 seeing large (10 deg or more) phase jumps frequently. | HARDWARE
136475 | 2017-08-27 | 1 | LI27 | 31 | 4 | 27-3 KMC faulted. | HARDWARE
136476 | 2017-08-27 | 2 | LI21 | 41 | 4 | 21-4 MOD faulted with no remote reset. Appears to be a tripped fire wire. | HARDWARE
136486 | 2017-08-28 |  | LI28 | 31 | 4 | 28-3: Tripping on SCR klixon interlock fault. | HARDWARE
135523 | 2017-06-07 | 1 | LI25 | 41 | 4 | 25-4: KMC fault | HARDWARE
135530 | 2017-06-07 | 0.5 | LI21 | 41 | 4 | Klys 21-4 showing BVLT jitter. | HARDWARE
136498 | 2017-08-29 |  | LI20 | 71 | 4 | 20-7 (aka L0A): Water leaking from PAD hose connection. | HARDWARE
136500 | 2017-08-29 | 2 | LI27 | 11 | 4 | 27-1: Seeing some beam volts jitter. | HARDWARE
134613 | 2017-04-20 | 1 | LI24 | 61 | 4 | 24-6 Modulator problems found: 2. Primary resisistors: R261, R262 bunrt lugs & wires. 3. HV filter caps missing two ground wires to EG. 4. Cable K" from the thyratron connect to the keep alive is touching the metal sharp edge.
134514 | 2017-04-09 | 0.7 | LI23 | 51 | 4 | 23-5 Modulator problems found: 1. Primary snubber resistors: lugs and wires need to be replaced. 2. HV filter caps missing two ground wires. 3. 120Vac T2" xfomer: loose ground.
134899 | 2017-05-10 | 2 | LI20 | 61 | 4 | LI20 klys focus magnet power supplies reading 0 V for 20-6 and 20-7. PS shows ON (normal color) Resis = -1 ohms (purple) voltage - 0V (purple) Current = 14.8 amps (red) | HARDWARE
136220 | 2017-07-28 | 1 | LI24 | 11 | 4 | 24-1 tripped and will not reset. PEM to reset please. | HARDWARE
136227 | 2017-07-28 | 0.2 | LI23 | 21 | 4 | 23-2: Phase jitter above spec limit. Amplitude jitter very low. | HARDWARE
136508 | 2017-08-30 | 2 | LI26 | 71 | 4 | 26-7: Seeing lots of beam volts jitter. | HARDWARE
134394 | 2017-03-23 | 1 | LI21 | 51 | 4 | 21-5: Primary snubber resistors: R261, R262 wires and lugs need to be replaced. | HARDWARE
134395 | 2017-03-23 | 0.5 | LI21 | 61 | 4 | 21-6: Primary snubber resistors: R261 wire and lug burnt; R262 loose screw | HARDWARE
134397 | 2017-03-23 | 1.2 | LI21 | 81 | 4 | 21-8: Problems found: 1. 480VAC cables are touching the metal sharp edges of xfomer. 2. The snubber resistors: R261 and R262 have wires and lugs burnt. | HARDWARE
134397 | 2017-03-23 | 2 | LI21 | 81 | 4 | 21-8: Problems found: 1. 480VAC cables are touching the metal sharp edges of xfomer. 2. The snubber resistors: R261 and R262 have wires and lugs burnt. | HARDWARE
135553 | 2017-06-09 | 1 | LI26 | 51 | 4 | Klystron LI26-51 has amplitude and beam volt jitter that exceeds tolerance levels. | HARDWARE
135554 | 2017-06-09 | 1.5 | LI29 | 11 | 4 | 29-1: Remove thyratron for use at 20-6. Install new thyratron in this station. | HARDWARE
135559 | 2017-06-09 |  | LI25 | 51 | 4 | 25-5C Repair klystron ion pump electrical leakage | HARDWARE
135681 | 2017-06-18 | 1 | LI24 | 31 | 4 | 24-3 has mod faulted and cannot be reset remotely | HARDWARE
135084 | 2017-05-16 | 1 | LI20 | 61 | 4 | Frequent Gun dropouts. Klystron 20-6 is having periodic bouts of extreme amplitude dropouts. See E-Log entry: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=890229 | HARDWARE
136513 | 2017-08-31 | 1 | LI27 | 41 | 4 | 27-4 is MOD faulted and won't reset | HARDWARE
136518 | 2017-08-31 | 1 | LI27 | 31 | 4 | Klystron 27-3 KMC faulted. | HARDWARE
136523 | 2017-09-01 | 1 | LI30 | 61 | 4 | 30-6: Seeing lots of beam volts jitter. | HARDWARE
136524 | 2017-09-01 |  | LI23 | 31 | 4 | station 23-3 has bad Cable status Verified it wasn't the PAD or MKSU PIOP swap doesn't seem to have helped - tried two spares just in case, left second spare in place. - M.Harms 5-Sep-17 | HARDWARE
136524 | 2017-09-01 |  | LI23 | 31 | 4 | station 23-3 has bad Cable status Verified it wasn't the PAD or MKSU PIOP swap doesn't seem to have helped - tried two spares just in case, left second spare in place. - M.Harms 5-Sep-17 | HARDWARE
136525 | 2017-09-01 | 0.1 | LI30 | 61 | 4 | station 30-6 MOD faults reset, but only lasts a couple of minutes before faulting again | HARDWARE
136532 | 2017-09-01 | 0.5 | LI24 | 61 | 4 | station 24-6 Klystron Waveguide fault | HARDWARE
134741 | 2017-05-03 | 1 | LI30 | 51 | 4 | 30-5: Keep alive power supply has no output. | HARDWARE
134748 | 2017-05-04 |  | LI20 | 51 | 4 | 20-5 PIOP appears in-op. repair or replace | HARDWARE
134749 | 2017-05-04 |  | LI21 | 21 | 4 | 21-2 PIOP appears in-op. repair or replace | HARDWARE
134750 | 2017-05-04 |  | LI23 | 41 | 4 | 23-4 PIOP appears in-op. repair or replace | HARDWARE
135128 | 2017-05-18 | 2 | LI28 | 31 | 4 | 28-3 tripping on thyratron trigger faults | HARDWARE
135135 | 2017-05-18 | 2.5 | LI23 | 11 | 4 | 23-1 getting intermittent accelerate triggers. Seems to mostly not trigger correctly but sometimes does. | HARDWARE
135142 | 2017-05-19 | 1 | LI26 | 41 | 4 | 26-4 trips on EOLCs | HARDWARE
135142 | 2017-05-19 | 1 | LI26 | 41 | 4 | 26-4 trips on EOLCs | HARDWARE
134497 | 2017-04-06 | 1.2 | LI22 | 31 | 4 | 22-3 Modulator: Problems found: 1. Primary snubber resistors: R261 and R262 burnt lugs & wires. 2. Fire wire lugs that connected to TB13, the lug is corrosion. | HARDWARE
134498 | 2017-04-06 | 1.2 | LI22 | 41 | 4 | 22-4 Modulator: problems found: 1. Primary snubber resistors: R260, R261, and R262 burnt lugs/wires 2. Fire wires TB12: lug is corrosion (see picture). | HARDWARE
134500 | 2017-04-06 | 1.2 | LI22 | 61 | 4 | 22-6 Modulator, Problems found: 1. Primary snubber resistors: R261 & R262 burnt lugs/wires. | HARDWARE
134500 | 2017-04-06 | 1.2 | LI22 | 61 | 4 | 22-6 Modulator, Problems found: 1. Primary snubber resistors: R261 & R262 burnt lugs/wires. | HARDWARE
134501 | 2017-04-06 | 1.2 | LI22 | 71 | 4 | 22-7 Modulator: problems found: 1. Primary snubber resistors: R260, R261 and R262 lugs/wires burnt. 2. Despiking Coil insulation burnt. | HARDWARE
134510 | 2017-04-09 | 1 | LI23 | 11 | 4 | 23-1 Modulator problems found: 1. T-20 : oil leaks 2. Primary snubber resistors: R260, R261, R262 lugs and wires burnt 3. HV filter caps missing two ground wires. 4. Core Bias Assembly: the two wires cross the AC/DC current xductor are over heat and melted the plastic of the AC/DC current xdu5ctor. 5. De-Spring Coil cover is cracked. | HARDWARE
134511 | 2017-04-09 | 0.75 | LI23 | 21 | 4 | 23-2 Modulator problems found: 1. Primary snubber resistors: R261, R262 lugs and wires burnt. 2. HV filter caps are missing two ground wires. | HARDWARE
134778 | 2017-05-07 | 1.5 | LI25 | 51 | 4 | 25-5 Modulator problems found: 1. T-20 leaks oil on cables and the T-20 base. 2. Primary snubber resistors: R261, R262 lugs and wires burnt. 3. HV filter caps missing ground wires. 4. Door #2: top 600VAC holes = 2 1/2 inches should be covered. 5. PFN strap #7 top (LHS) loose, strap needs to be replaced. | HARDWARE
134780 | 2017-05-07 | 2 | LI25 | 61 | 4 | 25-6 Modulator problems found: 1. Primary snubber resistor R261: lug &wire burnt. 2. HV filter caps missing ground wires. 3. Door#2: top 600VAC holes = 2 1/2 inches should be covered. | HARDWARE
134781 | 2017-05-07 | 2 | LI25 | 81 | 4 | 25-8 Modulator problems found: 1. T2-xformer 120VAC ground wire loose. 2. T-20 base ground wire loose. 3. Primary snubber resistor R260 burnt lug. 4. HV filter cap missing ground wires. 5. Door#2: top 600VAC holes = 2 1/2 inches should be covered. 6. TB1 phoenix connector 120VAC black wires should be re-terminated and re-insert properly. I think it is a Core bias output wire, not 120 VAC -p.stiles | HARDWARE
134550 | 2017-04-13 | 1 | LI24 | 11 | 4 | 24-1 Modulator problems found: 1. Primary snubber resistors: R261, R262 Lugs/wires need to be replaced. | HARDWARE
134551 | 2017-04-13 | 2 | LI24 | 21 | 4 | 24-2 Modulator problems found: 1. Primary snubber resistors: R261, R262 lugs / wires need to be replaced. 2. Door #5: Top right corner - ground wire loose. 3. Despiking coil cover is cracked. | HARDWARE
134553 | 2017-04-13 | 1 | LI24 | 31 | 4 | 24-3 Modulator problems found: 1. Primary snubber resistors: R261 lug /wire needs to be replaced. 2. HV filter caps missing ground wires. 3. T2- 120VAC xformer base : loose ground wire. 4. EOLC stack missing mounting screw. 5. Cable is touching the despiking coil body. | HARDWARE
134553 | 2017-04-13 | 1 | LI24 | 31 | 4 | 24-3 Modulator problems found: 1. Primary snubber resistors: R261 lug /wire needs to be replaced. 2. HV filter caps missing ground wires. 3. T2- 120VAC xformer base : loose ground wire. 4. EOLC stack missing mounting screw. 5. Cable is touching the despiking coil body. | HARDWARE
134557 | 2017-04-14 | 0.5 | LI23 | 21 | 4 | 23-2 Has a klystron vacuum fault. | HARDWARE
134770 | 2017-05-05 | 1 | LI25 | 11 | 4 | 25-1 KMC fault. Trim p/s breaker trips. | HARDWARE
135159 | 2017-05-19 | 1 | LI25 | 41 | 4 | 25-4 Replace HV regulator board. | HARDWARE
135163 | 2017-05-19 |  | LI24 | 11 | 4 | 24-1 tripping regularly; have had to reset several times in last two hours. | HARDWARE
135163 | 2017-05-19 | 1.5 | LI24 | 11 | 4 | 24-1 tripping regularly; have had to reset several times in last two hours. | HARDWARE
134512 | 2017-04-09 | 0.5 | LI23 | 31 | 4 | 23-3 Moduator problems found: 1. Primary snubber resistor R260: burnt lug 2. HV filter caps missing two ground wires. | HARDWARE
134513 | 2017-04-09 | 0.5 | LI23 | 41 | 4 | 23-4 Modulator problems found: 1. T-20 ground loose screw. 2. Primary snubber resistor R260 burnt lug; R261, R262 burnt wires. 3. HV filter caps missing ground wires 4. Cable C" below Thyratron, the insulation is cracked." | HARDWARE
134516 | 2017-04-09 | 0.5 | LI23 | 71 | 4 | 23-7 Modulator problem found: 1. Primary snubber resistors R261, R262: lugs/wires burnt 2. HV filter caps missing two ground wires. | HARDWARE
134517 | 2017-04-09 | 0.8 | LI23 | 81 | 4 | 23-8 Modulator problems found: 1. Primary snubber resistor R260: lug burnt 2. HV filter caps missing two ground wires. 3. Door #5: the ground wire that located on top right corner loose. | HARDWARE
135289 | 2017-05-25 |  | LI27 | 41 | 4 | Klystron LI27-4 will not go back to SLED mode remotely after changing to non-SLED mode. | HARDWARE
135297 | 2017-05-26 | 3 | LI22 | 61 | 4 | Klys 22-6 experiencing frequent KLV faults. | HARDWARE
135297 | 2017-05-26 | 1.2 | LI22 | 61 | 4 | Klys 22-6 experiencing frequent KLV faults. | HARDWARE
135297 | 2017-05-26 |  | LI22 | 61 | 4 | Klys 22-6 experiencing frequent KLV faults. | HARDWARE
135242 | 2017-05-23 | 1.5 | LI25 | 11 | 4 | 25-1 MOD faulted | HARDWARE
135331 | 2017-05-27 | 0.5 | LI24 | 31 | 4 | 24-3 was tripping on Reflected Energy | HARDWARE
135332 | 2017-05-27 | 1 | LI25 | 71 | 4 | 25-7 Reflected energy fault | HARDWARE
135450 | 2017-06-02 | 2 | LI21 | 21 | 4 | L1X occasionally shows MKT which trips the beam. If we are lucky, the fault goes away in 5 minutes. We are not sure what is causing this fault. 6/5/2017 edit: although L1X stopped going MKT after swapping the MKSU and SCR, it is now doing so again (has gone MKT three times in last two hours, for roughly ~8minutes each time). | HARDWARE
135496 | 2017-06-06 | 1 | LI23 | 61 | 4 | Klystron LI23-61 has high beam volts jitter. | HARDWARE
135499 | 2017-06-06 | 1 | LI23 | 31 | 4 | 23-3 HVR reset . | HARDWARE
135510 | 2017-06-06 | 2.5 | LI26 | 61 | 4 | 26-6 is KHD | HARDWARE
135300 | 2017-05-26 | 0.7 | LI28 | 71 | 4 | staion 28-7 Control Power Fault | HARDWARE
134826 | 2017-05-09 |  | LI30 | 21 | 4 | 30-2: Touch panel display blanks out and goes back to "HOME" while doing the set-up. It happens quite often during the settings. | HARDWARE
135686 | 2017-06-18 | 1 | LI21 | 51 | 4 | 21-5 MOD fault, no reset: NOTE: problem is Fire Wire short - Modulator was VERY hot, I staged a high voltage rectifier at the modulator but was sent to SPEAR and didn't make it back. | HARDWARE
135712 | 2017-06-20 |  | LI21 | 21 | 4 | 21-2 (L1X) modulator HV has a strong 16.6 Hz timing jitter, see http://physics-elog.slac.stanford.edu/lclselog/show.jsp?dir=/2017/25/20.06&pos=2017-06-20T17:35:31 | HARDWARE
135713 | 2017-06-20 | 8 | LI20 | 81 | 4 | 20-8 L0B MOD fault. Attempts to restart greeted immediately with ACOC fault. | HARDWARE
135722 | 2017-06-21 |  | LI28 | 21 | 4 | 28-2 KMC fault. | HARDWARE
135314 | 2017-05-26 | 4 | LI24 | 31 | 4 | 24-3 trips immediately after turn on with HV overcurrent fault showing. | HARDWARE
135320 | 2017-05-26 | 1 | LI28 | 41 | 4 | 28-4 has MASSIVE beam volts jitter. | HARDWARE
135830 | 2017-06-29 | 1.3 | LI29 | 61 | 4 | 29-6 has intermittent BV dropouts. | HARDWARE
135845 | 2017-06-30 |  | LI20 | 51 | 4 | LI20 TCAV0 trip on WG1 last night 2140 hrs. | HARDWARE
134854 | 2017-05-10 |  | LI22 | 21 | 4 | 22-2 klystron focus supply does not respond to on/off commands, shows purple status on EPICS. | HARDWARE
134873 | 2017-05-10 |  | LI21 | 81 | 4 | 21-8 has a LVPS fault that won't clear | HARDWARE
134874 | 2017-05-10 |  | LI21 | 31 | 4 | 21-3 has a T20 fan fault that won't clear | HARDWARE
135731 | 2017-06-22 | 0.3 | LI20 | 61 | 4 | 20-6 (GUN): showing occasional bouts of beam volts jitter Scope hooked up. Thyratron replace with good broken-in one from 29-1 Running OK, not great. ACR happy. | HARDWARE
135920 | 2017-07-06 | 2.5 | LI23 | 61 | 4 | 23-6 ACOCs on turn on. | HARDWARE
135921 | 2017-07-06 | 1 | LI29 | 61 | 4 | 29-6: Shows AMJ. causing 30% jitter on FEL intensity. DEACT'd, pulse stability improved immediately. | HARDWARE
135924 | 2017-07-06 | 2 | LI30 | 41 | 4 | 30-4: Thyratron becoming unstable, nearing EOL. | HARDWARE
135926 | 2017-07-06 | 2 | LI26 | 11 | 4 | 26-1: Thyratron becoming unstable, maybe near EOL. | HARDWARE
135927 | 2017-07-06 | 2 | LI27 | 81 | 4 | 27-8: Thyratron becoming unstable, maybe near EOL. | HARDWARE
135943 | 2017-07-07 | 1 | LI26 | 61 | 4 | 26-6 modulator T20 has oil around the base. Inspect bushings (input and output) for sighs of new leakage. Inspect oil level. | HARDWARE
135948 | 2017-07-07 |  | LI25 | 11 | 4 | 25-1: Leaking oil at bottom of pulse cable and oil dropping on the floor from the klystron. | HARDWARE
135954 | 2017-07-08 | 0.5 | LI21 | 51 | 4 | 21-5 tripled off on MKT fault could not reset remotely or at the station. | HARDWARE
135965 | 2017-07-10 | 2.5 | LI23 | 71 | 4 | Klystron 23-7 has a lot of phase jitter and spikes on the beam volts | HARDWARE
135982 | 2017-07-11 | 1.5 | LI26 | 61 | 4 | Klys 26-6 has VVS breaker fault. Be good if it can get fixed before 09:00 on 7/11. | HARDWARE
136029 | 2017-07-14 | 0.5 | LI28 | 21 | 4 | 28-2: Thyratron Compartment Fan Fault | HARDWARE
136106 | 2017-07-20 | 2 | LI21 | 81 | 4 | 21-8 amplitude is dropping out very frequently. | HARDWARE
136106 | 2017-07-20 | 1.3 | LI21 | 81 | 4 | 21-8 amplitude is dropping out very frequently. | HARDWARE
136139 | 2017-07-22 | 0.5 | LI25 | 21 | 4 | 25-2: Station keeps dropping out on external interlock. follow up: water trap shows many ACC1 trips, meter checking -24v suggests very fast trip | HARDWARE
136139 | 2017-07-22 | 1 | LI25 | 21 | 4 | 25-2: Station keeps dropping out on external interlock. follow up: water trap shows many ACC1 trips, meter checking -24v suggests very fast trip | HARDWARE
136139 | 2017-07-22 |  | LI25 | 21 | 4 | 25-2: Station keeps dropping out on external interlock. follow up: water trap shows many ACC1 trips, meter checking -24v suggests very fast trip | HARDWARE
136140 | 2017-07-22 | 1 | LI29 | 51 | 4 | 29-5 has frequent Amplitude drop outs. Ranging does not cure. Res = 6.5 volts. Edit 7/22/2017- swing shift, 29-5 is now tripped and trips off immediately on overcurrent if we try to bring back on. | HARDWARE
136140 | 2017-07-22 | 3 | LI29 | 51 | 4 | 29-5 has frequent Amplitude drop outs. Ranging does not cure. Res = 6.5 volts. Edit 7/22/2017- swing shift, 29-5 is now tripped and trips off immediately on overcurrent if we try to bring back on. | HARDWARE
136140 | 2017-07-22 | 1 | LI29 | 51 | 4 | 29-5 has frequent Amplitude drop outs. Ranging does not cure. Res = 6.5 volts. Edit 7/22/2017- swing shift, 29-5 is now tripped and trips off immediately on overcurrent if we try to bring back on. | HARDWARE
136001 | 2017-07-12 | 1 | LI23 | 11 | 4 | Klystron 23-1 having bouts of serious amplitude problems. Beam volts and phase appear fine but amplitude jumps up problematically. | HARDWARE
136009 | 2017-07-13 | 0.5 | LI24 | 61 | 4 | Klystron 24-6 tripping frequently on temperature faults. KLT faults. Temperature readback seems to drift low. See ELOG entry: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=900623 | HARDWARE
136019 | 2017-07-13 | 0.5 | LI20 | 61 | 4 | station 20-6 (GUN) is tripped and showing VVS fault | HARDWARE
135882 | 2017-07-01 | 1 | LI30 | 41 | 4 | 30-4: DQ'g current dropped to <1A. | HARDWARE
135883 | 2017-07-01 | 1 | LI21 | 11 | 4 | LCLS INJ: 21-1 and 21-2 did not reset remotely after PPS fault. | HARDWARE
135983 | 2017-07-11 | 4 | LI28 | 41 | 4 | Klys 28-4 tripped on modulator control power fault. | HARDWARE
135992 | 2017-07-11 | 0.5 | LI21 | 51 | 4 | 21-5 is in MKSU Trigger Enable Fault and will not reset remotely. | HARDWARE
134930 | 2017-05-11 | 3 | LI20 | 71 | 4 | Kly Mod 20-7, L0A: begin modulator startup. Process new E2V CX1937AX thyratron for HV ops. | HARDWARE
134930 | 2017-05-11 | 2.5 | LI20 | 71 | 4 | Kly Mod 20-7, L0A: begin modulator startup. Process new E2V CX1937AX thyratron for HV ops. | HARDWARE
134930 | 2017-05-11 | 2.5 | LI20 | 71 | 4 | Kly Mod 20-7, L0A: begin modulator startup. Process new E2V CX1937AX thyratron for HV ops. | HARDWARE
136199 | 2017-07-26 | 0.5 | LI25 | 71 | 4 | 25-7 MOD fault, won't reset remotely. | HARDWARE
136202 | 2017-07-27 | 1 | LI28 | 41 | 4 | 28-4 has high BV jitter. | HARDWARE
136203 | 2017-07-27 | 1 | LI29 | 71 | 4 | 29-7 has high amplitude jitter. | HARDWARE
136208 | 2017-07-27 | 0.5 | LI22 | 21 | 4 | 22-2 Has an intermittent T-20 fan fault. | HARDWARE
136209 | 2017-07-27 | 1 | LI22 | 21 | 4 | 22-2 Outboard filter cap looks like is bulging. May need replace. | HARDWARE
136210 | 2017-07-27 | 1 | LI25 | 21 | 4 | Klystron 25-2 MOD faulted and won't reset remotely | HARDWARE
136211 | 2017-07-27 |  | LI25 | 21 | 4 | 25-2 Trips on WG1. | HARDWARE
136214 | 2017-07-27 | 1 | LI29 | 71 | 4 | 29-7 Has a PFN fan fault. | HARDWARE
136216 | 2017-07-27 | 0.5 | LI26 | 61 | 4 | Klystron 26-6 has a MOD VVS fault and cannot be reset remotely. Also, check the J-Box on top of the modulator. | HARDWARE
136022 | 2017-07-13 | 0.25 | LI20 | 61 | 4 | LCLS gun klystron (20-6) has a VVS voltage fault after site-wide power dip. | HARDWARE
136025 | 2017-07-13 | 2 | LI20 | 61 | 4 | 20-6: Gun beam voltage is extremely unstable after recovery from power dip. Edit 07:20 7/14/2017- The beam volts still become unstable; it seems to come in spurts. Station still needs attention. | HARDWARE
136026 | 2017-07-13 | 1 | LI20 | 81 | 4 | L0B MOD faulted, won't reset. | HARDWARE
136035 | 2017-07-14 | 1.5 | LI23 | 21 | 4 | 23-2: Thyratron nearing EOL. | HARDWARE
136036 | 2017-07-14 | 1 | LI29 | 41 | 4 | 29-4: Keep alive power supply went from 280mA to 430mA after recent POMM work. | HARDWARE
136038 | 2017-07-14 |  | LI23 | 11 | 4 | 23-1: Tune PFN for less slope and ripple. | HARDWARE
136101 | 2017-07-19 |  | LI28 | 21 | 4 | 28-2 tripped on KMC after about 5 minutes of running. | HARDWARE
136082 | 2017-07-19 |  | LI25 | 21 | 4 | Klystron Li25-21 is Mod faulted and will not reset remotely. | HARDWARE
136091 | 2017-07-19 | 0.5 | LI25 | 41 | 4 | 25-4 station appears to be running fine but the MOD PLC EPICs PVs are not connecting. | HARDWARE
136093 | 2017-07-19 |  | LI25 | 81 | 4 | 25-8 frequent MOD PLC external interlock faults and PIOP MKSU trigger enable faults | HARDWARE
134949 | 2017-05-12 | 1.5 | LI22 | 51 | 4 | Kly Mod 22-5 stuck in klystron heater delay for the past 18 hours. | HARDWARE
134954 | 2017-05-12 |  | LI22 | 41 | 4 | 22-4 has low beam volts. | HARDWARE
134956 | 2017-05-12 |  | LI25 | 11 | 4 | 25-1 has low beam volts. | HARDWARE
134957 | 2017-05-12 |  | LI26 | 11 | 4 | 26-1 has low beam volts. | HARDWARE
136122 | 2017-07-21 | 1.5 | LI25 | 51 | 4 | Many trips from 25-5 this shift. Nothing visible on the error counter. | HARDWARE
134966 | 2017-05-12 |  | LI21 | 11 | 4 | 21-1 (L1S) only reaches 14 MV before hitting LLRF amplitude feedback limits (needs to get up to about 115 MV). Waveforms from the PAD look strange. | HARDWARE
134967 | 2017-05-12 | 0.5 | LI30 | 81 | 4 | RF Station 30-8 not getting triggers--shows CAMAC error on CUD | HARDWARE
134986 | 2017-05-12 | 2 | LI22 | 61 | 4 | 22-6 has a tripped CB7 and Safety Relay & VVS faults indicated in plc fault log | HARDWARE
135627 | 2017-06-14 |  | LI22 | 11 | 4 | 22-1 modulator panel shows a dequing under current fault. | HARDWARE
135628 | 2017-06-14 | 1 | LI26 | 51 | 4 | 26-5 modulator panel shows a dequing under current fault. | HARDWARE
135636 | 2017-06-14 |  | LI21 | 21 | 4 | intermittent Amplitude spikes on L1X observed. Directly correlated with FEL intensity jitter. Not sure if HV or LLRF is the problem. see elog entry: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=896147 shows HV spikes correlate with bunch length spikes. tripping MPS pics. | HARDWARE
135646 | 2017-06-14 | 1.5 | LI24 | 61 | 4 | 24-6 MKT fault. | HARDWARE
136138 | 2017-07-22 | 1.5 | LI30 | 11 | 4 | 30-1: After HV regulator swap, DQ'g at 33A, amplitude and phase jitter very high. | HARDWARE
136138 | 2017-07-22 | 1 | LI30 | 11 | 4 | 30-1: After HV regulator swap, DQ'g at 33A, amplitude and phase jitter very high. | HARDWARE
136159 | 2017-07-25 | 2 | LI28 | 81 | 4 | 28-8 T20 input snubber assembly has several burnt wires that need to be replaced. | HARDWARE
136160 | 2017-07-25 | 2.2 | LI27 | 11 | 4 | 27-1 Modulator problems found: 1. Primary snubber resistors: R260, R261 wires need to be replaced. 2. HV filter 4.7uF cap missing ground wires. 3. Core Bias TB1: pheonix connector (P1) Blk wire, the strands protruding from the connection. 4. Under thyratron at right side , the 115VAC cover came off and moved near the thyratron blower. | HARDWARE
136161 | 2017-07-25 | 1.5 | LI28 | 31 | 4 | 28-3 Modulator EOLC stack needs to be replaced | HARDWARE
136164 | 2017-07-25 | 0.5 | LI23 | 31 | 4 | 23-3: Door #3 ground hook wire too long. | HARDWARE
136165 | 2017-07-25 | 25 | LI27 | 31 | 4 | 27-3 Modulator problems found: 1. T20: oil leak around the base and T-20 terminal (R3). 2. Primary snubber resistor R262 wire need to be replaced. 3. HV filter cap 4.7uF missing ground wires. 4. Core Bias: Pheonix connector (P2) orange clip that secured the connector is broken. 5. Thyratron compartment light out. 6. Keep alive wires hanging out. they need to be secured. | HARDWARE
136167 | 2017-07-25 | 2 | LI23 | 51 | 4 | Grounding nut needed on main filter cap on the following stations | HARDWARE
136167 | 2017-07-25 | 3 | LI23 | 51 | 4 | Grounding nut needed on main filter cap on the following stations | HARDWARE
136167 | 2017-07-25 | 3 | LI23 | 51 | 4 | Grounding nut needed on main filter cap on the following stations | HARDWARE
136167 | 2017-07-25 | 2 | LI23 | 51 | 4 | Grounding nut needed on main filter cap on the following stations | HARDWARE
136167 | 2017-07-25 | 2 | LI23 | 51 | 4 | Grounding nut needed on main filter cap on the following stations | HARDWARE
136168 | 2017-07-25 | 0.5 | LI27 | 41 | 4 | 27-4 Modulator problem found: 1. Grounding nut needed on main filter cap. | HARDWARE
136171 | 2017-07-25 | 1.2 | LI27 | 71 | 4 | 27-7 Modulator problems found: 1. Primary snubber caps: the jumpers on the caps should be replaced. 2. Core Bias: 120VAC Pheonix connector (P2) the orange clip is broken. 3. Under Thyratron compartment: the big white cable (H1) insulation cut open by the sharp metal edge. | HARDWARE
136172 | 2017-07-25 | 0.5 | LI27 | 81 | 4 | 27-8 Modulator problem found: HV filter cap 4.7uF missing ground wires. | HARDWARE
136172 | 2017-07-25 | 1 | LI27 | 81 | 4 | 27-8 Modulator problem found: HV filter cap 4.7uF missing ground wires. | HARDWARE
136103 | 2017-07-19 | 0.5 | LI25 | 21 | 4 | Klystron LI25-21 is tripped on Mod Fault and will not reset remotely. | HARDWARE
136128 | 2017-07-21 | 0.5 | LI25 | 21 | 4 | Klystron 25-2 MOD faulted and won't reset remotely | HARDWARE
134987 | 2017-05-12 | 1 | LI22 | 41 | 4 | 22-4 is MOD faulted - indicates ACOC | HARDWARE
134988 | 2017-05-12 |  | LI22 | 41 | 4 | 22-4 is MOD faulted - indicates ACOC | HARDWARE
134989 | 2017-05-12 | 1 | LI23 | 71 | 4 | 23-7 is faulting on klystron O/C & O/V faults | HARDWARE
134990 | 2017-05-12 | 2 | LI28 | 41 | 4 | 28-4 tripped LVPS, F12 on relay board showing open, | HARDWARE
135006 | 2017-05-13 | 0.5 | LI20 | 51 | 4 | station 20-5 PIOP looks bad | HARDWARE
135651 | 2017-06-15 |  | LI22 | 51 | 4 | station 22-5 faulted HVR multiple times. We see 19 AMPL and 2 REFL faults on the error display since 08:00 yesterday. | HARDWARE
135663 | 2017-06-15 | 2 | LI28 | 61 | 4 | Klystron 28-6 MOD faulted and, when it is reset remotely, immediately trips off again. | HARDWARE
135767 | 2017-06-24 | 0.1 | LI25 | 21 | 4 | 25-2 WGV faulted. | HARDWARE
135767 | 2017-06-24 | 2 | LI25 | 21 | 4 | 25-2 WGV faulted. | HARDWARE
135769 | 2017-06-24 |  | LI20 | 81 | 4 | 20-8 L0B trips on restart HVOV | HARDWARE
135774 | 2017-06-26 | 0.5 | LI28 | 31 | 4 | 28-3 is getting OV faults. CB 7 tripped. | HARDWARE
135781 | 2017-06-26 | 1.3 | LI28 | 41 | 4 | Klys 28-4 showing Beam Volts jitter. | HARDWARE
136178 | 2017-07-25 | 2 | LI28 | 21 | 4 | 28-2 has a KMC fault that will not clear even though focus current is at 15A. | HARDWARE
136192 | 2017-07-26 | 1.5 | LI20 | 71 | 4 | Klystron 20-7 (L0A) does not have PPS permit. | HARDWARE
136192 | 2017-07-26 | 0.2 | LI20 | 71 | 4 | Klystron 20-7 (L0A) does not have PPS permit. | HARDWARE
135020 | 2017-05-13 | 0.5 | LI27 | 51 | 4 | 25-7 Klystron magnet power supply has no output. | HARDWARE
135022 | 2017-05-13 | 3 | LI24 | 61 | 4 | Station 24-6 not staying on. Frequently HVR faults. | HARDWARE
135022 | 2017-05-13 | 1 | LI24 | 61 | 4 | Station 24-6 not staying on. Frequently HVR faults. | HARDWARE
135023 | 2017-05-13 | 2 | LI24 | 11 | 4 | Klys 24-1 is MOD; no remote reset | HARDWARE
135674 | 2017-06-16 | 3.5 | LI20 | 81 | 4 | Klystron 20-8 (L0B) MOD faulted and won't reset remotely | HARDWARE
135675 | 2017-06-17 | 2.2 | LI30 | 61 | 4 | Klys 30-6 keeps going HVR | HARDWARE
135677 | 2017-06-18 | 1 | LI22 | 41 | 4 | Klys 22-4 is tripping frequently on over voltage. | HARDWARE
135684 | 2017-06-18 | 0 | LI20 | 81 | 4 | L0B's modulator interlocks will sometimes only reset via the PPS reset button, even when the interlock is not PPS related. See attachment: 28-2,20-5,6,7,8,21-1,21-2,b921.pdf and HV_boards_upgrade.xlsx for list of modulators upgraded. | HARDWARE
135684 | 2017-06-18 |  | LI20 | 81 | 4 | L0B's modulator interlocks will sometimes only reset via the PPS reset button, even when the interlock is not PPS related. See attachment: 28-2,20-5,6,7,8,21-1,21-2,b921.pdf and HV_boards_upgrade.xlsx for list of modulators upgraded. | HARDWARE
135689 | 2017-06-19 |  | LI21 | 21 | 4 | L1X amplitude jitter high; varying significantly shot to shot. Don't see correlation to beam volts, do see correlation to drive. | HARDWARE
135787 | 2017-06-27 | 1.3 | LI27 | 81 | 4 | 27-8 has high BV jitter. | HARDWARE
135788 | 2017-06-27 | 1.3 | LI26 | 11 | 4 | 26-1 has high BV jitter. | HARDWARE
135789 | 2017-06-27 | 0.1 | LI26 | 81 | 4 | 26-8 fast time plots aren't working. Tried plotting amplitude & BV jitter and got unrealistic-looking plots that didn't change from one plot request to the next. I don't think fast time plotting issues reflect a modulator problem. p.stiles | HARDWARE
135789 | 2017-06-27 | 1 | LI26 | 81 | 4 | 26-8 fast time plots aren't working. Tried plotting amplitude & BV jitter and got unrealistic-looking plots that didn't change from one plot request to the next. I don't think fast time plotting issues reflect a modulator problem. p.stiles | HARDWARE
135790 | 2017-06-27 | 1.5 | LI30 | 31 | 4 | 30-3 Has intermittent BV dropouts. | HARDWARE
135244 | 2017-05-23 | 2 | LI27 | 71 | 4 | 27-7 drops to HVR / loses HV CMD signal to regulator board with no faults showing. | HARDWARE
135381 | 2017-05-30 | 1.5 | LI21 | 11 | 4 | L1S beam volts jitter causing 20% jitter on FEL. Correlated well with phase jitter | HARDWARE
135398 | 2017-05-31 | 2 | LI27 | 41 | 4 | Station trips once on HVOC and it goes into Mod Fault lock-out | HARDWARE
134358 | 2017-03-20 | 1.2 | LI21 | 41 | 4 | 21-4: Core Bias Assembly, 120Vac "P2" secure clip broken. | HARDWARE
135826 | 2017-06-29 | 1 | LI27 | 61 | 4 | 27-6 has intermittent large BV dropouts. | HARDWARE
135827 | 2017-06-29 | 1.3 | LI27 | 11 | 4 | 27-1 has high BV jitter. | HARDWARE
135828 | 2017-06-29 | 1.3 | LI26 | 11 | 4 | 26-1 has very high phase jitter. | HARDWARE
135829 | 2017-06-29 | 1.3 | LI29 | 51 | 4 | 29-5 has high phase jitter. | HARDWARE
134574 | 2017-04-17 |  | LI25 | 41 | 4 | 25-4 Modulator's PLC will not remember settings. Suspect PLC processor module. Replace if needed. | HARDWARE
134575 | 2017-04-17 | 1 | LI25 | 81 | 4 | 25-8 Modulator Fire Wire Fault. Replace faulty fire wire. | HARDWARE
134578 | 2017-04-17 | 1 | LI30 | 61 | 4 | 30-6 modulator: Does not show an External Interlock. It should because MKSU is showing a KMC fault. | HARDWARE
135438 | 2017-06-02 | 3 | LI20 | 81 | 4 | Klystron 20-8 MOD faulted and won't reset remotely. | HARDWARE
134690 | 2017-05-02 | 1 | LI29 | 71 | 4 | 29-7 mod. has a PFN rooftop fan fault. | HARDWARE
135832 | 2017-06-29 | 1.3 | LI22 | 81 | 4 | 22-8 has intermittent single shot phase jumps. | HARDWARE
135834 | 2017-06-29 | 1.3 | LI28 | 31 | 4 | Klystron 28-3 MOD faulted with no remote reset. | HARDWARE
135834 | 2017-06-29 | 1.3 | LI28 | 31 | 4 | Klystron 28-3 MOD faulted with no remote reset. | HARDWARE
135843 | 2017-06-30 | 0.1 | LI26 | 31 | 4 | 26-3 has a Klystron Temperature fault. | HARDWARE
135464 | 2017-06-03 | 5 | LI24 | 81 | 4 | Kly Mod 24-8: install grid divider with grid wire unistalled. Attach BNC cable 143 to divider output to monitor PFN/Anode volts at PFN volts jack (J8) on front panel. Also, adjust set point and reservoir for stable mod ops @ 350 KV beam volts. | HARDWARE
135465 | 2017-06-03 | 2 | LI26 | 41 | 4 | 26-4 tripping off frequently. | HARDWARE
135465 | 2017-06-03 | 0.5 | LI26 | 41 | 4 | 26-4 tripping off frequently. | HARDWARE
135470 | 2017-06-04 | 2 | LI26 | 41 | 4 | Klystron LI26-41 is tripping off on EVOC faults. | HARDWARE
135470 | 2017-06-04 |  | LI26 | 41 | 4 | Klystron LI26-41 is tripping off on EVOC faults. | HARDWARE
135473 | 2017-06-04 | 0.5 | LI28 | 71 | 4 | station 28-7 Mod faulted-won't stay on after reset. | HARDWARE
134742 | 2017-05-03 | 1.5 | LI30 | 81 | 4 | 30-8: DQ'g divider has slow oil leak. | HARDWARE
135211 | 2017-05-22 | 1 | LI28 | 41 | 4 | 28-4 is faulted on KMC | HARDWARE
136194 | 2017-07-26 | 1.5 | LI30 | 81 | 4 | 30-8 Modulator shows HV but no current. Triggers are present. | HARDWARE
135474 | 2017-06-04 | 1 | LI26 | 61 | 4 | 26-6 Core Bias supply unstable, 15.0 amps +/- 0.1 amp at 5.65 volts | HARDWARE
135481 | 2017-06-05 | 1 | LI22 | 31 | 4 | 22-3 is MOD fault and won't reset | HARDWARE
135482 | 2017-06-05 | 1 | LI28 | 41 | 4 | 28-4 is MOD fault and won't reset | HARDWARE
135488 | 2017-06-05 | 2 | LI21 | 21 | 4 | 21-2 no focusing current read back on EPICS. SAM Module MODU:LI21:208-LI21 Cr 2 M08 - S21-F2A (21-2) shows zeros on EPICS. OK LED not flashing on the module. Suspect SAM module. | HARDWARE
135491 | 2017-06-05 | 1 | LI26 | 31 | 4 | Klystron 26-3 is fault lockout. Cannot reset remotely. | HARDWARE
135497 | 2017-06-06 | 1 | LI23 | 41 | 4 | Klystron LI23-41 has high amplitude jitter. | HARDWARE
135047 | 2017-05-14 |  | LI23 | 71 | 4 | Klystron 23-7 tripping off. Overvoltage errors, trips immediately when mod is on. | HARDWARE
135077 | 2017-05-16 | 3 | LI20 | 81 | 4 | Kly Mod 20-8 (L0B) has excessive phase and amplitude jitter. | HARDWARE
135086 | 2017-05-17 |  | LI26 | 61 | 4 | No 110vac control power to modulator 26-6 although feeder breaker at panel 2P26G1B - #2 is on, not tripped. Cycling breaker #2 has no effect. Other circuits in panel are ok. | HARDWARE
135917 | 2017-07-05 | 0.5 | LI28 | 41 | 4 | 28-4: Shows AMJ. causing 30% jitter on FEL intensity. DEACT'd, pulse stability improved immediately. | HARDWARE
134688 | 2017-05-02 | 2 | LI20 | 71 | 4 | Kly Mod 20-7 (L0A): turn on and adjust E2V thyratron filament and reservoir heater ckts. Adjust until there is stable 6.3 VAC on both heater ckts. | HARDWARE
134688 | 2017-05-02 | 3 | LI20 | 71 | 4 | Kly Mod 20-7 (L0A): turn on and adjust E2V thyratron filament and reservoir heater ckts. Adjust until there is stable 6.3 VAC on both heater ckts. | HARDWARE
134689 | 2017-05-02 | 1 | LI29 | 51 | 4 | 29-5 Mod. has a Thy Fan Flt. | HARDWARE
135511 | 2017-06-06 | 0 | LI21 | 51 | 4 | 21-5 having multiple Mod faults (water faults) | HARDWARE
135640 | 2017-06-14 |  | LI20 | 81 | 4 | L0B tripped, required PEM to lower res voltage to make stable. (-.5H) | HARDWARE
135096 | 2017-05-17 | 1 | LI29 | 71 | 4 | 29-7 tripping on thyratron filament O/C | HARDWARE
135104 | 2017-05-17 |  | LI27 | 51 | 4 | 27-5 drops to HVR / loses HV CMD signal to regulator board with no faults showing. | HARDWARE
135106 | 2017-05-17 | 2 | LI30 | 61 | 4 | 30-6 KMC | HARDWARE
135108 | 2017-05-17 |  | LI25 | 41 | 4 | TBR 25-4 Suspected klystron magnet short. | HARDWARE
135933 | 2017-07-06 | 1.5 | LI26 | 61 | 4 | 26-6 tripped on MOD fault; can't reset. | HARDWARE
135933 | 2017-07-06 | 2 | LI26 | 61 | 4 | 26-6 tripped on MOD fault; can't reset. | HARDWARE
134705 | 2017-05-02 | 1 | LI25 | 41 | 4 | 25-4 modulator PLC has zeros for all input values | HARDWARE
134708 | 2017-05-02 | 1 | LI20 | 71 | 4 | 20-7 Trim P/S Genesys needs replace. No output. | HARDWARE
134709 | 2017-05-02 | 1 | LI22 | 41 | 4 | 22-4 Trim P/S Display is flashing. Touch panel shows no output from unit. | HARDWARE
134710 | 2017-05-02 | 1 | LI22 | 71 | 4 | 22-7 Trim P/S Display is flashing. Touch panel shows no output from unit. | HARDWARE
136899 | 2017-09-28 | 0.5 | LI23 | 11 | 4 | 23-1 poor focus regulation see ACR ELOG entry: https://mccelog.slac.stanford.edu/elog/wbin/elog_item.php?elog_id=911103 | HARDWARE
136684 | 2017-09-13 |  | LI20 | 81 | 4 | RFS 20-8 L0B after 9/13/17 POMM 80kV thyratron replacement, we are seeing frequent instances of no HV for one pulse then slow (several pulses) recovery to normal. No interlock faults seen. | HARDWARE
136685 | 2017-09-13 | 1 | LI25 | 71 | 4 | 25-7: Thyratron nearing EOL. | HARDWARE
136692 | 2017-09-13 |  | LI20 | 81 | 4 | The symptom observed is a non-latching MPS fault. This fault was also observed when TD-11 is in. It is suspected that these faults are due to a combination of L0B, and a yet to be identified L2 station. Another issue is that the MPS fault does not persist for very long, and is not logged in the MPS fault log in the MPS GUI. | HARDWARE
136766 | 2017-09-20 |  | LI30 | 31 | 4 | 30-3 WGV fault. | HARDWARE
136804 | 2017-09-22 | 1.5 | LI23 | 81 | 4 | 23-8 has been in KHD for over an hour. | HARDWARE
136679 | 2017-09-13 |  | LI23 | 51 | 4 | 23-5 FLK attempts to reset it trips again on External Interlock lockout. I saw some PIOP messages during turn on attempt indicating klys over voltage. | HARDWARE
136700 | 2017-09-14 |  | LI20 | 81 | 4 | L0B showing frequent BVLT dropouts | HARDWARE
136716 | 2017-09-15 |  | LI20 | 81 | 4 | LOB (20-8) repeatedly tipping off on ACOC faults. Plan per PD and Axel- swap thyratron during PAMM next week, run with lower amplitude until then. | HARDWARE
136729 | 2017-09-17 |  | LI22 | 51 | 4 | 22-5 is experiencing frequent BV dropouts. | HARDWARE
136729 | 2017-09-17 | 0.5 | LI22 | 51 | 4 | 22-5 is experiencing frequent BV dropouts. | HARDWARE
136813 | 2017-09-23 | 3 | LI24 | 21 | 4 | 24-2: MOD faulted and will not reset remotely. Station getting no triggers. That problem fixed, still has instabilities. | HARDWARE
136815 | 2017-09-23 | 2 | LI20 | 81 | 4 | Kly Mod 20-8 (L0B): station again experiencing BV drop outs. | HARDWARE
136818 | 2017-09-24 |  | LI23 | 81 | 4 | station 27-8 had at least 8 AMPL faults in the last 15 minutes. | HARDWARE
136820 | 2017-09-24 | 4 | LI30 | 61 | 4 | station 30-6 MOD fault--won't reset remotely | HARDWARE
136821 | 2017-09-24 | 2 | LI27 | 81 | 4 | 27-8 ACOC interlocks. | HARDWARE
136821 | 2017-09-24 | 0.8 | LI27 | 81 | 4 | 27-8 ACOC interlocks. | HARDWARE
136825 | 2017-09-25 | 1 | LI21 | 81 | 4 | 21-8 MOD fault, no reset | HARDWARE
136798 | 2017-09-21 | 3 | LI26 | 81 | 4 | Klystron LI26-81 goes Mod Fault immediately after resetting. | HARDWARE
136838 | 2017-09-25 | 0.5 | LI27 | 61 | 4 | 27-6 reporting rapid trips on WA#1 flowswitch | HARDWARE
136868 | 2017-09-27 | 3 | LI20 | 81 | 4 | Kly Mod 20-8 (L0B): install 650 CFM fan pack on lower door vent of modulator thyratron cabinet. | HARDWARE
136831 | 2017-09-25 | 2 | LI24 | 21 | 4 | 24-2 at 300kv, DQing current = 265 amps indicated on PLC | HARDWARE
136839 | 2017-09-25 | 0.5 | LI27 | 61 | 4 | station 27-6 shows WA1 faults. Error display shows 1640 water faults since 08:00 this morning. | HARDWARE
136842 | 2017-09-26 | 0.5 | LI28 | 31 | 4 | Klys 28-3 is MOD; no remote reset. Ops tried resetting locally, but we know less than PEM about how to do that, so please draw no conclusions from us :) | HARDWARE
136843 | 2017-09-26 | 0.5 | LI22 | 81 | 4 | Klystron 22-8 experienced quite a few trips in a very short time span and then wouldn't reset | HARDWARE
136843 | 2017-09-26 | 0.5 | LI22 | 81 | 4 | Klystron 22-8 experienced quite a few trips in a very short time span and then wouldn't reset | HARDWARE
136844 | 2017-09-26 | 0.5 | LI24 | 21 | 4 | 24-2 had the MKSU changed and also the Beam voltage cable changed. Need to do PIOP ratios and change cable if necessary. | HARDWARE
136904 | 2017-09-29 | 2 | LI30 | 81 | 4 | 30-6 Modulator: replace thyratron. | HARDWARE
136859 | 2017-09-27 | 2 | LI28 | 31 | 4 | Klys 28-3 is MOD with no remote reset | HARDWARE
136860 | 2017-09-27 | 1 | LI22 | 81 | 4 | Klystron LI22-81 is Mod Faulted and won't reset remotely. Klystron had been dropping out sporadically before faulting off. Possible water problem; thought water summary fault flashed before tripping. Update: water kludge added and indicated an ACC1 fault - need flow / flow switch checked | HARDWARE
136862 | 2017-09-27 | 0 | LI27 | 81 | 4 | Modulator 27-8 MOD faults frequently. We are able to reset it, but it trips off again less than a minute later. | HARDWARE
136878 | 2017-09-28 | 0 | LI30 | 61 | 4 | Klystron LI30-61 has tripped off 6 times over the last shift and is now not able to reset remotely. Errors panel does not show reason for faulting. | HARDWARE
136741 | 2017-09-19 | 0.5 | LI22 | 51 | 4 | 22-5 PHJ due to near zero crossing on 360 scan | HARDWARE
136751 | 2017-09-20 |  | LI26 | 21 | 4 | 26-2 CAMAC blower chassis fan is noisy. | HARDWARE
136754 | 2017-09-20 | 0.5 | LI24 | 11 | 4 | 24-1: SCR snubbers have open resistors on phase A and phase C. | HARDWARE
136755 | 2017-09-20 | 0.5 | LI24 | 21 | 4 | 24-2: SCR snubbers have open resistor on phase A and blown resistor on phase C. | HARDWARE
136756 | 2017-09-20 | 0.8 | LI24 | 31 | 4 | 24-3: SCR snubbers have open resistors on phase B and phase C, and phase A board burnt. | HARDWARE
136730 | 2017-09-17 | 0.5 | LI28 | 21 | 4 | 28-2 has beam voltage jitter. | HARDWARE
136759 | 2017-09-20 | 1 | LI24 | 81 | 4 | 24-8: SCR snubbers have open resistors on phase B and phase C. | HARDWARE
136767 | 2017-09-20 | 3.5 | LI22 | 11 | 4 | 22-1 Modultor problems found: 1. Primary snubber resistors: R260, R262 lugs need to be replaced. A 2. Current transducer case of Core Bias Assembly is burnt. 3. AC wires and thyratron cables should be separated. Wires insulation burnt. 4. Cover for 120VAC should be secured with tie-wrap. | HARDWARE
136768 | 2017-09-20 | 2.5 | LI22 | 21 | 4 | 22-2 Modulator problems found: 1. Ground hook holder on the PFN door need to be replaced. 2. EOLC stack is missing the stand-off (see attached picture). 3. All ground hooks in this modulator, screw that hold cable and ground hook together need a lock washer to keep it tighten. | HARDWARE
136787 | 2017-09-21 | 1 | LI30 | 81 | 4 | 30-8: MOD faults, ACOC's, no reset. | HARDWARE
136788 | 2017-09-21 | 1 | LI23 | 81 | 4 | 23-8: Modulator still going into KHD. Klystron voltage dropping out, current staying steady. | HARDWARE
136624 | 2017-09-10 | 2 | LI20 | 81 | 4 | Klys 20-8 (L0B) showing beam volts dropouts. See 136243. | HARDWARE
136662 | 2017-09-13 |  | LI21 | 11 | 4 | Klystron 21-1 FLK and won't reset remotely | HARDWARE
136667 | 2017-09-13 | 1 | LI20 | 71 | 4 | Mod 20-7 KLT ( klystron delta temperature) fault | HARDWARE
136668 | 2017-09-13 |  | LI22 | 81 | 4 | 22-8 waveguide vacuum fault | HARDWARE
136669 | 2017-09-13 | 1 | LI23 | 31 | 4 | Modulator 23-3 has ACOC faults even when modulator high voltage is lowered to 1 kV or modulator triggers are removed. | HARDWARE
136671 | 2017-09-13 | 1 | LI22 | 81 | 4 | 22-8 tripped on WGV, wont reset. | HARDWARE
136672 | 2017-09-13 |  | LI24 | 31 | 4 | Klystron 24-3 MOD faulting frequently | HARDWARE
136911 | 2017-09-29 | 4 | LI28 | 31 | 4 | Modulator 28-3: from ACR it appears CB7 has again tripped. | HARDWARE
136913 | 2017-09-30 | 4 | LI25 | 51 | 4 | Klystron LI25-51 is tripped on "Mod fault" and will not reset remotely. | HARDWARE
136640 | 2017-09-12 | 2 | LI20 | 81 | 4 | 20-8 L0B several (3) HVOV trips in last half hour. | HARDWARE
136701 | 2017-09-14 | 3 | LI28 | 31 | 4 | 28-3 MOD fault with no remote reset. | HARDWARE
136710 | 2017-09-14 | 3 | LI23 | 81 | 4 | Klys 23-8 is on mystery KHD | HARDWARE
136727 | 2017-09-16 | 1 | LI30 | 71 | 4 | 30-7: MOD fault, tripping off randomly. | HARDWARE
136727 | 2017-09-16 | 0.5 | LI30 | 71 | 4 | 30-7: MOD fault, tripping off randomly. | HARDWARE
136801 | 2017-09-22 | 0.5 | LI26 | 31 | 4 | 26-3 has a KLT fault. | HARDWARE
136802 | 2017-09-22 | 0.2 | LI25 | 71 | 4 | 25-7 is FLK. | HARDWARE
136803 | 2017-09-22 | 1 | LI21 | 81 | 4 | Klystron 21-8 was experiencing frequent amplitude dropouts | HARDWARE
136808 | 2017-09-22 | 2 | LI20 | 81 | 4 | Klystron 20-8 frequent beam volt dropouts. See 136814. | HARDWARE